## Data
SP METROLOGY SYSTEM THAILAND | [9 empty] | Certificate No. : | [14 empty] | Page 1 of 1
[10 empty] | Receive Date : |  |  |  | 2016-01-11 |  |  |  |  | Calibration Date : |  |  |  | 2016-01-11
Equipment Name | [9 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [9 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Dial Thermometer | [6 empty] | Manufacturer : |  |  |  | SEKI | [5 empty] | Model : |  | T100
Serial No. : |  | 45678 | [6 empty] | ID No : |  | N/A |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution : |  |  | 0.1
Overall Inspection | [5 empty] | Good |  |  | Not Good | [16 empty] | 0.001
Referance Standard | [5 empty] | : | SP-ST-002 | [7 empty] | Due Date : |  | 2016-01-11 | [9 empty] | 0.1
Referance Standard | [5 empty] | : | [8 empty] | Due Date :
Calibration Procedure | [5 empty] | : | SP-CPT-04-02
| Nominal |  | Number of Measure |  |  | Actual Value | [9 empty] | Average STD |  |  | Average UUC |  |  | Error |  | Repeatability
| Value |  |  |  |  | STD |  | UUC |  |  | UUC |  | STD | [10 empty] | STD |  | UUC
| -15 |  | 1st |  |  | 0 |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 3rd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 4th |  |  | 0 |  | 0 |  |  | 0 |  | 0
| 0 |  | 1st |  |  | 0 |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 3rd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 4th |  |  | 0 |  | 0 |  |  | 0 |  | 0
| 15 |  | 1st |  |  | 0 |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 3rd |  |  | 0 |  | 0 |  |  | 0 |  | 0
|  |  | 4th |  |  | 0 |  | 0 |  |  | 0 |  | 0
| 350 |  | 1st |  |  | 350 |  | 350 |  |  | 350 |  | 350 |  |  | 350 |  |  | 350 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 350 |  | 350 |  |  | 350 |  | 350
|  |  | 3rd |  |  | 350 |  | 350 |  |  | 350 |  | 350
|  |  | 4th |  |  | 350 |  | 350 |  |  | 350 |  | 350
| 450 |  | 1st |  |  | 450 |  | 450 |  |  | 450 |  | 450 |  |  | 450 |  |  | 450 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 450 |  | 450 |  |  | 450 |  | 450
|  |  | 3rd |  |  | 450 |  | 450 |  |  | 450 |  | 450
|  |  | 4th |  |  | 450 |  | 450 |  |  | 450 |  | 450
Short Teram
[17 empty] | UUC 1 |  |  |  | UUC 2
[17 empty] | 0 |  |  |  | 0
0
Calibrated By : |  |  |  | Mr.Nirut Loha
|  |  | 3 |  | Mr.Nirut Loha
|  |  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | 0 | [6 empty] | Page : 1 of 3
| Customer |  |  |  | : | SP
0
| Equipment Name |  |  |  | : | Dial Thermometer
| Manufacturer |  |  |  | : | SEKI
| Model |  |  |  | : | T100
| Serial Number |  |  |  | : | 45678
| ID. Number |  |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  |  |  | : | 2016-01-11
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-01-11
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2017-01-10
| Calibration Procedure |  |  |  |  | : | SP-CPT-04-02
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-12 |  |  |  | Approved by :
| Calibrated by |  | : | Mr.Nirut Loha |  |  |  | 3 |  |  |  | ( Mr. Santi Hankitudomsuk )
Authorized Signatory
[10 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q38: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | 0 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  | Model |  | Serial No. | Certificate No. | Due. Date
| Digital Thermometer |  | GT11/3850-40-392 |  | 08000098/100288 | PSL T 572/58 | 2017-06-05
with PRT Sensor
| Digital Thermometer |  | 34970A |  | MY44010293 | TC 15/0326 | 2016-11-02
with SRT Sensor
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  | 0 | [6 empty] | Page : 3 of 3
[9 empty] | Unit : | °C
|  |  | Setting Value |  | Standard Reading | UUC Reading | Error | Uncertainty ( ± )
[12 empty] | UUC | Range | CMC
|  |  | -15 |  | 0 | 0 | 0 | 0.07 |  |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 0 |  | 0 | 0 | 0 | 0.07 |  |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 15 |  | 0 | 0 | 0 | 0.7 |  |  |  | 0.066 | 250 to <500 | 0.7
|  |  | 350 |  | 350 | 350 | 0 | 0.71 |  |  |  | 0.71 | 250 to <500 | 0.7
|  |  | 450 |  | 450 | 450 | 0 | 0.71 |  |  |  | 0.71 | 250 to <500 | 0.7
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -
Mr.Chainarong Matchayamat

## Uncertainty Budget
Uncertainty Budget Dial Thermometer & Temperature Gauge (-40 to 650°C)
Unit: °C
Nominal Value | Uncertainty of STD |  | Short term stability of UUC |  | Resolution of STD |  | Resolution of UUC |  | Repeatability STD |  | Repeatability UUC |  | Uc | Ui | Veff | K95 | U95%
°C | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
-15 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
0 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
15 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
350 | 0.7 | 0.35 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.35119 | 0 | ∞ | 2 | 0.70238
450 | 0.7 | 0.35 | 0 | 0 | 0.05 | 0.028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.35237 | 0 | ∞ | 2 | 0.70475
Note on C4: Certificate of Calubration Digithermometer with PRT senror
Note on E4: ได้จากค่า ผลการสอบเทียบที่จุดสอบเทียบแรกครั้งที่1 - ผลการสอบเทียบที่จุดสอบเทียบแรกครั้งที่2
Note on G4: ค่าความละเอียดในการอ่านของ STD = 0.001 C
Note on I4: ค่าความละเอียดในการอ่านของ UUC
Note on D5: Divisor = 2
Note on F5: Divisor = √3
Note on H5: Divisor = √3
Note on J5: Divisor = √3
Note on L5: Divisor = 1
Note on N5: Divisor = 1

## UCB New '2016
ใบแบบการประมาณค่าความไม่แน่นอนของการวัด
Function : | Temperature |  |  | Testpoint | -15 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 0 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 15 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 350 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.7 |  | Normal |  | 2 | 1 |  | 0.35 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.35119 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.71 | ¥
2
Function : | Temperature |  |  | Testpoint | 450 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.7 |  | Normal |  | 2 | 1 |  | 0.35 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.35237 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.71 | ¥
2
2

## Cert of STD
Digital Thermometer with TC PRT | [5 empty] | Digital Thermometer with TC SRT
SP-ST-002
Due Date |  |  | 2016-06-05 |  |  | SP-ST-003
Point | Uncert |  | Point | Uncert |  | Due Date | 2016-11-02
°C | CH 1 |  | °C | CH 2 |  | Point | Uncert
-40.0092 | 0.03 | °C | -80.0236 | 0.03 | °C | °C
-0.00145 | 0.03 | °C | -40.0092 | 0.03 | °C | 250 | 0.7 | °C
50.0427 | 0.03 | °C | -0.00145 | 0.03 | °C | 500 | 0.7 | °C
100.0512 | 0.03 | °C | 50.0427 | 0.03 | °C | 750 | 2.6 | °C
150.0512 | 0.03 | °C | 100.0518 | 0.03 | °C | 1200 | 2.6 | °C
200.0524 | 0.03 | °C
250.0772 | 0.03 | °C
Note on B6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I7: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B8: Nathaphol Boonmee: Normal Value
Note on C8: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on F8: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I9: Nathaphol Boonmee: Normal Value
Note on J9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง