## Data Record(Forward)
SP METROLOGY SYSTEM THAILAND | [5 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [6 empty] | Page |  | 1 | of | 2
[6 empty] | Receive Date : |  |  |  | 2016-01-12 |  |  |  |  | Calibration Date : |  |  | 2016-01-13
Equipment Name | [5 empty] | Temp & Humiduty : |  |  |  |  | 20 |  | oC | 50 |  | %RH
Dial Gauge Tester/ Calibration Tester | [5 empty] | Location |  |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  | Dial Gauge Tester/ Calibration Tester | [8 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  | 123 |  |  |  | Serial No. : |  |  | 456 |  |  |  |  | ID No : |  | 789
Range : |  | 0 | to | 10 | mm | Resolution : |  |  | 0.001 |  | mm
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : |  |  |  | SP-SD-028-1 | [8 empty] | Due Date :
1. Flatness Measurement : | [5 empty] | µm
2. Function Measurement
| Norminal Value |  | STD Reading (Forward) | [9 empty] | Average |  |  | Error |  | Repeatability
|  |  | X1 |  | X2 |  | X3 |  |  | X4
| 0 |  | 0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0
| 0.01 |  | 1 |  | 1 |  | 1 |  |  | 1 |  |  | 1 |  |  | 0.99 |  | 0
| 0.02 |  | 2 |  | 2 |  | 2 |  |  | 2 |  |  | 2 |  |  | 1.98 |  | 0
| 0.03 |  | 3 |  | 3 |  | 3 |  |  | 3 |  |  | 3 |  |  | 2.97 |  | 0
| 0.04 |  | 4 |  | 4 |  | 4 |  |  | 4 |  |  | 4 |  |  | 3.96 |  | 0
| 0.05 |  | 5 |  | 5 |  | 5 |  |  | 5 |  |  | 5 |  |  | 4.95 |  | 0
| 0.06 |  | 6 |  | 6 |  | 6 |  |  | 6 |  |  | 6 |  |  | 5.94 |  | 0
| 0.07 |  | 7 |  | 7 |  | 7 |  |  | 7 |  |  | 7 |  |  | 6.93 |  | 0
| 0.08 |  | 8 |  | 8 |  | 8 |  |  | 8 |  |  | 8 |  |  | 7.92 |  | 0
| 0.09 |  | 9 |  | 9 |  | 9 |  |  | 9 |  |  | 9 |  |  | 8.91 |  | 0
| 0.1 |  | 10 |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 9.9 |  | 0
| 0.11 |  | 0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | -0.11 |  | 0
| 0.12 |  | 1 |  | 1 |  | 1 |  |  | 1 |  |  | 1 |  |  | 0.88 |  | 0
| 0.13 |  | 2 |  | 2 |  | 2 |  |  | 2 |  |  | 2 |  |  | 1.87 |  | 0
| 0.14 |  | 3 |  | 3 |  | 3 |  |  | 3 |  |  | 3 |  |  | 2.86 |  | 0
| 0.15 |  | 4 |  | 4 |  | 4 |  |  | 4 |  |  | 4 |  |  | 3.85 |  | 0
| 0.16 |  | 5 |  | 5 |  | 5 |  |  | 5 |  |  | 5 |  |  | 4.84 |  | 0
| 0.17 |  | 6 |  | 6 |  | 6 |  |  | 6 |  |  | 6 |  |  | 5.83 |  | 0
| 0.18 |  | 7 |  | 7 |  | 7 |  |  | 7 |  |  | 7 |  |  | 6.82 |  | 0
| 0.19 |  | 8 |  | 8 |  | 8 |  |  | 8 |  |  | 8 |  |  | 7.81 |  | 0
| 0.2 |  | 9 |  | 9 |  | 9 |  |  | 9 |  |  | 9 |  |  | 8.8 |  | 0
| 0.3 |  | 10 |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 9.7 |  | 0
| 0.4 |  | 11 |  | 11 |  | 11 |  |  | 11 |  |  | 11 |  |  | 10.6 |  | 0
| 0.5 |  | 12 |  | 12 |  | 12 |  |  | 12 |  |  | 12 |  |  | 11.5 |  | 0
| 0.6 |  | 13 |  | 13 |  | 13 |  |  | 13 |  |  | 13 |  |  | 12.4 |  | 0
| 0.7 |  | 14 |  | 14 |  | 14 |  |  | 14 |  |  | 14 |  |  | 13.3 |  | 0
| 0.8 |  | 15 |  | 15 |  | 15 |  |  | 15 |  |  | 15 |  |  | 14.2 |  | 0
| 0.9 |  | 16 |  | 16 |  | 16 |  |  | 16 |  |  | 16 |  |  | 15.1 |  | 0
| 1 |  | 17 |  | 17 |  | 17 |  |  | 17 |  |  | 17 |  |  | 16 |  | 0
| 1.5 |  | 18 |  | 18 |  | 18 |  |  | 18 |  |  | 18 |  |  | 16.5 |  | 0
| 2 |  | 19 |  | 19 |  | 19 |  |  | 19 |  |  | 19 |  |  | 17 |  | 0
| 2.5 |  | 20 |  | 20 |  | 20 |  |  | 20 |  |  | 20 |  |  | 17.5 |  | 0
| 3 |  | 21 |  | 21 |  | 21 |  |  | 21 |  |  | 21 |  |  | 18 |  | 0 |  |  |  | 0.1
| 3.5 |  | 22 |  | 22 |  | 22 |  |  | 22 |  |  | 22 |  |  | 18.5 |  | 0 |  |  |  | 50
| 4 |  | 23 |  | 23 |  | 23 |  |  | 23 |  |  | 23 |  |  | 19 |  | 0 |  |  |  | 5
| 4.5 |  | 24 |  | 24 |  | 24 |  |  | 24 |  |  | 24 |  |  | 19.5 |  | 0
| 5 |  | 25 |  | 25 |  | 25 |  |  | 25 |  |  | 25 |  |  | 20 |  | 0
| 6 |  | 12 |  | 12 |  | 12 |  |  | 12 |  |  | 12 |  |  | 6 |  | 0
| 7 |  | 13 |  | 13 |  | 13 |  |  | 13 |  |  | 13 |  |  | 6 |  | 0
| 8 |  | 14 |  | 14 |  | 14 |  |  | 14 |  |  | 14 |  |  | 6 |  | 0
| 9 |  | 15 |  | 15 |  | 15 |  |  | 15 |  |  | 15 |  |  | 6 |  | 0
| 10 |  | 16 |  | 16 |  | 16 |  |  | 16 |  |  | 16 |  |  | 6 |  | 0
| 15 |  | 17 |  | 17 |  | 17 |  |  | 17 |  |  | 17 |  |  | 2 |  | 0
| 20 |  | 18 |  | 18 |  | 18 |  |  | 18 |  |  | 18 |  |  | -2 |  | 0
| 25 |  | 19 |  | 19 |  | 19 |  |  | 19 |  |  | 19 |  |  | -6 |  | 0
| 30 |  | 20 |  | 20 |  | 20 |  |  | 20 |  |  | 20 |  |  | -10 |  | 0
| 35 |  | 21 |  | 21 |  | 21 |  |  | 21 |  |  | 21 |  |  | -14 |  | 0
| 40 |  | 22 |  | 22 |  | 22 |  |  | 22 |  |  | 22 |  |  | -18 |  | 0
| 45 |  | 23 |  | 23 |  | 23 |  |  | 23 |  |  | 23 |  |  | -22 |  | 0
| 50 |  | 24 |  | 24 |  | 24 |  |  | 24 |  |  | 24 |  |  | -26 |  | 0

## Data Record(Backward)
SP METROLOGY SYSTEM THAILAND | [5 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [6 empty] | Page |  | 2 | of | 2
[6 empty] | Receive Date : |  |  |  | 2016-01-12 |  |  |  |  | Calibration Date : |  |  | 2016-01-13
Equipment Name | [5 empty] | Temp & Humiduty : |  |  |  |  | 20 |  | oC | 50 |  | %RH
Dial Gauge Tester/ Calibration Tester | [5 empty] | Location |  |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  | Dial Gauge Tester/ Calibration Tester | [8 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  | 123 |  |  |  | Serial No. : |  |  | 456 |  |  |  |  | ID No : |  | 789
Range : |  | 0 | to | 10 | mm | Resolution : |  |  | 0.001 |  | mm
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : | [12 empty] | Due Date :
1. Flatness Measurement : |  |  |  |  | 0 | µm
2. Function Measurement
| Norminal Value |  | STD Reading (Backward) | [9 empty] | Average |  |  | Error |  | Repeatability
|  |  | X1 |  | X2 |  | X3 |  |  | X4
| 0 |  | 0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0
| 0.01 |  | 1 |  | 1 |  | 1 |  |  | 1 |  |  | 1 |  |  | 0.99 |  | 0
| 0.02 |  | 2 |  | 2 |  | 2 |  |  | 2 |  |  | 2 |  |  | 1.98 |  | 0
| 0.03 |  | 3 |  | 3 |  | 3 |  |  | 3 |  |  | 3 |  |  | 2.97 |  | 0
| 0.04 |  | 4 |  | 4 |  | 4 |  |  | 4 |  |  | 4 |  |  | 3.96 |  | 0
| 0.05 |  | 5 |  | 5 |  | 5 |  |  | 5 |  |  | 5 |  |  | 4.95 |  | 0
| 0.06 |  | 6 |  | 6 |  | 6 |  |  | 6 |  |  | 6 |  |  | 5.94 |  | 0
| 0.07 |  | 7 |  | 7 |  | 7 |  |  | 7 |  |  | 7 |  |  | 6.93 |  | 0
| 0.08 |  | 8 |  | 8 |  | 8 |  |  | 8 |  |  | 8 |  |  | 7.92 |  | 0
| 0.09 |  | 9 |  | 9 |  | 9 |  |  | 9 |  |  | 9 |  |  | 8.91 |  | 0
| 0.1 |  | 10 |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 9.9 |  | 0
| 0.2 |  | 11 |  | 11 |  | 11 |  |  | 11 |  |  | 11 |  |  | 10.8 |  | 0
| 0.25 |  | 12 |  | 12 |  | 12 |  |  | 12 |  |  | 12 |  |  | 11.75 |  | 0
| 0.3 |  | 13 |  | 13 |  | 13 |  |  | 13 |  |  | 13 |  |  | 12.7 |  | 0
| 0.35 |  | 14 |  | 14 |  | 14 |  |  | 14 |  |  | 14 |  |  | 13.65 |  | 0
| 0.4 |  | 15 |  | 15 |  | 15 |  |  | 15 |  |  | 15 |  |  | 14.6 |  | 0
| 0.45 |  | 16 |  | 16 |  | 16 |  |  | 16 |  |  | 16 |  |  | 15.55 |  | 0
| 0.5 |  | 17 |  | 17 |  | 17 |  |  | 17 |  |  | 17 |  |  | 16.5 |  | 0
| 1 |  | 18 |  | 18 |  | 18 |  |  | 18 |  |  | 18 |  |  | 17 |  | 0
| 2 |  | 19 |  | 19 |  | 19 |  |  | 19 |  |  | 19 |  |  | 17 |  | 0
| 3 |  | 20 |  | 20 |  | 20 |  |  | 20 |  |  | 20 |  |  | 17 |  | 0
| 4 |  | 21 |  | 21 |  | 21 |  |  | 21 |  |  | 21 |  |  | 17 |  | 0
| 5 |  | 22 |  | 22 |  | 22 |  |  | 22 |  |  | 22 |  |  | 17 |  | 0
| 10 |  | 23 |  | 23 |  | 23 |  |  | 23 |  |  | 23 |  |  | 13 |  | 0
| 15 |  | 24 |  | 24 |  | 24 |  |  | 24 |  |  | 24 |  |  | 9 |  | 0
| 20 |  | 25 |  | 25 |  | 25 |  |  | 25 |  |  | 25 |  |  | 5 |  | 0
| 25 |  | 26 |  | 26 |  | 26 |  |  | 26 |  |  | 26 |  |  | 1 |  | 0
Calibration by :

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR15120012-1 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
88/115, Moo 3, T. Klongsam
A. Klongluang, Pathumthani 12120
| Equipment Name |  |  | : | Dial Gauge Tester/ Calibration Tester
| Manufacturer |  |  | : | 0
| Model |  |  | : | 123
| Serial Number |  |  | : | 0
| ID. Number |  |  | : | 0
| Received Date |  |  | : | 2016-01-12
| Calibration Date |  |  | : | 2016-01-13
| Recommended Due Date |  |  | : | 2017-01-13
Environmental Conditions
| Ambient Temperature |  |  | : | 20 °C | ± 1 °C
| Relative Humidity |  |  | : | 50% | ± 15 %
| Location of Calibration |  |  | : | In-Lab
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. standards used to perform this
calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. the result reported herein apply only to the calibration of the item described above.
| all calibrations are performed manufacture's specifications, full, without the expressed written consent of | [9 empty] | 1 | Mr.Sombut Srikampa
| SP Metrology System (Thailand). | [9 empty] | 3 | Mr. Natthaphol Boonmee
[11 empty] | 8 | Mr. Vichan Ananta
[11 empty] | 9 | Mr.Kittikorn Kingmali
| Date of Issue : |  | 2016-01-14 |  |  |  | Approved by : |  |  |  | 10 | Mr.Chainarong Matchayamat
[8 empty] | 3 | ( Mr. Natthaphol Boonmee ) |  | 11 | Ms. Arunkamon Raramanus
| Calibrated by : |  | Ms. Arunkamon Raramanus | [5 empty] | Authorized Signatory
Note on O39: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120011 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Linear gauge |  |  | MT2501/ ND281B | 25347859 P/ 16126167A | MTL152769 | 2016-10-21
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  | SPR15120011 | [8 empty] | Page : 3 of 3
|  | Scale Range |  | : | 0 | to | 10 | mm. | Resolution : | 0.001 |  | mm.
[14 empty] | Unit : | mm
|  | Nominal Value | STD Reading Forward |  |  |  | Error |  | STD Reading Backward |  | Error |  |  | Uncertainty (±) µm
|  | 0 | 0 |  |  |  | 0 |  | 0 |  | 0 |  |  | 0.65067
|  | 0.01 | 1 |  |  |  | 0.99 |  | 1 |  | 0.99 |  |  | 0.65067
|  | 0.02 | 2 |  |  |  | 1.98 |  | 2 |  | 1.98 |  |  | 0.65067
|  | 0.03 | 3 |  |  |  | 2.97 |  | 3 |  | 2.97 |  |  | 0.65067
|  | 0.04 | 4 |  |  |  | 3.96 |  | 4 |  | 3.96 |  |  | 0.65067
|  | 0.05 | 5 |  |  |  | 4.95 |  | 5 |  | 4.95 |  |  | 0.65067
|  | 0.06 | 6 |  |  |  | 5.94 |  | 6 |  | 5.94 |  |  | 0.65067
|  | 0.07 | 7 |  |  |  | 6.93 |  | 7 |  | 6.93 |  |  | 0.65067
|  | 0.08 | 8 |  |  |  | 7.92 |  | 8 |  | 7.92 |  |  | 0.65067
|  | 0.09 | 9 |  |  |  | 8.91 |  | 9 |  | 8.91 |  |  | 0.65067
|  | 0.1 | 10 |  |  |  | 9.9 |  | 10 |  | 9.9 |  |  | 0.65067
|  | 0.4 | 11 |  |  |  | 10.6 |  | 11 |  | 10.8 |  |  | 0.65069
|  | 0.5 | 12 |  |  |  | 11.5 |  | 12 |  | 11.75 |  |  | 0.6507
|  | 0.6 | 13 |  |  |  | 12.4 |  | 13 |  | 12.7 |  |  | 0.65072
|  | 0.7 | 14 |  |  |  | 13.3 |  | 14 |  | 13.65 |  |  | 0.65073
|  | 0.8 | 15 |  |  |  | 14.2 |  | 15 |  | 14.6 |  |  | 0.65075
|  | 0.9 | 16 |  |  |  | 15.1 |  | 16 |  | 15.55 |  |  | 0.65078
|  | 1 | 17 |  |  |  | 16 |  | 17 |  | 16.5 |  |  | 0.6508
|  | 1.5 | 18 |  |  |  | 16.5 |  | 18 |  | 17 |  |  | 0.65097
|  | 2 | 19 |  |  |  | 17 |  | 19 |  | 17 |  |  | 0.65121
|  | 2.5 | 20 |  |  |  | 17.5 |  | 20 |  | 17 |  |  | 0.65151
|  | 3 | 21 |  |  |  | 18 |  | 21 |  | 17 |  |  | 0.65188
|  | 3.5 | 22 |  |  |  | 18.5 |  | 22 |  | 17 |  |  | 0.65232
|  | 4 | 23 |  |  |  | 19 |  | 23 |  | 13 |  |  | 0.65283
|  | 4.5 | 24 |  |  |  | 19.5 |  | 24 |  | 9 |  |  | 0.6534
|  | 5 | 25 |  |  |  | 20 |  | 25 |  | 5 |  |  | 0.65405
|  | 6 | 12 |  |  |  | 6 |  | 26 |  | 1 |  |  | 0.65553
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget Dial Gauge Tester/ Calibration Tester
Nominal Value | Repeatability |  | Uncertainty of STD |  | Deformation |  | Temperature Effect |  | Resolution of STD |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 0 | 0 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.01 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.15e-07 | 6.6395e-08 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.02 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 2.3e-07 | 1.3279e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.03 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 3.45e-07 | 1.9919e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.04 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 4.6e-07 | 2.6558e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.05 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 5.75e-07 | 3.3198e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.06 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 6.9e-07 | 3.9837e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.07 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 8.05e-07 | 4.6477e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.08 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 9.2e-07 | 5.3116e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.09 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.035e-06 | 5.9756e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.1 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.15e-06 | 6.6395e-07 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032533 | 0.00032533 | ∞ | 2 | 0.65067
0.4 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 4.6e-06 | 2.6558e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032534 | 0.00032534 | ∞ | 2 | 0.65069
0.5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 5.75e-06 | 3.3198e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032535 | 0.00032535 | ∞ | 2 | 0.6507
0.6 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 6.9e-06 | 3.9837e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032536 | 0.00032536 | ∞ | 2 | 0.65072
0.7 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 8.05e-06 | 4.6477e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032537 | 0.00032537 | ∞ | 2 | 0.65073
0.8 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 9.2e-06 | 5.3116e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032538 | 0.00032538 | ∞ | 2 | 0.65075
0.9 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.035e-05 | 5.9756e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032539 | 0.00032539 | ∞ | 2 | 0.65078
1 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.15e-05 | 6.6395e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.0003254 | 0.0003254 | ∞ | 2 | 0.6508
1.5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 1.725e-05 | 9.9593e-06 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032549 | 0.00032549 | ∞ | 2 | 0.65097
2 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 2.3e-05 | 1.3279e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.0003256 | 0.0003256 | ∞ | 2 | 0.65121
2.5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 2.875e-05 | 1.6599e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032576 | 0.00032576 | ∞ | 2 | 0.65151
3 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 3.45e-05 | 1.9919e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032594 | 0.00032594 | ∞ | 2 | 0.65188
3.5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 4.025e-05 | 2.3238e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032616 | 0.00032616 | ∞ | 2 | 0.65232
4 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 4.6e-05 | 2.6558e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032642 | 0.00032642 | ∞ | 2 | 0.65283
4.5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 5.175e-05 | 2.9878e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.0003267 | 0.0003267 | ∞ | 2 | 0.6534
5 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 5.75e-05 | 3.3198e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032702 | 0.00032702 | ∞ | 2 | 0.65405
6 | 0 | 0 | 0.0003 | 0.00015 | 8.3532e-08 | 4.8227e-08 | 6.9e-05 | 3.9837e-05 | 5e-06 | 2.8868e-06 | 0.0005 | 0.00028868 | 0.00032776 | 0.00032776 | ∞ | 2 | 0.65553
[5 empty] | A2 | 8.3532e-08
|  |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  |  | 0.137
|  |  |  | dD = | Ball diameter |  | 0.003
|  |  |  | V = | Poisson-coefficient |  | Steel = | 0.3
|  |  |  | E = | Elasticity modulus |  | Steel = | 2*10^11
Note on F4: Certificate of Calubration Linear Gauge
Note on J4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3

## Uncert of STD
Certificate of Calibration (Length Gauge)
SP-SD-028
Due Date |  | 2016-10-21
0 | 0.3 | µm | 0.0003 | mm
2.5 | 0.3 | µm | 0.0003 | mm
5 | 0.3 | µm | 0.0003 | mm
7.5 | 0.3 | µm | 0.0003 | mm
10 | 0.3 | µm | 0.0003 | mm
12.5 | 0.3 | µm | 0.0003 | mm
15 | 0.3 | µm | 0.0003 | mm
17.5 | 0.3 | µm | 0.0003 | mm
20 | 0.3 | µm | 0.0003 | mm
22.5 | 0.3 | µm | 0.0003 | mm
25 | 0.3 | µm | 0.0003 | mm
Note on D5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on C6: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Dial Test Indicator
Note on E6: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน