## Data
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR17xxxx | [5 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 42459 |  |  |  | Calibration Date : |  |  |  | 42459
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 | %RH
Balance | [7 empty] | Location |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  |  | ABC Ltd.
Adress |  | :
Equipment Name : |  |  |  | Electronic Balance | [6 empty] | Manufacturer : |  |  |  |  | AA
Model : | AB | [6 empty] | Serial No. : |  |  | 1234567 | [5 empty] | ID No : |  | N/A
Range : | 0 |  | to | 2500 |  | g |  |  | Resolution : | 0.01 |  |  |  | g
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Reference Standard : |  |  |  |  | SP-SM-026 | [7 empty] | Due Date : |  |  | 2017-11-22
Reference Standard : | [12 empty] | Due Date :
Calibration Procedure No.: |  |  |  |  | SP-CPM-04-01
Measurement Result
Repeatability ( n = 10 number of measurement ) | [8 empty] | (Max Capacity)
Value | 1 |  |  |  | 2 |  |  | 3 |  |  | 4 |  |  |  | 5 |  |  | STDEV
2500 | 2500 |  |  |  | 2500 |  |  | 2500 |  |  | 2500 |  |  |  | 2500
| 3 |  |  |  | 7 |  |  | 8 |  |  | 9 |  |  |  | 10 |  |  | 0
| 2500 |  |  |  | 2500 |  |  | 2500 |  |  | 2500 |  |  |  | 2500
Departure of indication from nominal Value | [8 empty] | (Step up to 10%)
Nominal | UUC Reading | [13 empty] | Average |  |  | Repeatability |  |  | Error
g | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
0 | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0
250 | 250 |  |  |  | 250 |  |  | 250 |  |  | 250 |  |  |  | 250 |  |  | 0 |  |  | 0
500 | 500 |  |  |  | 500 |  |  | 500 |  |  | 500 |  |  |  | 500 |  |  | 0 |  |  | 0
750 | 750 |  |  |  | 750 |  |  | 750 |  |  | 750 |  |  |  | 750 |  |  | 0 |  |  | 0
1000 | 1000 |  |  |  | 1000 |  |  | 1000 |  |  | 1000 |  |  |  | 1000 |  |  | 0 |  |  | 0
1250 | 1250 |  |  |  | 1250 |  |  | 1250 |  |  | 1250 |  |  |  | 1250 |  |  | 0 |  |  | 0
1500 | 1500 |  |  |  | 1500 |  |  | 1500 |  |  | 1500 |  |  |  | 1500 |  |  | 0 |  |  | 0
1750 | 1750 |  |  |  | 1750 |  |  | 1750 |  |  | 1750 |  |  |  | 1750 |  |  | 0 |  |  | 0
2000 | 2000 |  |  |  | 2000 |  |  | 2000 |  |  | 2000 |  |  |  | 2000 |  |  | 0 |  |  | 0
2250 | 2250 |  |  |  | 2250 |  |  | 2250 |  |  | 2250 |  |  |  | 2250 |  |  | 0 |  |  | 0
2500 | 2500 |  |  |  | 2500 |  |  | 2500 |  |  | 2500 |  |  |  | 2500 |  |  | 0 |  |  | 0
Off - Center Loading | [6 empty] | (1/4 of Capacity) | [13 empty] | Page 2 of 2
Load test | lndicator
625 | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
Center | 625 |  |  |  | 625 |  |  | 625 |  |  | 625
Front | 625 |  |  |  | 625 |  |  | 625 |  |  | 625
Back | 625 |  |  |  | 625 |  |  | 625 |  |  | 625
Left | 625 |  |  |  | 625 |  |  | 625 |  |  | 625
Right | 625 |  |  |  | 625 |  |  | 625 |  |  | 625
Diff. Error Max. |  | 0
Calibrated By : |  |  | Mr.Nirut Loha
Mr.Nirut Loha
Mr.Pakapon Nammontree
Mr.Prayoon Topart
Mr.Santi Thonghlor
Mr.Werayut Jampol

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR17xxxx | [6 empty] | Page : 1 of 3
| Customer |  |  |  | : | ABC Ltd.
0
| Equipment Name |  |  |  | : | Electronic Balance
| Manufacturer |  |  |  | : | AA
| Model |  |  |  | : | AB
| Serial Number |  |  |  | : | 1234567
| ID. Number |  |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  |  |  | : | 42459
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 42459
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 42824
| Calibration Procedure |  |  |  |  | : | SP-CPM-04-01 |  | Date of Issue |  |  |  | : | 42460
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Calibrated by |  | : | Mr.Nirut Loha |  |  |  | Approved by :
|  |  |  | Calibration Officer |  |  |  | 3 |  |  |  | (Mr. Santi Hankitudomsuk)
Authorized Signatory
[10 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q38: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR17xxxx |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Standard Weight Set |  |  | Class F1 | SS6K3324-16 | CM160770 | 22 Nov 2017
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR17xxxx | [7 empty] | Page : 3 of 3
Repeatability ( n = 10 number of measurement )
|  | Standard Weight |  | Standard Deviation
|  | 2500 | g | 0 |  | g
| Departure of indication from nominal Value | [7 empty] | Unit : | g
|  | Nominal Value |  |  | UUC Reading |  | Error | Uncertainty ( ± ) mg |  |  |  | mg
[12 empty] | UUC | Range | CMC
|  | 0 |  |  | 0 |  | 0 | 9.56 |  |  |  | 9.56 | 0 - 1000 g | 0.04
|  | 250 |  |  | 250 |  | 0 | 9.57 |  |  |  | 9.57 |  | 0.4 | 3.31
|  | 500 |  |  | 500 |  | 0 | 9.61 |  |  |  | 9.61 |  | 1 | 3.31
|  | 750 |  |  | 750 |  | 0 | 9.67 |  |  |  | 9.67 |  | 1 | 3.31
|  | 1000 |  |  | 1000 |  | 0 | 9.75 |  |  |  | 9.75 |  | 1 | 3.31
|  | 1250 |  |  | 1250 |  | 0 | 9.85 |  |  |  | 9.85 |  | 2 | 3.31
|  | 1500 |  |  | 1500 |  | 0 | 9.57 |  |  |  | 9.57 |  | 2 | 3.31
|  | 1750 |  |  | 1750 |  | 0 | 9.61 |  |  |  | 9.61 |  | 2 | 3.31
|  | 2000 |  |  | 2000 |  | 0 | 9.67 |  |  |  | 9.67 |  | 2 | 3.31
|  | 2250 |  |  | 2250 |  | 0 | 9.75 |  |  |  | 9.75 |  | 3 | 3.31
|  | 2500 |  |  | 2500 |  | 0 | 9.85 |  |  |  | 9.85 |  | 3 | 3.31
Off - Center Loading
|  | Center |  | 625 |  | g
|  | Front |  | 625 |  | g
|  | Back |  | 625 |  | g
|  | Left |  | 625 |  | g
|  | Right |  | 625 |  | g
|  | Maximum difference |  | 0 |  | g
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%
- End of Certificate -

## Uncert Budget
Uncertainty Budget of Electronic Balance
Nominal | Uncertainty |  |  | Air Buoyancy effect |  | Resolution of |  | Repeatability |  | u(δIecc) |  | u(δItemp) |  | Uc | Ui | Veff | K95 | U95%
| Standard Weight. |  |  | ± 1 ppm 10^6 |  | UUC |  | of UUC
g | g |  |  | g |  | g |  | g
Value | Value |  | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
250 |  |  | 0 | 4.3832e-06 | 2.5306e-06 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.001 | 0.00028868 | 0.0029012 | 0 | ∞ | 2 | 0.0058023
500 |  |  | 0 | 8.7663e-06 | 5.0612e-06 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.002 | 0.00057735 | 0.0029439 | 0 | ∞ | 2 | 0.0058878
750 |  |  | 0 | 1.3149e-05 | 7.5919e-06 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.003 | 0.00086603 | 0.0030139 | 0 | ∞ | 2 | 0.0060277
1000 |  |  | 0 | 1.7533e-05 | 1.0122e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.004 | 0.0011547 | 0.0031091 | 0 | ∞ | 2 | 0.0062183
1250 |  |  | 0 | 2.1916e-05 | 1.2653e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.005 | 0.0014434 | 0.0032275 | 0 | ∞ | 2 | 0.006455
1500 |  |  | 0 | 2.6299e-05 | 1.5184e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.006 | 0.0017321 | 0.0033665 | 0 | ∞ | 2 | 0.0067331
1750 |  |  | 0 | 3.0682e-05 | 1.7714e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.007 | 0.0020207 | 0.0035238 | 0 | ∞ | 2 | 0.0070475
2000 |  |  | 0 | 3.5065e-05 | 2.0245e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.008 | 0.0023094 | 0.0036969 | 0 | ∞ | 2 | 0.0073938
2250 |  |  | 0 | 3.9448e-05 | 2.2776e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.009 | 0.0025981 | 0.0038838 | 0 | ∞ | 2 | 0.0077676
2500 |  |  | 0 | 4.3832e-05 | 2.5306e-05 | 0.005 | 0.0028868 | 0 | 0 | 0 | 0 | 0.01 | 0.0028868 | 0.0040826 | 0 | ∞ | 2 | 0.0081651
Temperature | 23
Humidity | 50
Range | 0-2500 g.
ρa | 1.1777
Resolution | 0.01 | g.
ปี 2012
Nominal |  |  | Variation
Length |  |  | of GB.
( mm ) |  |  | ( mm )
1 |  |  | 0.03
1.005 |  |  | 0.05
1.01 |  |  | 0.04
1.02 |  |  | 0.07
1.03 |  |  | 0.04
1.04 |  |  | 0.04
1.05 |  |  | 0.06
1.06 |  |  | 0.07
1.07 |  |  | 0.08
1.08 |  |  | 0.04
1.09 |  |  | 0.04
1.1 |  |  | 0.04
1.2 |  |  | 0.14
1.3 |  |  | 0.03
1.4 |  |  | 0.07
1.5 |  |  | 0.02
1.6 |  |  | 0.07
1.7 |  |  | 0.07
1.8 |  |  | 0.06
1.9 |  |  | 0.04
2 |  |  | 0.03
3 |  |  | 0.08
4 |  |  | 0.06
5 |  |  | 0.07
6 |  |  | 0.05
7 |  |  | 0.07
8 |  |  | 0.01
9 |  |  | 0.07
10 |  |  | 0.06
11 |  |  | 0.06
12 |  |  | 0.02
13 |  |  | 0.04
14 |  |  | 0.05
15 |  |  | 0.05
16 |  |  | 0.07
17 |  |  | 0.04
18 |  |  | 0.05
19 |  |  | 0.09
20 |  |  | 0.08
21 |  |  | 0.06
22 |  |  | 0.04
23 |  |  | 0.04
24 |  |  | 0.13
25 |  |  | 0.05
50 |  |  | 0.14
75 |  |  | 0.14
100 |  |  | 0.08
125 |  |  | 0.03
150 |  |  | 0
175 |  |  | 0
200 |  |  | 0
250 |  |  | 0
300 |  |  | 0
400 |  |  | 0
500 |  |  | 0
600 |  |  | 0
700 |  |  | 0
800 |  |  | 0
900 |  |  | 0
1000 |  |  | 0
Note on N3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การวางน้ำหนักไม่ตรงกลาง
Note on P3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การเปลี่ยนแปลงอุณภูมิ
Note on B6: (1ppm*W)/SQRT3
Note on C6: ค่าจากใบรายงานผลการสอบเทียบ STD Weight
Note on H6: (1ppm*W)/SQRT3
Note on J6: ค่าความละเอียดในการอ่านของเครื่องชั่ง
Note on L6: ได้จากการวัดซ้ำ
Note on D7: ใช้ตุ้มน้ำหนัก2ค่ารวมกัน
Note on E7: ใช้ตุ้มน้ำหนัก3ค่ารวมกัน
Note on F7: ใช้ตุ้มน้ำหนัก4ค่ารวมกัน
Note on B21: SP Metrology: The current density of an Air.

## Uncert Budget '2017
Uncertainty Budget of Electronic Balance
Function : | Mass |  |  | Testpoint | 0 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0028868 | 3
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00956 | 3
[5 empty] | 3.31 |  |  |  |  | 3
Function : | Mass |  |  | Testpoint | 250 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00025 |  | Rectangular |  | 1.7321 | 1 |  | 0.00014434 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0028904 | 3.015
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00957 | 3.015
[5 empty] | 3.31 |  |  |  |  | 3.015
Function : | Mass |  |  | Testpoint | 500 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0029011 | 3.0603
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00961 | 3.0603
[5 empty] | 3.31 |  |  |  |  | 3.0603
Function : | Mass |  |  | Testpoint | 750 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00075 |  | Rectangular |  | 1.7321 | 1 |  | 0.00043301 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.002919 | 3.1365
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00967 | 3.1365
[5 empty] | 3.31 |  |  |  |  | 3.1365
Function : | Mass |  |  | Testpoint | 1000 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.001 |  | Rectangular |  | 1.7321 | 1 |  | 0.00057735 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0029439 | 3.2448
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00975 | 3.2448
[5 empty] | 3.31 |  |  |  |  | 3.2448
Function : | Mass |  |  | Testpoint | 1250 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00125 |  | Rectangular |  | 1.7321 | 1 |  | 0.00072169 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0029756 | 3.3867
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00985 | 3.3867
[5 empty] | 3.31 |  |  |  |  | 3.3867
Function : | Mass |  |  | Testpoint | 1500 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.0015 |  | Rectangular |  | 1.7321 | 1 |  | 0.00086603 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0030139 | 3.5643
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00998 | 3.5643
[5 empty] | 3.31 |  |  |  |  | 3.5643
Function : | Mass |  |  | Testpoint | 1750 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00175 |  | Rectangular |  | 1.7321 | 1 |  | 0.0010104 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0030585 | 3.78
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.01013 | 3.78
[5 empty] | 3.31 |  |  |  |  | 3.78
Function : | Mass |  |  | Testpoint | 2000 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.002 |  | Rectangular |  | 1.7321 | 1 |  | 0.0011547 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0031091 | 4.0368
U | Expaned uncertainty |  |  |  | k = | 2.87 |  |  |  | 0.00893 | 4.0368
[5 empty] | 2.87 |  |  |  |  | 4.0368
Function : | Mass |  |  | Testpoint | 2250 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00225 |  | Rectangular |  | 1.7321 | 1 |  | 0.001299 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0031656 | 4.338
U | Expaned uncertainty |  |  |  | k = | 2.87 |  |  |  | 0.00909 | 4.338
[5 empty] | 2.87 |  |  |  |  | 4.338
Function : | Mass |  |  | Testpoint | 2500 |  | g | Range | 2500 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.0025 |  | Rectangular |  | 1.7321 | 1 |  | 0.0014434 | ¥
UB2 | Resolution of UUC |  | 0.005 |  | Rectangular |  | 1.7321 | 1 |  | 0.0028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0032275 | 4.6875
U | Expaned uncertainty |  |  |  | k = | 2.87 |  |  |  | 0.00927 | 4.6875
[5 empty] | 2.87 |  |  |  |  | 4.6875

## Cert of STD
Weight 10 mg - 500 mg |  |  |  | Weight 1g - 5 kg |  |  |  | Weight 10 g - 20 kg |  |  |  | Weight 5 kg - 20 kg
SP-SM-001 |  |  |  | SP-SM-002 |  |  |  | SP-SM-003 |  |  |  | SP-SM-004
Due Date |  |  |  | Due Date |  |  |  | Due Date |  |  |  | No. | Due Date
42550 |  |  |  | 42550 |  |  |  | 42601 |  |  |  |  | 42706
10 | mg | 0.01 | mg | 1 | g | 0.02 | mg | 10 | g | 0.035 | mg | 1 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg | 2 | g* | 0.03 | mg* | 20 | g | 0.042 | mg | 2 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg* | 2 | g | 0.03 | mg | 50 | g | 0.049 | mg | 1 | 10 | kg | 31 | mg
50 | mg | 0.01 | mg | 5 | g | 0.04 | mg | 100 | g | 0.11 | mg | 2 | 10 | kg | 31 | mg
100 | mg | 0.01 | mg | 10 | g | 0.04 | mg | 200 | g | 0.17 | mg | 1 | 20 | kg | 54 | mg
200 | mg | 0.01 | mg | 20 | g | 0.05 | mg | 500 | g | 0.4 | mg | 2 | 20 | kg | 54 | mg
500 | mg | 0.01 | mg | 20 | g* | 0.05 | mg* | 1 | kg | 1.7 | mg | 3 | 20 | kg | 54 | mg
|  |  |  | 50 | g | 0.1 | mg | 2 | kg | 2 | mg | 4 | 20 | kg | 54 | mg
|  |  |  | 100 | g | 0.2 | mg | 10 | kg | 12 | mg | 5 | 20 | kg | 54 | mg
|  |  |  | 200 | g | 0.4 | mg | 20 | kg | 18 | mg | 6 | 20 | kg | 54 | mg
|  |  |  | 200 | g* | 0.4 | mg* |  |  |  |  | 7 | 20 | kg | 54 | mg
|  |  |  | 500 | g | 2 | mg |  |  |  |  | 8 | 20 | kg | 54 | mg
|  |  |  | 1 | kg | 2 | mg |  |  |  |  | 9 | 20 | kg | 54 | mg
|  |  |  | 1 | kg* | 2 | mg |  |  |  |  | 10 | 20 | kg | 54 | mg
|  |  |  | 2 | kg | 4 | mg |  |  |  |  | 11 | 20 | kg | 54 | mg
|  |  |  | 5 | kg | 10 | mg |  |  |  |  | 12 | 20 | kg | 54 | mg
[12 empty] | 13 | 20 | kg | 54 | mg
[12 empty] | 14 | 20 | kg | 54 | mg
[12 empty] | 15 | 20 | kg | 54 | mg
[12 empty] | 16 | 20 | kg | 54 | mg
[12 empty] | 17 | 20 | kg | 54 | mg
[12 empty] | 18 | 20 | kg | 54 | mg
[12 empty] | 19 | 20 | kg | 54 | mg
[12 empty] | 20 | 20 | kg | 54 | mg
[12 empty] | 21 | 20 | kg | 54 | mg
[12 empty] | 22 | 20 | kg | 54 | mg
[12 empty] | 23 | 20 | kg | 54 | mg
[12 empty] | 24 | 20 | kg | 54 | mg
[12 empty] | 25 | 20 | kg | 54 | mg
[12 empty] | 26 | 20 | kg | 54 | mg
[12 empty] | 27 | 20 | kg | 54 | mg
[12 empty] | 28 | 20 | kg | 54 | mg
[12 empty] | 29 | 20 | kg | 54 | mg
[12 empty] | 30 | 20 | kg | 54 | mg
[12 empty] | 31 | 20 | kg | 54 | mg
[12 empty] | 32 | 20 | kg | 54 | mg
[12 empty] | 33 | 20 | kg | 54 | mg
[12 empty] | 34 | 20 | kg | 54 | mg
[12 empty] | 35 | 20 | kg | 54 | mg
[12 empty] | 36 | 20 | kg | 54 | mg
[12 empty] | 37 | 20 | kg | 54 | mg
[12 empty] | 38 | 20 | kg | 54 | mg
[12 empty] | 39 | 20 | kg | 54 | mg
[12 empty] | 40 | 20 | kg | 54 | mg
[12 empty] | 41 | 20 | kg | 54 | mg
[12 empty] | 42 | 20 | kg | 54 | mg
[12 empty] | 43 | 20 | kg | 54 | mg
[12 empty] | 44 | 20 | kg | 54 | mg
[12 empty] | 45 | 20 | kg | 54 | mg
[12 empty] | 46 | 20 | kg | 54 | mg
[12 empty] | 47 | 20 | kg | 54 | mg
[12 empty] | 48 | 20 | kg | 54 | mg
[12 empty] | 49 | 20 | kg | 54 | mg
[12 empty] | 50 | 20 | kg | 54 | mg
Note on B6: MPE Class F1
Note on G6: MPE Class F1
Note on L6: MPE Class E2
Note on Q6: MPE Class M1
Note on B8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B9: Nathaphol Boonmee: Normal Value
Note on D9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G9: Nathaphol Boonmee: Normal Value
Note on I9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L9: Nathaphol Boonmee: Normal Value
Note on N9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R9: Nathaphol Boonmee: Normal Value
Note on T9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง