## Data
SP METROLOGY SYSTEM THAILAND | [6 empty] | Certificate No. : |  |  |  | SPR17020004-1 | [9 empty] | Page 1 of 1
[7 empty] | Receive Date : |  |  |  | 2017-02-01 |  |  |  |  | Calibration Date : |  |  |  | 2017-02-02
Equipment Name | [6 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 |  | %RH
Duro Meter | [6 empty] | Location |  |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  | Luckytex (Thailand)
Adress |  | :
Equipment Name : |  |  | Duro Meter | [7 empty] | Manufacturer : |  |  |  | HOTO
Model : | EX-C | [5 empty] | Serial No. : |  |  | 050228 | [7 empty] | ID No : | AB-FN-0105
Range : | Shore C | [7 empty] | Resolution : | 1
Overall Inspection |  |  |  |  | Good |  | Not Good
Reference Standard |  |  |  | : | SP-SM-001 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-002 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-002-1 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-003 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-003-1 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-003-2 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-004 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-026 | [7 empty] | Due Date : |  |  | 2017-11-22
Calibration Procedure |  |  |  | : | SP-CPM-04-03 | [17 empty] | 10 | 1.3 |  | 132.56 |  | 124.91 |  | 140.21 | g
[22 empty] | Type | Point | Load Nominal |  | Load Nominal |  | Limits
Measurement Result | [21 empty] | A, B and O | 20 | 2.05 | N | 209.04 | g | 201.39 | to | 216.69 | g
Normal | UUC Reading | [12 empty] | Error |  |  | STDV |  |  | Repeatability |  |  | 30 | 2.8 | N | 285.52 | g | 277.87 | to | 293.17 | g
Value | X1 |  |  | X2 |  | X3 |  | X4 |  |  | Average | [11 empty] | 40 | 3.55 | N | 362 | g | 354.35 | to | 369.65 | g
10 | 10 |  |  | 10 |  | 10 |  | 10 |  |  | 10 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 50 | 4.3 | N | 438.48 | g | 430.83 | to | 446.13 | g
20 | 20 |  |  | 20 |  | 20 |  | 20 |  |  | 20 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 60 | 5.05 | N | 514.96 | g | 507.31 | to | 522.61 | g
30 | 30 |  |  | 30 |  | 30 |  | 30 |  |  | 30 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 70 | 5.8 | N | 591.44 | g | 583.79 | to | 599.09 | g
40 | 40 |  |  | 40 |  | 40 |  | 40 |  |  | 40 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 80 | 6.55 | N | 667.91 | g | 660.26 | to | 675.56 | g
50 | 50 |  |  | 50 |  | 50 |  | 50 |  |  | 50 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 90 | 7.3 | N | 744.39 | g | 736.74 | to | 752.04 | g
60 | 60 |  |  | 60 |  | 60 |  | 60 |  |  | 60 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 100 | 8.05 | N | 820.87 | g | 813.22 | to | 828.52 | g
70 | 70 |  |  | 70 |  | 70 |  | 70 |  |  | 70 |  |  | 0 |  |  | 0 |  |  | 0
80 | 80 |  |  | 80 |  | 80 |  | 80 |  |  | 80 |  |  | 0 |  |  | 0 |  |  | 0 |  | Type | Point | Load Nominal |  | Load Nominal |  | Limits
90 | 90 |  |  | 90 |  | 90 |  | 90 |  |  | 90 |  |  | 0 |  |  | 0 |  |  | 0 |  | C, D and DO | 20 | 8.89 | N | 906.53 | g | 861.2 | to | 951.86 g
[20 empty] | 0 |  |  | 30 | 13.335 | N | 1.36 | kg | 1.315 | to | 1.405 kg
[23 empty] | 40 | 17.78 | N | 1.813 | kg | 1.768 | to | 1.858 kg
[23 empty] | 50 | 22.225 | N | 2.266 | kg | 2.221 | to | 2.311 kg
[23 empty] | 60 | 26.67 | N | 2.72 | kg | 2.675 | to | 2.765 kg
[23 empty] | 70 | 31.115 | N | 3.173 | kg | 3.128 | to | 3.218 kg
[23 empty] | 80 | 35.56 | N | 3.626 | kg | 3.581 | to | 3.671 kg
[23 empty] | 90 | 40.005 | N | 4.079 | kg | 4.034 | to | 4.124 kg
[23 empty] | 100 | 44.45 | N | 4.53 | kg | 4.485 | to | 4.575 kg
Calibrated By : |  |  | Mr. Pakapon Nammontree
| 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR17020004-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | : | Luckytex (Thailand)
0
| Equipment Name |  |  |  | : | Duro Meter
| Manufacturer |  |  |  | : | HOTO
| Model |  |  |  | : | EX-C
| Serial Number |  |  |  | : | 050228
| ID. Number |  |  |  | : | AB-FN-0105
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  |  |  | : | 2017-02-01
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2017-02-02
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2018-02-02
| Calibration Procedure |  |  |  |  | : | SP-CPM-04-04
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2017-02-03 |  |  |  | Approved by :
| Calibrated by |  | : | Mr. Pakapon Nammontree |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
[10 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR17020004-1 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Standard Weight |  |  | N/A | N/A | SPR16110161-3 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | SPR16110161-4 | 01 Dec 2017
| Standard Weight |  |  | 5 kg | N/A | 60-210057-1 | 10 Feb 2018
| Standard Weight Hook |  |  | N/A | N/A | SPR16110161-5 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | 60-210057-2 | 10 Feb 2018
| Standard Weight |  |  | N/A | N/A | 60-210057-3 | 10 Feb 2018
| Standard Weight |  |  | N/A | N/A | SPR16110161-2 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | SS6K3324-16 | 22 Nov 2017
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR17020004-1 |  |  |  | Page : 3 of 3
| Measurement Result |  | : | Shore C
Unit : Deg
|  | Nominal Value | UUC Reading |  | Error | Uncertainty ( ± )
|  | 10 | 10 |  | 0 | 0.57735
|  | 20 | 20 |  | 0 | 0.57735
|  | 30 | 30 |  | 0 | 0.57735
|  | 40 | 40 |  | 0 | 0.57736
|  | 50 | 50 |  | 0 | 0.57735
|  | 60 | 60 |  | 0 | 0.57736
|  | 70 | 70 |  | 0 | 0.57736
|  | 80 | 80 |  | 0 | 0.57737
|  | 90 | 90 |  | 0 | 0.57738
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%
- End of Certificate -

## Uncert Budget(A, B, O)
Uncertainty Budget of Durometer
Nominal | Repeatability |  | Uncertainty |  |  |  |  | Air Buoyancy effect |  | Resolution of |  | Uc | Ui | Veff | K95 | U95%
| of UUC |  | Standard Weight. |  |  |  |  | ± 1 ppm 10^6 |  | UUC
Point | Value | Ui | Value |  |  |  | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
20 | 0 | 0 | 0.17 | 0.035 |  |  | 0.1025 | 2e-05 | 1.1547e-05 | 0.5 | 0.28868 | 0.30633 | 0 | ∞ | 2 | 0.61266
30 | 0 | 0 | 0.17 | 0.049 | 0.042 | 0.035 | 0.148 | 3e-05 | 1.7321e-05 | 0.5 | 0.28868 | 0.3244 | 0 | ∞ | 2 | 0.64881
40 | 0 | 0 | 0.17 | 0.11 | 0.049 | 0.035 | 0.182 | 4e-05 | 2.3094e-05 | 0.5 | 0.28868 | 0.34126 | 0 | ∞ | 2 | 0.68252
50 | 0 | 0 | 0.17 | 0.4 | 0.042 | 0.05 | 0.331 | 5e-05 | 2.8868e-05 | 0.5 | 0.28868 | 0.4392 | 0 | ∞ | 2 | 0.87839
60 | 0 | 0 | 0.4 | 0.04 | 0.04 |  | 0.24 | 6e-05 | 3.4641e-05 | 0.5 | 0.28868 | 0.37541 | 0 | ∞ | 2 | 0.75082
70 | 0 | 0 | 0.4 | 0.049 | 0.042 | 0.05 | 0.2705 | 7e-05 | 4.0415e-05 | 0.5 | 0.28868 | 0.39561 | 0 | ∞ | 2 | 0.79121
80 | 0 | 0 | 0.4 | 0.11 | 0.049 | 0.042 | 0.3005 | 8e-05 | 4.6188e-05 | 0.5 | 0.28868 | 0.41669 | 0 | ∞ | 2 | 0.83339
90 | 0 | 0 | 0.4 | 0.17 | 0.042 | 0.05 | 0.331 | 9e-05 | 5.1962e-05 | 0.5 | 0.28868 | 0.4392 | 0 | ∞ | 2 | 0.87839
100 | 0 | 0 | 0.4 | 0.17 | 0.11 | 0.042 | 0.361 | 0.0001 | 5.7735e-05 | 0.5 | 0.28868 | 0.46223 | 0 | ∞ | 2 | 0.92446
|  |  | Table 2. | T.O. 33K6-4-1362-1
|  |  | Type |  | Point | Load Nominal |  | Load Nominal |  | Limits
|  |  | A, B and O |  | 20 | 2.05 | N | 209.04 | g | 201.39 | to | 216.69 g
[5 empty] | 30 | 2.8 | N | 285.52 | g | 277.87 | to | 293.17 g
[5 empty] | 40 | 3.55 | N | 362 | g | 354.35 | to | 369.65 g
[5 empty] | 50 | 4.3 | N | 438.48 | g | 430.83 | to | 446.13 g
[5 empty] | 60 | 5.05 | N | 514.96 | g | 507.31 | to | 522.61 g
[5 empty] | 70 | 5.8 | N | 591.44 | g | 583.79 | to | 599.09 g
[5 empty] | 80 | 6.55 | N | 667.91 | g | 660.26 | to | 675.56 g
[5 empty] | 90 | 7.3 | N | 744.39 | g | 736.74 | to | 752.04 g
[5 empty] | 100 | 8.05 | N | 820.87 | g | 813.22 | to | 828.52 g
ปี 2012
Nominal | [6 empty] | Variation
Length | [6 empty] | of GB.
( mm ) | [6 empty] | ( mm )
1 | [6 empty] | 0.03
1.005 | [6 empty] | 0.05
1.01 | [6 empty] | 0.04
1.02 | [6 empty] | 0.07
1.03 | [6 empty] | 0.04
1.04 | [6 empty] | 0.04
1.05 | [6 empty] | 0.06
1.06 | [6 empty] | 0.07
1.07 | [6 empty] | 0.08
1.08 | [6 empty] | 0.04
1.09 | [6 empty] | 0.04
1.1 | [6 empty] | 0.04
1.2 | [6 empty] | 0.14
1.3 | [6 empty] | 0.03
1.4 | [6 empty] | 0.07
1.5 | [6 empty] | 0.02
1.6 | [6 empty] | 0.07
1.7 | [6 empty] | 0.07
1.8 | [6 empty] | 0.06
1.9 | [6 empty] | 0.04
2 | [6 empty] | 0.03
3 | [6 empty] | 0.08
4 | [6 empty] | 0.06
5 | [6 empty] | 0.07
6 | [6 empty] | 0.05
7 | [6 empty] | 0.07
8 | [6 empty] | 0.01
9 | [6 empty] | 0.07
10 | [6 empty] | 0.06
11 | [6 empty] | 0.06
12 | [6 empty] | 0.02
13 | [6 empty] | 0.04
14 | [6 empty] | 0.05
15 | [6 empty] | 0.05
16 | [6 empty] | 0.07
17 | [6 empty] | 0.04
18 | [6 empty] | 0.05
19 | [6 empty] | 0.09
20 | [6 empty] | 0.08
21 | [6 empty] | 0.06
22 | [6 empty] | 0.04
23 | [6 empty] | 0.04
24 | [6 empty] | 0.13
25 | [6 empty] | 0.05
50 | [6 empty] | 0.14
75 | [6 empty] | 0.14
100 | [6 empty] | 0.08
125 | [6 empty] | 0.03
150 | [6 empty] | 0
175 | [6 empty] | 0
200 | [6 empty] | 0
250 | [6 empty] | 0
300 | [6 empty] | 0
400 | [6 empty] | 0
500 | [6 empty] | 0
600 | [6 empty] | 0
700 | [6 empty] | 0
800 | [6 empty] | 0
900 | [6 empty] | 0
1000 | [6 empty] | 0
Note on C6: ได้จากการวัดซ้ำ
Note on E6: ค่าจากใบรายงานผลการสอบเทียบ STD Weight
Note on J6: (1ppm*W)/SQRT3
Note on L6: ค่าความละเอียดในการอ่านของเครื่องชั่ง
Note on F7: ใช้ตุ้มน้ำหนัก2ค่ารวมกัน
Note on G7: ใช้ตุ้มน้ำหนัก3ค่ารวมกัน
Note on H7: ใช้ตุ้มน้ำหนัก4ค่ารวมกัน

## Uncert Budget(C, D, DO)
Uncertainty Budget of Durometer
Nominal | Repeatability |  | Uncertainty |  |  |  |  | Air Buoyancy effect |  | Resolution of |  | Uc | Ui | Veff | K95 | U95%
| of UUC |  | Standard Weight. |  |  |  |  | ± 1 ppm 10^6 |  | UUC
Point | Value | Ui | Value |  |  |  | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
20 | 0 | 0 | 0.4 | 0.17 | 0.4 | 0.035 | 0.0005025 | 2e-05 | 1.1547e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62500378959.9495 | 2.00 | 0.57735
30 | 0 | 0 | 1.7 | 0.17 | 0.11 | 0.084 | 0.001032 | 3e-05 | 1.7321e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62501597996.2143 | 2.00 | 0.57735
40 | 0 | 0 | 1.7 | 0.4 | 0.17 | 0.11 | 0.00119 | 4e-05 | 2.3094e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62502124968.0617 | 2.00 | 0.57736
50 | 0 | 0 | 2 | 0.17 | 0.049 | 0.042 | 0.0011305 | 5e-05 | 2.8868e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62501918310.0945 | 2.00 | 0.57735
60 | 0 | 0 | 2 | 0.4 | 0.17 | 0.042 | 0.001306 | 6e-05 | 3.4641e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62502560280.2197 | 2.00 | 0.57736
70 | 0 | 0 | 2 | 1.7 | 0.11 | 0.091 | 0.0019505 | 7e-05 | 4.0415e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62505709255.7515 | 2.00 | 0.57736
80 | 0 | 0 | 2 | 1.7 | 0.51 | 0.077 | 0.0021435 | 8e-05 | 4.6188e-05 | 0.5 | 0.28868 | 0.28868 | 0 | 62506895278.544 | 2.00 | 0.57737
90 | 0 | 0 | 2 | 4 | 0.049 | 0.042 | 0.0030455 | 9e-05 | 5.1962e-05 | 0.5 | 0.28868 | 0.28869 | 0 | 62513917430.0682 | 2.00 | 0.57738
100 | 0 | 0 | 2 | 4 | 0.4 | 0.077 | 0.0032385 | 0.0001 | 5.7735e-05 | 0.5 | 0.28868 | 0.28869 | 0 | 62515737813.9654 | 2.00 | 0.57739
|  |  | Table 2. | T.O. 33K6-4-1362-1
|  |  | Type |  | Point | Load Nominal |  | Load Nominal |  | Limits
|  |  | C, D and DO |  | 20 | 8.89 | N | 906.53 | g | 861.2 | to | 951.86 g
[5 empty] | 30 | 13.335 | N | 1.36 | kg | 1.315 | to | 1.405 kg
[5 empty] | 40 | 17.78 | N | 1.813 | kg | 1.768 | to | 1.858 kg
[5 empty] | 50 | 22.225 | N | 2.266 | kg | 2.221 | to | 2.311 kg
[5 empty] | 60 | 26.67 | N | 2.72 | kg | 2.675 | to | 2.765 kg
[5 empty] | 70 | 31.115 | N | 3.173 | kg | 3.128 | to | 3.218 kg
[5 empty] | 80 | 35.56 | N | 3.626 | kg | 3.581 | to | 3.671 kg
[5 empty] | 90 | 40.005 | N | 4.079 | kg | 4.034 | to | 4.124 kg
[5 empty] | 100 | 44.45 | N | 4.53 | kg | 4.485 | to | 4.575 kg
ปี 2012
Nominal | [6 empty] | Variation
Length | [6 empty] | of GB.
( mm ) | [6 empty] | ( mm )
1 | [6 empty] | 0.03
1.005 | [6 empty] | 0.05
1.01 | [6 empty] | 0.04
1.02 | [6 empty] | 0.07
1.03 | [6 empty] | 0.04
1.04 | [6 empty] | 0.04
1.05 | [6 empty] | 0.06
1.06 | [6 empty] | 0.07
1.07 | [6 empty] | 0.08
1.08 | [6 empty] | 0.04
1.09 | [6 empty] | 0.04
1.1 | [6 empty] | 0.04
1.2 | [6 empty] | 0.14
1.3 | [6 empty] | 0.03
1.4 | [6 empty] | 0.07
1.5 | [6 empty] | 0.02
1.6 | [6 empty] | 0.07
1.7 | [6 empty] | 0.07
1.8 | [6 empty] | 0.06
1.9 | [6 empty] | 0.04
2 | [6 empty] | 0.03
3 | [6 empty] | 0.08
4 | [6 empty] | 0.06
5 | [6 empty] | 0.07
6 | [6 empty] | 0.05
7 | [6 empty] | 0.07
8 | [6 empty] | 0.01
9 | [6 empty] | 0.07
10 | [6 empty] | 0.06
11 | [6 empty] | 0.06
12 | [6 empty] | 0.02
13 | [6 empty] | 0.04
14 | [6 empty] | 0.05
15 | [6 empty] | 0.05
16 | [6 empty] | 0.07
17 | [6 empty] | 0.04
18 | [6 empty] | 0.05
19 | [6 empty] | 0.09
20 | [6 empty] | 0.08
21 | [6 empty] | 0.06
22 | [6 empty] | 0.04
23 | [6 empty] | 0.04
24 | [6 empty] | 0.13
25 | [6 empty] | 0.05
50 | [6 empty] | 0.14
75 | [6 empty] | 0.14
100 | [6 empty] | 0.08
125 | [6 empty] | 0.03
150 | [6 empty] | 0
175 | [6 empty] | 0
200 | [6 empty] | 0
250 | [6 empty] | 0
300 | [6 empty] | 0
400 | [6 empty] | 0
500 | [6 empty] | 0
600 | [6 empty] | 0
700 | [6 empty] | 0
800 | [6 empty] | 0
900 | [6 empty] | 0
1000 | [6 empty] | 0
Note on C6: ได้จากการวัดซ้ำ
Note on E6: ค่าจากใบรายงานผลการสอบเทียบ STD Weight
Note on J6: (1ppm*W)/SQRT3
Note on L6: ค่าความละเอียดในการอ่านของเครื่องชั่ง
Note on F7: ใช้ตุ้มน้ำหนัก2ค่ารวมกัน
Note on G7: ใช้ตุ้มน้ำหนัก3ค่ารวมกัน
Note on H7: ใช้ตุ้มน้ำหนัก4ค่ารวมกัน

## Cert of STD
Weight 10 mg - 500 mg |  |  |  | Weight 1g - 5 kg |  |  |  | Weight 10 g - 20 kg |  |  |  | Weight 5 kg - 20 kg
SP-SM-001 |  |  |  | SP-SM-002 |  |  |  | SP-SM-003 |  |  |  | SP-SM-004
Due Date |  |  |  | Due Date |  |  |  | Due Date |  |  |  | No. | Due Date
2016-06-29 |  |  |  | 2016-06-29 |  |  |  | 2016-08-19 |  |  |  |  | 2016-12-02
10 | mg | 0.01 | mg | 1 | g | 0.02 | mg | 10 | g | 0.035 | mg | 1 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg | 2 | g* | 0.03 | mg* | 20 | g | 0.042 | mg | 2 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg* | 2 | g | 0.03 | mg | 50 | g | 0.049 | mg | 1 | 10 | kg | 31 | mg
50 | mg | 0.01 | mg | 5 | g | 0.04 | mg | 100 | g | 0.11 | mg | 2 | 10 | kg | 31 | mg
100 | mg | 0.01 | mg | 10 | g | 0.04 | mg | 200 | g | 0.17 | mg | 1 | 20 | kg | 54 | mg
200 | mg | 0.01 | mg | 20 | g | 0.05 | mg | 500 | g | 0.4 | mg | 2 | 20 | kg | 54 | mg
500 | mg | 0.01 | mg | 20 | g* | 0.05 | mg* | 1 | kg | 1.7 | mg | 3 | 20 | kg | 54 | mg
|  |  |  | 50 | g | 0.1 | mg | 2 | kg | 2 | mg | 4 | 20 | kg | 54 | mg
|  |  |  | 100 | g | 0.2 | mg | 10 | kg | 12 | mg | 5 | 20 | kg | 54 | mg
|  |  |  | 200 | g | 0.4 | mg | 20 | kg | 18 | mg | 6 | 20 | kg | 54 | mg
|  |  |  | 200 | g* | 0.4 | mg* |  |  |  |  | 7 | 20 | kg | 54 | mg
|  |  |  | 500 | g | 2 | mg |  |  |  |  | 8 | 20 | kg | 54 | mg
|  |  |  | 1 | kg | 2 | mg |  |  |  |  | 9 | 20 | kg | 54 | mg
|  |  |  | 1 | kg* | 2 | mg |  |  |  |  | 10 | 20 | kg | 54 | mg
|  |  |  | 2 | kg | 4 | mg |  |  |  |  | 11 | 20 | kg | 54 | mg
|  |  |  | 5 | kg | 10 | mg |  |  |  |  | 12 | 20 | kg | 54 | mg
[12 empty] | 13 | 20 | kg | 54 | mg
[12 empty] | 14 | 20 | kg | 54 | mg
[12 empty] | 15 | 20 | kg | 54 | mg
[12 empty] | 16 | 20 | kg | 54 | mg
[12 empty] | 17 | 20 | kg | 54 | mg
[12 empty] | 18 | 20 | kg | 54 | mg
[12 empty] | 19 | 20 | kg | 54 | mg
[12 empty] | 20 | 20 | kg | 54 | mg
[12 empty] | 21 | 20 | kg | 54 | mg
[12 empty] | 22 | 20 | kg | 54 | mg
[12 empty] | 23 | 20 | kg | 54 | mg
[12 empty] | 24 | 20 | kg | 54 | mg
[12 empty] | 25 | 20 | kg | 54 | mg
[12 empty] | 26 | 20 | kg | 54 | mg
[12 empty] | 27 | 20 | kg | 54 | mg
[12 empty] | 28 | 20 | kg | 54 | mg
[12 empty] | 29 | 20 | kg | 54 | mg
[12 empty] | 30 | 20 | kg | 54 | mg
[12 empty] | 31 | 20 | kg | 54 | mg
[12 empty] | 32 | 20 | kg | 54 | mg
[12 empty] | 33 | 20 | kg | 54 | mg
[12 empty] | 34 | 20 | kg | 54 | mg
[12 empty] | 35 | 20 | kg | 54 | mg
[12 empty] | 36 | 20 | kg | 54 | mg
[12 empty] | 37 | 20 | kg | 54 | mg
[12 empty] | 38 | 20 | kg | 54 | mg
[12 empty] | 39 | 20 | kg | 54 | mg
[12 empty] | 40 | 20 | kg | 54 | mg
[12 empty] | 41 | 20 | kg | 54 | mg
[12 empty] | 42 | 20 | kg | 54 | mg
[12 empty] | 43 | 20 | kg | 54 | mg
[12 empty] | 44 | 20 | kg | 54 | mg
[12 empty] | 45 | 20 | kg | 54 | mg
[12 empty] | 46 | 20 | kg | 54 | mg
[12 empty] | 47 | 20 | kg | 54 | mg
[12 empty] | 48 | 20 | kg | 54 | mg
[12 empty] | 49 | 20 | kg | 54 | mg
[12 empty] | 50 | 20 | kg | 54 | mg
Note on B6: MPE Class F1
Note on G6: MPE Class F1
Note on L6: MPE Class E2
Note on Q6: MPE Class M1
Note on B8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B9: Nathaphol Boonmee: Normal Value
Note on D9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G9: Nathaphol Boonmee: Normal Value
Note on I9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L9: Nathaphol Boonmee: Normal Value
Note on N9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R9: Nathaphol Boonmee: Normal Value
Note on T9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง