## Data Record
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [8 empty] | Page | 1 | of | 1
[9 empty] | Receive Date : |  |  |  | 2016-12-12 |  |  |  |  | Calibration Date : |  |  | 2016-12-12
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Vernier Dial and Digital Calipers | [8 empty] | Location | [5 empty] | In Lab | [6 empty] | On Site
Customer Name : |  |  | SCG
Equipment Name : |  |  |  | Vernier Dial | [5 empty] | Manufacturer : |  |  |  | Mitotoyo
Model : |  | 123 |  |  |  |  | Serial No. : |  | 456 | [6 empty] | ID No : |  | 789
Range : |  | 0 | to | 150 |  | mm |  |  |  | Resolution : | 0.01 |  | mm |  |  | Readability : |  |  | 0.005 | mm
Overall Inspection | [5 empty] | Good |  |  | Not Good
Referance Standard : | [11 empty] | Due Date :
Referance Standard : | [11 empty] | Due Date :
1. Parallelism of measuring faces
1.1 External jaw
|  | Normal |  | UUC Reading |  |  |  |  | Max-Min
|  | Value |  | Root |  | Center |  | Tip
|  | 15 |  | 15 |  | 14.99 |  | 14.99 | 0.01
|  | 75 |  | 75 |  | 75 |  | 75 | 0
|  | 150 |  | 150 |  | 150.01 |  | 150.01 | 0.01
|  | Max | [6 empty] | 0.01
2. External Measurement
|  | Normal |  | UUC Reading | [13 empty] | Repeatability |  |  | Error
|  | Value |  | X1 |  |  | X2 |  | X3 |  |  | X4 |  |  | Average
|  | 0 |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
|  | 1 |  | 1 |  |  | 1 |  | 1 |  |  | 1 |  |  | 1 |  |  | 0 |  |  | 0
|  | 1.5 |  | 1.5 |  |  | 1.5 |  | 1.5 |  |  | 1.5 |  |  | 1.5 |  |  | 0 |  |  | 0
|  | 5 |  | 5 |  |  | 5 |  | 5 |  |  | 5 |  |  | 5 |  |  | 0 |  |  | 0
|  | 10 |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 0 |  |  | 0
|  | 20 |  | 20 |  |  | 20 |  | 20 |  |  | 20 |  |  | 20 |  |  | 0 |  |  | 0
|  | 50 |  | 50 |  |  | 50 |  | 50 |  |  | 50 |  |  | 50 |  |  | 0 |  |  | 0
|  | 100 |  | 100 |  |  | 100 |  | 100 |  |  | 100 |  |  | 100 |  |  | 0 |  |  | 0
|  | 150 |  | 150 |  |  | 150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  |  | 0
|  | 200 |  | 200 |  |  | 200 |  | 200 |  |  | 200 |  |  | 200 |  |  | 0 |  |  | 0
|  | 250 |  | 250 |  |  | 250 |  | 250 |  |  | 250 |  |  | 250 |  |  | 0 |  |  | 0
|  | 300 |  | 300 |  |  | 300 |  | 300 |  |  | 300 |  |  | 300 |  |  | 0 |  |  | 0
|  | 400 |  | 400 |  |  | 400 |  | 400 |  |  | 400 |  |  | 400 |  |  | 0 |  |  | 0
|  | 500 |  | 500 |  |  | 500 |  | 500 |  |  | 500 |  |  | 500 |  |  | 0 |  |  | 0
|  | 600 |  | 600 |  |  | 600 |  | 600 |  |  | 600 |  |  | 600 |  |  | 0 |  |  | 0
|  | 700 |  | 700 |  |  | 700 |  | 700 |  |  | 700 |  |  | 700 |  |  | 0 |  |  | 0
|  | 800 |  | 800 |  |  | 800 |  | 800 |  |  | 800 |  |  | 800 |  |  | 0 |  |  | 0
|  | 900 |  | 900 |  |  | 900 |  | 900 |  |  | 900 |  |  | 900 |  |  | 0 |  |  | 0
|  | 1000 |  | 1000 |  |  | 1000 |  | 1000 |  |  | 1000 |  |  | 1000 |  |  | 0 |  |  | 0
|  | 1500 |  | 1500.01 |  |  | 1500.01 |  | 1500.01 |  |  | 1500.01 |  |  | 1500.01 |  |  | 0 |  |  | 0.01
Calibrated By : |  |  |  |  | Ms. Arunkamon Raramanus
|  |  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120012-1 | [7 empty] | Page : 1 of 5
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Vernier Dial
| Manufacturer |  |  |  | : | Mitotoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 456
| ID. Number |  |  |  | : | 789
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-12-12
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2016-12-12
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2017-12-12
| Calibration Procedure |  |  |  | : | SP-CPT-04-02
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-12-13 |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 10 | Mr.Chainarong Matchayamat | [7 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | #REF! |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  | #REF! | [9 empty] | Page : 3 of 3
|  | Range : | 0 | to | 150 | mm |  | Resolution : |  |  | 0.01 |  | mm
|  |  | Measurement Result | [6 empty] | Unit : |  | mm
|  |  | Normal Value |  | UUC Reading |  | Error |  | Uncertainty ( ± )
|  |  | 0 |  | 0 |  | 0 |  | 0.0057738
|  |  | 1 |  | 1 |  | 0 |  | 0.0057738
|  |  | 1.5 |  | 1.5 |  | 0 |  | 0.0057738
|  |  | 5 |  | 5 |  | 0 |  | 0.0057742
|  |  | 10 |  | 10 |  | 0 |  | 0.0057753
|  |  | 20 |  | 20 |  | 0 |  | 0.00578
|  |  | 50 |  | 50 |  | 0 |  | 0.0058123
|  |  | 100 |  | 100 |  | 0 |  | 0.0059255
|  |  | 150 |  | 150 |  | 0 |  | 0.0061255
|  |  | 200 |  | 200 |  | 0 |  | 0.0063788
|  |  | 250 |  | 250 |  | 0 |  | 0.0066896
|  |  | 300 |  | 300 |  | 0 |  | 0.0070503
|  |  | 400 |  | 400 |  | 0 |  | 0.0078955
|  |  | 500 |  | 500 |  | 0 |  | 0.0088672
|  |  | 600 |  | 600 |  | 0 |  | 0.0099147
|  |  | 700 |  | 700 |  | 0 |  | 0.011066
|  |  | 800 |  | 800 |  | 0 |  | 0.012225
|  |  | 900 |  | 900 |  | 0 |  | 0.013421
|  |  | 1000 |  | 1000 |  | 0 |  | 0.014633
|  |  | 1500 |  | 1500.01 |  | 0.01 |  | 0.02101
Parallelism between external measuring face by measure at the root,
|  |  | center and tip of callipers was found to be | [6 empty] | 0.01 | mm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -

## Uncertainty Budget (Ext)
Uncertainty Budget of Vernier, Dial and Digital Callipers
Nominal Value | Uncertainty of STD |  | Temperature Effect |  | Resolution of UUC |  | Repeatability |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (mm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 6e-05 | 3e-05 | 0 | 0 | 0.005 | 0.0028868 | 0 | 0 | 0.0028869 | 0 | ∞ | 2 | 0.0057738
1 | 6e-05 | 3e-05 | 1.15e-05 | 6.6395e-06 | 0.005 | 0.0028868 | 0 | 0 | 0.0028869 | 0 | ∞ | 2 | 0.0057738
1.5 | 6e-05 | 3e-05 | 1.725e-05 | 9.9593e-06 | 0.005 | 0.0028868 | 0 | 0 | 0.0028869 | 0 | ∞ | 2 | 0.0057738
5 | 6e-05 | 3e-05 | 5.75e-05 | 3.3198e-05 | 0.005 | 0.0028868 | 0 | 0 | 0.0028871 | 0 | ∞ | 2 | 0.0057742
10 | 6e-05 | 3e-05 | 0.000115 | 6.6395e-05 | 0.005 | 0.0028868 | 0 | 0 | 0.0028877 | 0 | ∞ | 2 | 0.0057753
20 | 7e-05 | 3.5e-05 | 0.00023 | 0.00013279 | 0.005 | 0.0028868 | 0 | 0 | 0.00289 | 0 | ∞ | 2 | 0.00578
50 | 9e-05 | 4.5e-05 | 0.000575 | 0.00033198 | 0.005 | 0.0028868 | 0 | 0 | 0.0029061 | 0 | ∞ | 2 | 0.0058123
100 | 0.00012 | 6e-05 | 0.00115 | 0.00066395 | 0.005 | 0.0028868 | 0 | 0 | 0.0029627 | 0 | ∞ | 2 | 0.0059255
150 | 0.00047 | 0.000235 | 0.001725 | 0.00099593 | 0.005 | 0.0028868 | 0 | 0 | 0.0030627 | 0 | ∞ | 2 | 0.0061255
200 | 0.00055 | 0.000275 | 0.0023 | 0.0013279 | 0.005 | 0.0028868 | 0 | 0 | 0.0031894 | 0 | ∞ | 2 | 0.0063788
250 | 0.00063 | 0.000315 | 0.002875 | 0.0016599 | 0.005 | 0.0028868 | 0 | 0 | 0.0033448 | 0 | ∞ | 2 | 0.0066896
300 | 0.00071 | 0.000355 | 0.00345 | 0.0019919 | 0.005 | 0.0028868 | 0 | 0 | 0.0035252 | 0 | ∞ | 2 | 0.0070503
400 | 0.00089 | 0.000445 | 0.0046 | 0.0026558 | 0.005 | 0.0028868 | 0 | 0 | 0.0039477 | 0 | ∞ | 2 | 0.0078955
500 | 0.0011 | 0.00055 | 0.00575 | 0.0033198 | 0.005 | 0.0028868 | 0 | 0 | 0.0044336 | 0 | ∞ | 2 | 0.0088672
600 | 0.00122 | 0.00061 | 0.0069 | 0.0039837 | 0.005 | 0.0028868 | 0 | 0 | 0.0049574 | 0 | ∞ | 2 | 0.0099147
700 | 0.00165 | 0.000825 | 0.00805 | 0.0046477 | 0.005 | 0.0028868 | 0 | 0 | 0.0055331 | 0 | ∞ | 2 | 0.011066
800 | 0.00181 | 0.000905 | 0.0092 | 0.0053116 | 0.005 | 0.0028868 | 0 | 0 | 0.0061127 | 0 | ∞ | 2 | 0.012225
900 | 0.00199 | 0.000995 | 0.01035 | 0.0059756 | 0.005 | 0.0028868 | 0 | 0 | 0.0067105 | 0 | ∞ | 2 | 0.013421
1000 | 0.00211 | 0.001055 | 0.0115 | 0.0066395 | 0.005 | 0.0028868 | 0 | 0 | 0.0073164 | 0 | ∞ | 2 | 0.014633
1500 | 0.00337 | 0.001685 | 0.01725 | 0.0099593 | 0.005 | 0.0028868 | 0 | 0 | 0.010505 | 0 | ∞ | 2 | 0.02101
Note on D4: Certificate of Calibration Gauge Block
Note on F4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 2
Note on G6: Divisor = √3
Note on I6: Divisor = √3
Note on K6: Divisor = 1

## Cert of STD
Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block)
SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[5 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[5 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[5 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[5 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[5 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[5 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[5 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[5 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[5 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[5 empty] | 14 | 0.07 | µm | 7e-05 | mm
[5 empty] | 15 | 0.07 | µm | 7e-05 | mm
[5 empty] | 16 | 0.07 | µm | 7e-05 | mm
[5 empty] | 17 | 0.07 | µm | 7e-05 | mm
[5 empty] | 18 | 0.07 | µm | 7e-05 | mm
[5 empty] | 19 | 0.07 | µm | 7e-05 | mm
[5 empty] | 20 | 0.07 | µm | 7e-05 | mm
[5 empty] | 21 | 0.07 | µm | 7e-05 | mm
[5 empty] | 22 | 0.07 | µm | 7e-05 | mm
[5 empty] | 23 | 0.07 | µm | 7e-05 | mm
[5 empty] | 24 | 0.07 | µm | 7e-05 | mm
[5 empty] | 25 | 0.07 | µm | 7e-05 | mm
[5 empty] | 50 | 0.09 | µm | 9e-05 | mm
[5 empty] | 75 | 0.1 | µm | 0.0001 | mm
[5 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on C5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on E5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on K5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน