## Data Record
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  | SPR16050023-8 | [5 empty] | Page |  |  | of
[8 empty] | Receive Date : |  |  | 2016-03-21 |  |  |  | Calibration Date : |  |  |  | 2016-03-21
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  | 20 | oC | 50 |  | %RH
Steel Ruler | [7 empty] | Location |  |  | In Lab | [6 empty] | On Site
Customer Name :
Equipment Name : |  |  |  | Steel Ruler | [5 empty] | Manufacturer : |  |  |  | EKA
Model : |  | 123 | [5 empty] | Serial No. : | 123456 | [6 empty] | ID No : |  | N/A
Range : |  | 0 |  | to | 2000 |  | mm |  |  |  | Resolution : | 10 | cm
Overall Inspection | [5 empty] | Good |  | Not Good
Referance Standard : | [11 empty] | Due Date :
Referance Standard : | [11 empty] | Due Date :
Nominal Value | Position | Standard Reading | [10 empty] | Average |  |  |  | Repeatability |  |  | Error
|  | X1 |  |  |  | X2 |  | X3 |  | X4
0 | L1 | 0 |  |  |  | 0 |  | 0 |  | 0 |  |  | 0.00015 |  |  |  | 8.0178e-05 |  |  | -0.00015
| L2 | 0.0003 |  |  |  | 0.0003 |  | 0.0003 |  | 0.0003
200 | L1 | 199.9988 |  |  |  | 199.9988 |  | 199.9988 |  | 199.9988 |  |  | 199.9994 |  |  |  | 0.00034744 |  |  | 0.00055
| L2 | 200.0001 |  |  |  | 200.0001 |  | 200.0001 |  | 200.0001
400 | L1 | 399.9996 |  |  |  | 399.9994 |  | 399.9996 |  | 399.9995 |  |  | 399.9996 |  |  |  | 7.0711e-05 |  |  | 0.0004
| L2 | 399.9996 |  |  |  | 399.9996 |  | 399.9999 |  | 399.9996
600 | L1 | 59.9992 |  |  |  | 59.9992 |  | 59.9992 |  | 59.999 |  |  | 59.9992 |  |  |  | 4.6291e-05 |  |  | 540.0009
| L2 | 59.9992 |  |  |  | 59.9992 |  | 59.9992 |  | 59.999
800 | L1 | 79.9986 |  |  |  | 79.9985 |  | 79.9982 |  | 79.9988 |  |  | 79.9985 |  |  |  | 0.00011573 |  |  | 720.0015
| L2 | 79.9986 |  |  |  | 79.9985 |  | 79.9982 |  | 79.9988
1000 | L1 | 99.998 |  |  |  | 99.9982 |  | 99.9979 |  | 99.9982 |  |  | 99.9981 |  |  |  | 6.9437e-05 |  |  | 900.0019
| L2 | 99.998 |  |  |  | 99.9982 |  | 99.9979 |  | 99.9982
1200 | L1 | 1199.983 |  |  |  | 1199.982 |  | 1199.982 |  | 1199.981 |  |  | 1199.9909 |  |  |  | 0.0047914 |  |  | 0.00905
| L2 | 1199.9999 |  |  |  | 1199.9999 |  | 1199.9999 |  | 1199.9999
1400 | L1 | 1399.9984 |  |  |  | 1399.9984 |  | 1399.9984 |  | 1399.9984 |  |  | 1399.9994 |  |  |  | 0.0005078 |  |  | 0.00065
| L2 | 1400.0003 |  |  |  | 1400.0003 |  | 1400.0003 |  | 1400.0003
1600 | L1 | 1599.998 |  |  |  | 1599.998 |  | 1599.998 |  | 1599.9979 |  |  | 1600.0001 |  |  |  | 0.001137 |  |  | -0.0001
| L2 | 1600.0023 |  |  |  | 1600.0021 |  | 1600.0024 |  | 1600.0021
1800 | L1 | 1799.9983 |  |  |  | 1799.9983 |  | 1799.9983 |  | 1799.9983 |  |  | 1799.9998 |  |  |  | 0.00079796 |  |  | 0.0002125
| L2 | 1800.0013 |  |  |  | 1800.0013 |  | 1800.0015 |  | 1800.001
2000 | L1 | 1999.998 |  |  |  | 1999.998 |  | 1999.998 |  | 1999.998 |  |  | 1999.9982 |  |  |  | 0.00014392 |  |  | 0.0018
| L2 | 1999.9981 |  |  |  | 1999.9984 |  | 1999.9988 |  | 1999.9983
Calibrated By : |  |  | Ms. Arunkamon Raramanus
|  |  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16050023-8 | [5 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Steel Ruler
| Manufacturer |  |  |  | : | 0
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 123456
| ID. Number |  |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  | : | 2016-03-21
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  | : | 2016-03-21
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  | : | 2017-03-21
| Calibration Procedure |  |  |  |  | : | SP-CPT-04-13
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-03-22 |  |  |  | Approved by :
| Calibrated by |  | : | #REF! |  |  |  |  | 3 |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16050023-8 |  |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  |  | Model | Serial No. | Certificate No. | Due. Date
| 3D Vision Measuring Machine |  |  |  | D3020-T | 110021 | MTO150453-1 | 2016-06-26
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR16050023-8 |  |  |  |  | Page : 3 of 3
[7 empty] | Unit : |  | mm
|  | Nominal Value |  | STD Reading | Error | Uncertainty ( ± ) µm
|  | 0 |  | 0.00015 | 0.00015 | 3.0006
|  | 200 |  | 199.9994 | -0.00055 | 4.0071
|  | 400 |  | 399.9996 | -0.0004 | 6.1015
|  | 600 |  | 59.9992 | -540.0009 | 8.5143
|  | 800 |  | 79.9985 | -720.0015 | 11.0417
|  | 1000 |  | 99.9981 | -900.0019 | 13.6147
|  | 1200 |  | 1199.9909 | -0.00905 | 16.2201
|  | 1400 |  | 1399.9994 | -0.00065 | 18.8313
|  | 1600 |  | 1600.0001 | 0.0001 | 21.4575
|  | 1800 |  | 1799.9998 | -0.0002125 | 24.0902
|  | 2000 |  | 1999.9982 | -0.0018 | 26.7275
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Steel Ruler
Nominal Value | Repeatability |  | Uncert of 3D Vision |  | Temperature Effect |  | Resolution of 3D Vision |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.003 | 0.0015 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0015003 | 0 | ∞ | 2 | 3.0006
200 | 0 | 0 | 0.003 | 0.0015 | 0.0023 | 0.0013279 | 5e-05 | 2.8868e-05 | 0.0020035 | 0 | ∞ | 2 | 4.0071
400 | 5.4006e-05 | 5.4006e-05 | 0.003 | 0.0015 | 0.0046 | 0.0026558 | 5e-05 | 2.8868e-05 | 0.0030508 | 5.4006e-05 | 30547443.0009 | 2 | 6.1015
600 | 5e-05 | 5e-05 | 0.003 | 0.0015 | 0.0069 | 0.0039837 | 5e-05 | 2.8868e-05 | 0.0042572 | 5e-05 | 157658501.3348 | 2 | 8.5143
800 | 0.000125 | 0.000125 | 0.003 | 0.0015 | 0.0092 | 0.0053116 | 5e-05 | 2.8868e-05 | 0.0055209 | 0.000125 | 11415769.4978 | 2 | 11.0417
1000 | 7.5e-05 | 7.5e-05 | 0.003 | 0.0015 | 0.0115 | 0.0066395 | 5e-05 | 2.8868e-05 | 0.0068073 | 7.5e-05 | 203603085.452 | 2 | 13.6147
1200 | 0.00020412 | 0.00020412 | 0.003 | 0.0015 | 0.0138 | 0.0079674 | 5e-05 | 2.8868e-05 | 0.00811 | 0.00020412 | 7475365.5955 | 2 | 16.2201
1400 | 0 | 0 | 0.003 | 0.0015 | 0.0161 | 0.0092953 | 5e-05 | 2.8868e-05 | 0.0094156 | 0 | ∞ | 2 | 18.8313
1600 | 4.5644e-05 | 4.5644e-05 | 0.003 | 0.0015 | 0.0184 | 0.010623 | 5e-05 | 2.8868e-05 | 0.010729 | 4.5644e-05 | 9158019012.1952 | 2 | 21.4575
1800 | 5.1539e-05 | 5.1539e-05 | 0.003 | 0.0015 | 0.0207 | 0.011951 | 5e-05 | 2.8868e-05 | 0.012045 | 5.1539e-05 | 8949923951.7633 | 2 | 24.0902
2000 | 7.3598e-05 | 7.3598e-05 | 0.003 | 0.0015 | 0.023 | 0.013279 | 5e-05 | 2.8868e-05 | 0.013364 | 7.3598e-05 | 3261138790.615 | 2 | 26.7275
Note
The uncertainty that has multiplied the number of times the Set STD 3D Vision as the distance 2000 mm Number Set STD. Electronic Scale 2000 mm at a distance of 25 times, so the uncertainty of × 7.
[14 empty] | Nominal Value | Position |  |  |  |  | σn-1
[16 empty] | 1 | 2 | 3 | 4
[14 empty] | 0 | L1 | 0.00015 | 0.00015 | 0.00015 | 0.00015 | 0
L2
[14 empty] | 200 | L1 | 199.9994 | 199.9994 | 199.9994 | 199.9994 | 0
L2
[14 empty] | 400 | L1 | 399.9996 | 399.9995 | 399.9998 | 399.9995 | 0.00010801
L2
[14 empty] | 600 | L1 | 59.9992 | 59.9992 | 59.9992 | 59.999 | 0.0001
L2
[14 empty] | 800 | L1 | 79.9986 | 79.9985 | 79.9982 | 79.9988 | 0.00025
L2
[14 empty] | 1000 | L1 | 99.998 | 99.9982 | 99.9979 | 99.9982 | 0.00015
L2
[14 empty] | 1200 | L1 | 1199.9914 | 1199.9909 | 1199.9909 | 1199.9905 | 0.00040825
L2
[14 empty] | 1400 | L1 | 1399.9994 | 1399.9994 | 1399.9994 | 1399.9994 | 0
L2
[14 empty] | 1600 | L1 | 1600.0002 | 1600.0001 | 1600.0002 | 1600 | 9.1287e-05
L2
[14 empty] | 1800 | L1 | 1799.9998 | 1799.9998 | 1799.9999 | 1799.9996 | 0.00010308
L2
[14 empty] | 2000 | L1 | 1999.9981 | 1999.9982 | 1999.9984 | 1999.9981 | 0.0001472
L2
Note on F4: Certificate of Calibration 3D Vision
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3

## Uncert of STD
Certificate of Calibration (3D) |  |  |  |  | Certificate of Calibration (3D)
SP-SD-005 |  |  |  |  | SP-SD-005
Due Date |  | 2015-06-26 |  |  | Due Date |  | 2015-06-26
X - Axis |  |  |  |  | Y - Axis
0 | 3 | µm | 0.003 | mm | 0 | 3 | µm | 0.003 | mm
5 | 3 | µm | 0.003 | mm | 5 | 3 | µm | 0.003 | mm
10 | 3 | µm | 0.003 | mm | 10 | 3 | µm | 0.003 | mm
30 | 3 | µm | 0.003 | mm | 30 | 3 | µm | 0.003 | mm
50 | 3 | µm | 0.003 | mm | 50 | 3 | µm | 0.003 | mm
80 | 3 | µm | 0.003 | mm | 80 | 3 | µm | 0.003 | mm
100 | 3 | µm | 0.003 | mm | 100 | 3 | µm | 0.003 | mm
150 | 3 | µm | 0.003 | mm | 150 | 3 | µm | 0.003 | mm
200 | 3 | µm | 0.003 | mm | 200 | 3 | µm | 0.003 | mm
250 | 3 | µm | 0.003 | mm
300 | 3 | µm | 0.003 | mm
Note on A5: Grade 0
Note on G5: Grade 0
Note on C6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B8: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on D8: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on H8: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on J8: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน