## Data Record (Lenght)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR15120045-1 | [6 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Measuring Micro Scope | [7 empty] | Location |  |  |  |  | In Lab |  |  |  | On Site
Customer Name : |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  | Measuring Micro Scope | [6 empty] | Manufacturer : |  |  |  | OPI
Model : | 123 |  |  |  | Serial No. : |  |  | 456 |  |  |  | ID No : |  | 789 |  |  |  | Resolution : |  |  | 0.0001 |  | mm
Range : | X-axis |  | 0 | to | 300 |  | mm |  | Y-axis |  | 0 | to | 300 |  | mm |  | Z-axis |  | 0 | to | 300 |  | mm
Overall Inspection | [5 empty] | Good |  |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Referance Standard : | [13 empty] | Due Date :
1. Function Scale measurement X-axis direction.
Nominal Value. |  | UUC Reading | [11 empty] | Average |  |  | Error |  |  | Repeatability
|  | X1 |  |  | X2 |  |  | X3 |  |  | X4
0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
30 |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 0.001 |  |  | 0
60 |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 0.001 |  |  | 0
90 |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 0.001 |  |  | 0
120 |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 0.001 |  |  | 0
150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  |  | 0
180 |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 0.003 |  |  | 0
210 |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 0.003 |  |  | 0
240 |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 0.003 |  |  | 0
270 |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 0.003 |  |  | 0
300 |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 0.005 |  |  | 0
2. Function Scale measurement Y-axis direction.
Nominal Value. |  | UUC Reading | [11 empty] | Average |  |  | Error |  |  | Repeatability
|  | X1 |  |  | X2 |  |  | X3 |  |  | X4
0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
30 |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 0.001 |  |  | 0
60 |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 0.001 |  |  | 0
90 |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 0.001 |  |  | 0
120 |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 0.001 |  |  | 0
150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  |  | 0
180 |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 0.003 |  |  | 0
210 |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 0.003 |  |  | 0
240 |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 0.003 |  |  | 0
270 |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 0.003 |  |  | 0
300 |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 0.005 |  |  | 0
3. Function Scale measurement Z-axis direction.
Nominal Value. |  | UUC Reading | [11 empty] | Average |  |  | Error |  |  | Repeatability
|  | X1 |  |  | X2 |  |  | X3 |  |  | X4
0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
30 |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 30.001 |  |  | 0.001 |  |  | 0
60 |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 60.001 |  |  | 0.001 |  |  | 0
90 |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 90.001 |  |  | 0.001 |  |  | 0
120 |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 120.001 |  |  | 0.001 |  |  | 0
150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  |  | 0
180 |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 180.003 |  |  | 0.003 |  |  | 0
210 |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 210.003 |  |  | 0.003 |  |  | 0
240 |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 240.003 |  |  | 0.003 |  |  | 0
270 |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 270.003 |  |  | 0.003 |  |  | 0
300 |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 300.005 |  |  | 0.005 |  |  | 0
Calibrated By : |  |  | Mr.

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120045-1 | [6 empty] | Page : 1 of 4
| Customer |  |  |  | : | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
| Equipment Name |  |  |  | : | Measuring Micro Scope
| Manufacturer |  |  |  | : | OPI
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 456
| ID. Number |  |  |  | : | 789
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  |  | : | 2016-01-08
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-01-13
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2017-01-12
| Calibration Procedure |  |  |  |  | : | SP-CPD-04-17
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-14 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
Certificate Number |  | : | SPR16070004-24 |  |  |  | Page : 2 of 4
Reference Standards
Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
Standard Glass Scale |  |  | GS010410 | 000400 | DL-0050-15 | 2017-05-15
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. |  | : | SPR16070004-1 |  |  |  |  | Page : 3 of 4
|  | 1. Function Measurement : Measuring X - Axis Measure | [5 empty] | Unit : | mm
|  |  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  |  | 0 |  | 0 | 0 | 2.0008
|  |  | 30.001 |  | 30.001 | 0.001 | 2.0401
|  |  | 60.001 |  | 60.001 | 0.001 | 2.1536
|  |  | 90.001 |  | 90.001 | 0.001 | 2.3306
|  |  | 120.001 |  | 120.001 | 0.001 | 2.5578
|  |  | 150 |  | 150 | 0 | 2.8233
|  |  | 180.003 |  | 180.003 | 0.003 | 3.1171
|  |  | 210.003 |  | 210.003 | 0.003 | 5.7989
|  |  | 240.003 |  | 240.003 | 0.003 | 6.0125
|  |  | 270.003 |  | 270.003 | 0.003 | 6.2328
|  |  | 300.005 |  | 300.005 | 0.005 | 6.4812
|  | 2. Function Measurement : Measuring Y - Axis Measure | [5 empty] | Unit : | mm
|  |  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  |  | 0 |  | 0 | 0 | Value
|  |  | 30.001 |  | 30.001 | 0.001 | 2.0008
|  |  | 60.001 |  | 60.001 | 0.001 | 2.0401
|  |  | 90.001 |  | 90.001 | 0.001 | 2.1536
|  |  | 120.001 |  | 120.001 | 0.001 | 2.3306
|  |  | 150 |  | 150 | 0 | 2.5578
|  |  | 180.003 |  | 180.003 | 0.003 | 2.8233
|  |  | 210.003 |  | 210.003 | 0.003 | 3.1171
|  |  | 240.003 |  | 240.003 | 0.003 | 5.7989
|  |  | 270.003 |  | 270.003 | 0.003 | 6.0006
|  |  | 300.005 |  | 300.005 | 0.005 | 6.2328
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying

## Result (2)
Result of Calibration
|  | Certificate No. |  | : | SPR16070004-1 |  |  |  |  | Page : 4 of 4
|  | 3. Function Measurement : Measuring Z - Axis Measure | [5 empty] | Unit : | mm
|  |  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  |  | 0 |  | 0 | 0 | 2.0008
|  |  | 30 |  | 30.001 | 0.001 | 2.0401
|  |  | 60 |  | 60.001 | 0.001 | 2.1536
|  |  | 90 |  | 90.001 | 0.001 | 2.3306
|  |  | 120 |  | 120.001 | 0.001 | 2.5578
|  |  | 150 |  | 150 | 0 | 2.8233
|  |  | 180 |  | 180.003 | 0.003 | 3.1171
|  |  | 210 |  | 210.003 | 0.003 | 5.7989
|  |  | 240 |  | 240.003 | 0.003 | 6.0006
|  |  | 270 |  | 270.003 | 0.003 | 6.2328
|  |  | 300 |  | 300.005 | 0.005 | 6.4812
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying

## Uncertainty Budget
Uncertainty Budget of Measuring Microscope
X
Nominal Value | Repeatability |  | Uncert of Glass Scale |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.002 | 0.001 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0010004 | 0 | ∞ | 2 | 2.0008
30 | 0 | 0 | 0.002 | 0.001 | 0.000345 | 0.00019919 | 5e-05 | 2.8868e-05 | 0.0010201 | 0 | ∞ | 2 | 2.0401
60 | 0 | 0 | 0.002 | 0.001 | 0.00069 | 0.00039837 | 5e-05 | 2.8868e-05 | 0.0010768 | 0 | ∞ | 2 | 2.1536
90 | 0 | 0 | 0.002 | 0.001 | 0.001035 | 0.00059756 | 5e-05 | 2.8868e-05 | 0.0011653 | 0 | ∞ | 2 | 2.3306
120 | 0 | 0 | 0.002 | 0.001 | 0.00138 | 0.00079674 | 5e-05 | 2.8868e-05 | 0.0012789 | 0 | ∞ | 2 | 2.5578
150 | 0 | 0 | 0.002 | 0.001 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0014116 | 0 | ∞ | 2 | 2.8233
180 | 0 | 0 | 0.002 | 0.001 | 0.00207 | 0.0011951 | 5e-05 | 2.8868e-05 | 0.0015586 | 0 | ∞ | 2 | 3.1171
210 | 0 | 0 | 0.005084 | 0.002542 | 0.002415 | 0.0013943 | 5e-05 | 2.8868e-05 | 0.0028994 | 0 | ∞ | 2 | 5.7989
240 | 0 | 0 | 0.005098 | 0.002549 | 0.00276 | 0.0015935 | 5e-05 | 2.8868e-05 | 0.0030062 | 0 | ∞ | 2 | 6.0125
270 | 0 | 0 | 0.005098 | 0.002549 | 0.003105 | 0.0017927 | 5e-05 | 2.8868e-05 | 0.0031164 | 0 | ∞ | 2 | 6.2328
300 | 0 | 0 | 0.005112 | 0.002556 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0032406 | 0 | ∞ | 2 | 6.4812
Y
Nominal Value | Repeatability |  | Uncert of Glass Scale |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.002 | 0.001 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0010004 | 0 | ∞ | 2 | 2.0008
30 | 0 | 0 | 0.002 | 0.001 | 0.000345 | 0.00019919 | 5e-05 | 2.8868e-05 | 0.0010201 | 0 | ∞ | 2 | 2.0401
60 | 0 | 0 | 0.002 | 0.001 | 0.00069 | 0.00039837 | 5e-05 | 2.8868e-05 | 0.0010768 | 0 | ∞ | 2 | 2.1536
90 | 0 | 0 | 0.002 | 0.001 | 0.001035 | 0.00059756 | 5e-05 | 2.8868e-05 | 0.0011653 | 0 | ∞ | 2 | 2.3306
120 | 0 | 0 | 0.002 | 0.001 | 0.00138 | 0.00079674 | 5e-05 | 2.8868e-05 | 0.0012789 | 0 | ∞ | 2 | 2.5578
150 | 0 | 0 | 0.002 | 0.001 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0014116 | 0 | ∞ | 2 | 2.8233
180 | 0 | 0 | 0.002 | 0.001 | 0.00207 | 0.0011951 | 5e-05 | 2.8868e-05 | 0.0015586 | 0 | ∞ | 2 | 3.1171
210 | 0 | 0 | 0.005084 | 0.002542 | 0.002415 | 0.0013943 | 5e-05 | 2.8868e-05 | 0.0028994 | 0 | ∞ | 2 | 5.7989
240 | 0 | 0 | 0.005084 | 0.002542 | 0.00276 | 0.0015935 | 5e-05 | 2.8868e-05 | 0.0030003 | 0 | ∞ | 2 | 6.0006
270 | 0 | 0 | 0.005098 | 0.002549 | 0.003105 | 0.0017927 | 5e-05 | 2.8868e-05 | 0.0031164 | 0 | ∞ | 2 | 6.2328
300 | 0 | 0 | 0.005112 | 0.002556 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0032406 | 0 | ∞ | 2 | 6.4812
Z
Nominal Value | Repeatability |  | Uncert of Glass Scale |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.002 | 0.001 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0010004 | 0 | ∞ | 2 | 2.0008
30 | 0 | 0 | 0.002 | 0.001 | 0.000345 | 0.00019919 | 5e-05 | 2.8868e-05 | 0.0010201 | 0 | ∞ | 2 | 2.0401
60 | 0 | 0 | 0.002 | 0.001 | 0.00069 | 0.00039837 | 5e-05 | 2.8868e-05 | 0.0010768 | 0 | ∞ | 2 | 2.1536
90 | 0 | 0 | 0.002 | 0.001 | 0.001035 | 0.00059756 | 5e-05 | 2.8868e-05 | 0.0011653 | 0 | ∞ | 2 | 2.3306
120 | 0 | 0 | 0.002 | 0.001 | 0.00138 | 0.00079674 | 5e-05 | 2.8868e-05 | 0.0012789 | 0 | ∞ | 2 | 2.5578
150 | 0 | 0 | 0.002 | 0.001 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0014116 | 0 | ∞ | 2 | 2.8233
180 | 0 | 0 | 0.002 | 0.001 | 0.00207 | 0.0011951 | 5e-05 | 2.8868e-05 | 0.0015586 | 0 | ∞ | 2 | 3.1171
210 | 0 | 0 | 0.005084 | 0.002542 | 0.002415 | 0.0013943 | 5e-05 | 2.8868e-05 | 0.0028994 | 0 | ∞ | 2 | 5.7989
240 | 0 | 0 | 0.005084 | 0.002542 | 0.00276 | 0.0015935 | 5e-05 | 2.8868e-05 | 0.0030003 | 0 | ∞ | 2 | 6.0006
270 | 0 | 0 | 0.005098 | 0.002549 | 0.003105 | 0.0017927 | 5e-05 | 2.8868e-05 | 0.0031164 | 0 | ∞ | 2 | 6.2328
300 | 0 | 0 | 0.005112 | 0.002556 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0032406 | 0 | ∞ | 2 | 6.4812
Note on F4: Certificate of Calibration STD Glass Scale
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on F20: Certificate of Calibration STD Glass Scale
Note on H20: (Normal Value * 11.5*10^-6 * 1)
Note on E22: Divisor = 1
Note on G22: Divisor = 2
Note on I22: Divisor = √3
Note on K22: Divisor = √3
Note on F36: Certificate of Calibration STD Glass Scale
Note on H36: (Normal Value * 11.5*10^-6 * 1)
Note on E38: Divisor = 1
Note on G38: Divisor = 2
Note on I38: Divisor = √3
Note on K38: Divisor = √3

## Uncert of STD
Certificate of Calibration (Std. Scale) |  |  |  | Certificate of Calibration (Std. Scale) |  |  |  | Certificate of Calibration (Std. Scale)
SP-SD-020 |  |  |  | SP-SD-021 |  |  |  | SP-SD-022
Due Date | 2016-05-15 |  |  | Due Date | 2016-05-15 |  |  | Due Date | 2016-05-15
0 | µm | 0.002 | mm | 0 | µm | 0.002 | mm | 40 | µm | 0.005014 | mm
1 | µm | 0.002 | mm | 5 | µm | 0.002 | mm | 80 | µm | 0.005028 | mm
3 | µm | 0.002 | mm | 10 | µm | 0.002 | mm | 120 | µm | 0.005042 | mm
5 | µm | 0.002 | mm | 15 | µm | 0.002 | mm | 160 | µm | 0.005056 | mm
10 | µm | 0.002 | mm | 20 | µm | 0.002 | mm | 200 | µm | 0.00507 | mm
15 | µm | 0.002 | mm | 30 | µm | 0.002 | mm | 240 | µm | 0.005084 | mm
20 | µm | 0.002 | mm | 50 | µm | 0.002 | mm | 280 | µm | 0.005098 | mm
25 | µm | 0.002 | mm | 80 | µm | 0.002 | mm | 320 | µm | 0.005112 | mm
30 | µm | 0.002 | mm | 100 | µm | 0.002 | mm | 360 | µm | 0.005126 | mm
35 | µm | 0.002 | mm | 120 | µm | 0.002 | mm | 400 | µm | 0.00514 | mm
40 | µm | 0.002 | mm | 150 | µm | 0.002 | mm | 410 | µm | 0.0051435 | mm
45 | µm | 0.002 | mm | 200 | µm | 0.002 | mm
50 | µm | 0.002 | mm
[12 empty] | 50 | 0.09 | µm | 9e-05 | mm
[12 empty] | 75 | 0.1 | µm | 0.0001 | mm
[12 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on B5: Grade 0
Note on H5: Grade 0
Note on N5: Grade 0
Note on D6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on P6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง