## Data Record
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  | SPR15120023-1 | [7 empty] | Page | 1 | of | 1
[9 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  | 2015-12-12
Equipment Name | [8 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
High Master | [8 empty] | Location |  | In Lab
Customer Name : |  |  |  | LG
Equipment Name : |  |  |  | High Master |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  | 123
Serial No. : |  |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 | mm
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : | [10 empty] | Due Date :
Referance Standard : | [10 empty] | Due Date :
1. Block Pitch Accuracy and Paralellism
Nominal Lengh |  | Position Point |  | STD Reading | [7 empty] | Average |  |  |  | Deviation
|  |  |  | X1 |  | X2 |  |  | X3 | X4
20 |  | C |  | 0 |  | 0 |  |  | 0 | 0 |  | 0 |  |  |  | 0.00013038
|  | 1 |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001
|  | 2 |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001
|  | 3 |  | -0.0001 |  | -0.0001 |  |  | -0.0001 | -0.0001 |  | -0.0001
|  | 4 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
Nominal Lengh |  | UUC Error | [7 empty] | STD Reading |  |  | Error |  |  |  | Repeatability
|  | X1 |  | X2 |  | X3 |  |  | X4
10 |  | 0.0001 |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 10.0001 |  |  | 0.0001 |  |  |  | 0
20 |  | -0.0002 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | 19.9998 |  |  | -0.0002 |  |  |  | 0
30 |  | 0 |  | 0 |  | 0 |  |  | 0 | 30 |  |  | 0 |  |  |  | 0
50 |  | 0 |  | 0 |  | 0 |  |  | 0 | 50 |  |  | 0 |  |  |  | 0
75 |  | 0 |  | 0 |  | 0 |  |  | 0 | 75 |  |  | 0 |  |  |  | 0
100 |  | 0 |  | 0 |  | 0 |  |  | 0 | 100 |  |  | 0 |  |  |  | 0
125 |  | 0 |  | 0 |  | 0 |  |  | 0 | 125 |  |  | 0 |  |  |  | 0
150 |  | 0 |  | 0 |  | 0 |  |  | 0 | 150 |  |  | 0 |  |  |  | 0
200 |  | 0 |  | 0 |  | 0 |  |  | 0 | 200 |  |  | 0 |  |  |  | 0
250 |  | 0 |  | 0 |  | 0 |  |  | 0 | 250 |  |  | 0 |  |  |  | 0
300 |  | 0 |  | 0 |  | 0 |  |  | 0 | 300 |  |  | 0 |  |  |  | 0
400 |  | 0 |  | 0 |  | 0 |  |  | 0 | 400 |  |  | 0 |  |  |  | 0
500 |  | 0 |  | 0 |  | 0 |  |  | 0 | 500 |  |  | 0 |  |  |  | 0
Calibrate By :
| 11 | Ms. Arunkamon Raramanus
1
4
2
3
C

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120023-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | : | LG
| Equipment Name |  |  |  | : | High Master
| Manufacturer |  |  |  | : | Mittutoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | sdfdthj
| ID. Number |  |  |  | : | 123
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  |  | : | 2016-05-05
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-05-06
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2017-05-06
| Calibration Procedure |  |  |  |  | : | SP-CPD-04-33
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-05-07 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120023-1 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  |  | SPR15120023-1 | [5 empty] | Page : 3 of 3
|  |  | Nominal Value |  | STD Reading | Error | Uncertainty ( ± ) µm
|  |  | 10 |  | 10.0001 | 0.0001 | 2.3147
|  |  | 20 |  | 19.9998 | -0.0002 | 2.3264
|  |  | 30 |  | 30 | 0 | 2.3468
|  |  | 50 |  | 50 | 0 | 2.4053
|  |  | 75 |  | 75 | 0 | 2.5176
|  |  | 100 |  | 100 | 0 | 2.6673
|  |  | 125 |  | 125 | 0 | 2.8755
|  |  | 150 |  | 150 | 0 | 3.0863
|  |  | 200 |  | 200 | 0 | 3.5627
|  |  | 250 |  | 250 | 0 | 4.0932
|  |  | 300 |  | 300 | 0 | 4.6595
|  |  | 400 |  | 400 | 0 | 5.8602
|  |  | 500 |  | 500 | 0 | 7.1155
|  |  | Block Pitch Accuracy and Paralellism Devation |  |  |  |  | 0.00013038 | mm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Height Master
Nominal Value | Repeatability |  | Uncertainty of Lenier |  | Uncertainty of GB |  | Contact Deformation |  | Temperature Effect |  | Resolution of STD |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
10 | 0 | 0 | 0.002 | 0.0011547 | 6e-05 | 3e-05 | 0 | 0 | 0.000115 | 6.6395e-05 | 5e-05 | 2.8868e-05 | 0.0011574 | 0 | ∞ | 2 | 2.3147
20 | 0 | 0 | 0.002 | 0.0011547 | 7e-05 | 3.5e-05 | 0 | 0 | 0.00023 | 0.00013279 | 5e-05 | 2.8868e-05 | 0.0011632 | 0 | ∞ | 2 | 2.3264
30 | 0 | 0 | 0.002 | 0.0011547 | 0.00011 | 5.5e-05 | 0 | 0 | 0.000345 | 0.00019919 | 5e-05 | 2.8868e-05 | 0.0011734 | 0 | ∞ | 2 | 2.3468
50 | 0 | 0 | 0.002 | 0.0011547 | 9e-05 | 4.5e-05 | 0 | 0 | 0.000575 | 0.00033198 | 5e-05 | 2.8868e-05 | 0.0012027 | 0 | ∞ | 2 | 2.4053
75 | 0 | 0 | 0.002 | 0.0011547 | 0.0001 | 5e-05 | 0 | 0 | 0.0008625 | 0.00049796 | 5e-05 | 2.8868e-05 | 0.0012588 | 0 | ∞ | 2 | 2.5176
100 | 0 | 0 | 0.002 | 0.0011547 | 0.00012 | 6e-05 | 0 | 0 | 0.00115 | 0.00066395 | 5e-05 | 2.8868e-05 | 0.0013336 | 0 | ∞ | 2 | 2.6673
125 | 0 | 0 | 0.002 | 0.0011547 | 0.00042 | 0.00021 | 0 | 0 | 0.0014375 | 0.00082994 | 5e-05 | 2.8868e-05 | 0.0014377 | 0 | ∞ | 2 | 2.8755
150 | 0 | 0 | 0.002 | 0.0011547 | 0.00047 | 0.000235 | 0 | 0 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0015431 | 0 | ∞ | 2 | 3.0863
200 | 0 | 0 | 0.002 | 0.0011547 | 0.00055 | 0.000275 | 0 | 0 | 0.0023 | 0.0013279 | 5e-05 | 2.8868e-05 | 0.0017813 | 0 | ∞ | 2 | 3.5627
250 | 0 | 0 | 0.002 | 0.0011547 | 0.00063 | 0.000315 | 0 | 0 | 0.002875 | 0.0016599 | 5e-05 | 2.8868e-05 | 0.0020466 | 0 | ∞ | 2 | 4.0932
300 | 0 | 0 | 0.002 | 0.0011547 | 0.00071 | 0.000355 | 0 | 0 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0023297 | 0 | ∞ | 2 | 4.6595
400 | 0 | 0 | 0.002 | 0.0011547 | 0.00089 | 0.000445 | 0 | 0 | 0.0046 | 0.0026558 | 5e-05 | 2.8868e-05 | 0.0029301 | 0 | ∞ | 2 | 5.8602
500 | 0 | 0 | 0.002 | 0.0011547 | 0.0011 | 0.00055 | 0 | 0 | 0.00575 | 0.0033198 | 5e-05 | 2.8868e-05 | 0.0035577 | 0 | ∞ | 2 | 7.1155
Contact Deformation
[7 empty] | STD | UUC
[7 empty] | Steel | Steel | 0 | µm
[7 empty] | Steel | Ceramic | 0.0004 | µm
[7 empty] | Steel | Tungsten | 0.0137 | µm
Note on F4: Certificate of Calubration Linear Height Gauge
Note on H4: Certificate of Calubration Gauge Block
Note on L4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3

## Uncert of STD
Certificate of Calibration (Linear) |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
SP-SD-006 |  |  |  |  | SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
Due Date |  | 2015-06-26 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09
0 | 2 | µm | 0.002 | mm | 2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
20 | 2 | µm | 0.002 | mm | 5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
50 | 2 | µm | 0.002 | mm | 7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
100 | 2 | µm | 0.002 | mm | 10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
150 | 2 | µm | 0.002 | mm | 12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
200 | 2 | µm | 0.002 | mm | 15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
250 | 2 | µm | 0.002 | mm | 17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
300 | 2 | µm | 0.002 | mm | 20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
400 | 2 | µm | 0.002 | mm | 22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
500 | 2 | µm | 0.002 | mm | 25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
600 | 2 | µm | 0.002 | mm | [5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
700 | 2 | µm | 0.002 | mm | Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
[5 empty] | Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[15 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[15 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[15 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[15 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[15 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[15 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[15 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[15 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[15 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[15 empty] | 14 | 0.07 | µm | 7e-05 | mm
[15 empty] | 15 | 0.07 | µm | 7e-05 | mm
[15 empty] | 16 | 0.07 | µm | 7e-05 | mm
[15 empty] | 17 | 0.07 | µm | 7e-05 | mm
[15 empty] | 18 | 0.07 | µm | 7e-05 | mm
[15 empty] | 19 | 0.07 | µm | 7e-05 | mm
[15 empty] | 20 | 0.07 | µm | 7e-05 | mm
[15 empty] | 21 | 0.07 | µm | 7e-05 | mm
[15 empty] | 22 | 0.07 | µm | 7e-05 | mm
[15 empty] | 23 | 0.07 | µm | 7e-05 | mm
[15 empty] | 24 | 0.07 | µm | 7e-05 | mm
[15 empty] | 25 | 0.07 | µm | 7e-05 | mm
[15 empty] | 50 | 0.09 | µm | 9e-05 | mm
[15 empty] | 75 | 0.1 | µm | 0.0001 | mm
[15 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on H3: Grade 0
Note on N3: Grade 0
Note on T3: Grade 0
Note on Z3: Grade 0
Note on AF3: Grade 0
Note on D4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on P4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on V4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AB4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AH4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on C5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Linear Hight Gauge
Note on E5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on K5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AA5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AC5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AG5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AI5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน