## Data Record
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR15120045-1 | [6 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Bevel Protractor | [7 empty] | Location |  |  |  |  | In Lab |  |  |  | On Site
Customer Name : |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  | Bevel Protractor | [8 empty] | Manufacturer : |  |  |  | INZER
Model : | 123 | [6 empty] | Serial No. : |  |  | 456 | [5 empty] | ID No : |  | 789
Range : | Scale |  | 0 | to | 300 |  | mm |  | Angle |  | 0 | to | 300 |  | Deg |  | Resolution : |  |  | 0.001 |  | Deg
Overall Inspection | [5 empty] | Good |  |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Referance Standard : | [13 empty] | Due Date :
Function Scale measurement
Nominal Value. |  | UUC Reading | [11 empty] | Average |  |  | Error |  |  | Repeatability
|  | X1 |  |  | X2 |  |  | X3 |  |  | X4
0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
30 |  | 30 |  |  | 30 |  |  | 30 |  |  | 30 |  |  | 30 |  |  | 0 |  |  | 0
60 |  | 60 |  |  | 60 |  |  | 60 |  |  | 60 |  |  | 60 |  |  | 0 |  |  | 0
90 |  | 90 |  |  | 90 |  |  | 90 |  |  | 90 |  |  | 90 |  |  | 0 |  |  | 0
120 |  | 120 |  |  | 120 |  |  | 120 |  |  | 120 |  |  | 120 |  |  | 0 |  |  | 0
150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  |  | 0
180 |  | 180 |  |  | 180 |  |  | 180 |  |  | 180 |  |  | 180 |  |  | 0 |  |  | 0
210 |  | 210 |  |  | 210 |  |  | 210 |  |  | 210 |  |  | 210 |  |  | 0 |  |  | 0
240 |  | 240 |  |  | 240 |  |  | 240 |  |  | 240 |  |  | 240 |  |  | 0 |  |  | 0
270 |  | 270 |  |  | 270 |  |  | 270 |  |  | 270 |  |  | 270 |  |  | 0 |  |  | 0
300 |  | 300 |  |  | 300 |  |  | 300 |  |  | 300 |  |  | 300 |  |  | 0 |  |  | 0
Function angle measurement
| Nominal Value. |  |  | UUC Reading | [7 empty] | Average |  |  | Error |  |  | Repeatability
|  |  |  | X1 |  | X2 |  | X3 |  | X4
| 1 |  | ° | 1 |  | 1 |  | 1 |  | 1 |  | 1 |  |  | 0 |  |  | 0
| 2 |  | ° | 2 |  | 2 |  | 2 |  | 2 |  | 2 |  |  | 0 |  |  | 0
| 3 |  | ° | 3 |  | 3 |  | 3 |  | 3 |  | 3 |  |  | 0 |  |  | 0
| 4 |  | ° | 4 |  | 4 |  | 4 |  | 4 |  | 4 |  |  | 0 |  |  | 0
| 5 |  | ° | 5 |  | 5 |  | 5 |  | 5 |  | 5 |  |  | 0 |  |  | 0
| 10 |  | ° | 10 |  | 10 |  | 10 |  | 10 |  | 10 |  |  | 0 |  |  | 0
| 15 |  | ° | 15 |  | 15 |  | 15 |  | 15 |  | 15 |  |  | 0 |  |  | 0
| 20 |  | ° | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  |  | 0 |  |  | 0
| 25 |  | ° | 25 |  | 25 |  | 25 |  | 25 |  | 25 |  |  | 0 |  |  | 0
| 30 |  | ° | 30 |  | 30 |  | 30 |  | 30 |  | 30 |  |  | 0 |  |  | 0
Calibrated By : |  |  | Mr.

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR15120045-1 |  |  |  |  | Page : 1 of 4
| Customer |  |  | : | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
| Equipment Name |  |  | : | Bevel Protractor
| Manufacturer |  |  | : | INZER
| Model |  |  | : | 123
| Serial Number |  |  | : | 456
| ID. Number |  |  | : | 789
| Received Date |  |  | : | 2016-01-01
| Calibration Date |  |  | : | 2016-01-01
| Recommended Due Date |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  | : | 20 °C | ± 1 °C
| Relative Humidity |  |  | : | 50% | ± 15 %
| Location of Calibration |  |  | : | In-Lab
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. standards used to perform this
calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. the result reported herein apply only to the calibration of the item described above.
| all calibrations are performed manufacture's specifications, full, without the expressed written consent of | [9 empty] | 1 | Mr.Sombut Srikampa
| SP Metrology System (Thailand). | [9 empty] | 3 | Mr. Natthaphol Boonmee
[11 empty] | 8 | Mr. Vichan Ananta
[11 empty] | 9 | Mr.Kittikorn Kingmali
| Date of Issue : |  | 2016-01-02 |  |  |  | Approved by : |  |  |  | 10 | Mr.Chainarong Matchayamat
[8 empty] | 3 | ( Mr. Natthaphol Boonmee ) |  | 11 | Ms. Arunkamon Raramanus
| Calibrated by : |  | Ms. Arunkamon Raramanus | [5 empty] | Authorized Signatory
Note on O39: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120045-1 |  |  | Page : 2 of 4
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| 3D Vision Measuring Machine |  |  | D3020-T | 110021 | MTO 150453-1 | 2016-06-26
| Angle block set |  |  | AB-130 | 50136 | MC 1506930 | 2016-05-25
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  |  | SPR15120045-1 | [7 empty] | Page : 3 of 3
|  | Function Scale measurement | [8 empty] | Unit : |  | mm
|  |  | Nominal Value |  |  | UUC Reading |  | Error |  | Uncertainty ( ± ) µm
|  |  | 0 |  |  | 0 |  | 0 |  | 3.0006
|  |  | 30 |  |  | 30 |  | 0 |  | 3.0269
|  |  | 60 |  |  | 60 |  | 0 |  | 3.1045
|  |  | 90 |  |  | 90 |  | 0 |  | 3.2298
|  |  | 120 |  |  | 120 |  | 0 |  | 3.3974
|  |  | 150 |  |  | 150 |  | 0 |  | 3.6015
|  |  | 180 |  |  | 180 |  | 0 |  | 3.8362
|  |  | 210 |  |  | 210 |  | 0 |  | 4.0963
|  |  | 240 |  |  | 240 |  | 0 |  | 4.3772
|  |  | 270 |  |  | 270 |  | 0 |  | 4.6753
|  |  | 300 |  |  | 300 |  | 0 |  | 4.9873
|  | Function angle measurement | [7 empty] | Unit : |  | Deegree
|  |  | Nominal Value |  |  | UUC Reading |  | Error |  | Uncertainty ( ± ) "
|  |  | 1 | ° |  | 1 | ° | 0 | ° | 12.0771
|  |  | 2 | ° |  | 2 | ° | 0 | ° | 12.0771
|  |  | 3 | ° |  | 3 | ° | 0 | ° | 12.0771
|  |  | 4 | ° |  | 4 | ° | 0 | ° | 12.0771
|  |  | 5 | ° |  | 5 | ° | 0 | ° | 12.0771
|  |  | 10 | ° |  | 10 | ° | 0 | ° | 12.0771
|  |  | 15 | ° |  | 15 | ° | 0 | ° | 12.0771
|  |  | 20 | ° |  | 20 | ° | 0 | ° | 12.0218
|  |  | 25 | ° |  | 25 | ° | 0 | ° | 12.0218
|  |  | 30 | ° |  | 30 | ° | 0 | ° | 12.0218
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget(Angle)
Uncertainty Budget of Bevel Protractor
Nominal Value | Repeatability |  | Uncert of Angle Block |  | flatness of Angle Blocks |  | Temperature Effect |  | Mechanical Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
( o ) | ( o ) |  | ( o ) |  | ( o ) |  | ( o ) |  | ( o ) |  | ( o ) | [5 empty] | ( " )
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
1 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
2 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
3 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
4 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
5 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
10 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
15 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0386 | 0 | ∞ | 2 | 12.0771
20 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0109 | 0 | ∞ | 2 | 12.0218
25 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0109 | 0 | ∞ | 2 | 12.0218
30 | 0 | 0 | 12 | 6 | 0.576 | 0.33255 | 0.144 | 0.083138 | 1.0002 | 0.57747 | 0.20004 | 0.11549 | 6.0109 | 0 | ∞ | 2 | 12.0218
Temp Effect
Temperature Variation from height of Base Angle Blocks = 6 x (11.5x10-6) x 1
= 0.000069 mm
Temperature Variation of Angle Blocks = tan -1 ( 0.000069/100)
= | 4e-05 | Deg
= | 0.144 | Second
Flatness of Angle Blocks
Angle from flatness of Angle Blocks = tan-1
= | tan-1 (0.00029/100)
= | 0.00016 | Deg
= | 0.576 | Second
Resolution
0.05' | 4.9998 | Second
0.01' | 1.0002 | Second
Mechanical Effect
0.01' | 1.0002 | Second
Note on E4: Certificate of Calibration Angle Block
Note on I4: (Normal Value * 11.5*10^-6 * 1)
Note on D6: Divisor = 1
Note on F6: Divisor = 2
Note on H6: Divisor = 2
Note on J6: Divisor = √3
Note on L6: Divisor = √3
Note on N6: Divisor = √3

## Uncertainty Budget(Length)
Uncertainty Budget of Bevel Protractor
Nominal Value | Repeatability |  | Uncert of 3D Vision |  | Temperature Effect |  | Resolution of 3D Vision |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.003 | 0.0015 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0015003 | 0 | 45.5963 | 2.00 | 3.0006
30 | 0 | 0 | 0.003 | 0.0015 | 0.000345 | 0.00019919 | 5e-05 | 2.8868e-05 | 0.0015134 | 0 | 47.2179 | 2.00 | 3.0269
60 | 0 | 0 | 0.003 | 0.0015 | 0.00069 | 0.00039837 | 5e-05 | 2.8868e-05 | 0.0015523 | 0 | 52.2527 | 2.00 | 3.1045
90 | 0 | 0 | 0.003 | 0.0015 | 0.001035 | 0.00059756 | 5e-05 | 2.8868e-05 | 0.0016149 | 0 | 61.2107 | 2.00 | 3.2298
120 | 0 | 0 | 0.003 | 0.0015 | 0.00138 | 0.00079674 | 5e-05 | 2.8868e-05 | 0.0016987 | 0 | 74.9419 | 2.00 | 3.3974
150 | 0 | 0 | 0.003 | 0.0015 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0018008 | 0 | 94.6364 | 2.00 | 3.6015
180 | 0 | 0 | 0.003 | 0.0015 | 0.00207 | 0.0011951 | 5e-05 | 2.8868e-05 | 0.0019181 | 0 | 121.8242 | 2.00 | 3.8362
210 | 0 | 0 | 0.003 | 0.0015 | 0.002415 | 0.0013943 | 5e-05 | 2.8868e-05 | 0.0020481 | 0 | 158.3753 | 2.00 | 4.0963
240 | 0 | 0 | 0.003 | 0.0015 | 0.00276 | 0.0015935 | 5e-05 | 2.8868e-05 | 0.0021886 | 0 | 206.4998 | 2.00 | 4.3772
270 | 0 | 0 | 0.003 | 0.0015 | 0.003105 | 0.0017927 | 5e-05 | 2.8868e-05 | 0.0023376 | 0 | 268.7477 | 2.00 | 4.6753
300 | 0 | 0 | 0.003 | 0.0015 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0024937 | 0 | 348.009 | 2.00 | 4.9873
Note on F4: Certificate of Calibration 3D Vision
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3

## Cert STD
Certificate of Calibration (Angle Block) |  |  | Certificate of Calibration (3D) |  |  |  |  | Certificate of Calibration (3D)
SP-SD-025 |  |  | SP-SD-005 |  |  |  |  | SP-SD-005
1/4 | 12 | " | X - Axis |  |  |  |  | Y - Axis
1/2 | 12 | " | 0 | 3 | µm | 0.003 | mm | 0 | 3 | µm | 0.003 | mm
1 | 12 | " | 5 | 3 | µm | 0.003 | mm | 5 | 3 | µm | 0.003 | mm
2 | 12 | " | 10 | 3 | µm | 0.003 | mm | 10 | 3 | µm | 0.003 | mm
3 | 12 | " | 30 | 3 | µm | 0.003 | mm | 30 | 3 | µm | 0.003 | mm
4 | 12 | " | 50 | 3 | µm | 0.003 | mm | 50 | 3 | µm | 0.003 | mm
5 | 12 | " | 80 | 3 | µm | 0.003 | mm | 80 | 3 | µm | 0.003 | mm
10 | 12 | " | 100 | 3 | µm | 0.003 | mm | 100 | 3 | µm | 0.003 | mm
15 | 12 | " | 150 | 3 | µm | 0.003 | mm | 150 | 3 | µm | 0.003 | mm
20 | 12 | " | 200 | 3 | µm | 0.003 | mm | 200 | 3 | µm | 0.003 | mm
25 | 12 | " | 250 | 3 | µm | 0.003 | mm
30 | 12 | " | 300 | 3 | µm | 0.003 | mm
Note on C4: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Angle
Note on D4: Unit ° degree, ' minute, " second
Note on E4: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on K5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on Q5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน