## Data Record
SP METROLOGY SYSTEM THAILAND | [5 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [6 empty] | Page |  | 1 | of | 1
[6 empty] | Receive Date : |  |  |  | 2015-01-11 |  |  |  | Calibration Date : |  |  |  | 2015-01-12
Equipment Name | [5 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Pitch Gauge | [5 empty] | Location |  |  |  |  | In Lab |  |  |  |  | On Site
Customer Name : |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  | Pitch Gauge | [6 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : | 123 |  |  |  | Serial No. : |  |  | 456 | [6 empty] | ID No : |  | 789
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [11 empty] | Due Date :
Nominal Value |  | Positions |  | Standard Reading | [6 empty] | Average |  |  |  | Error |  |  | Repeatability
|  |  |  | 1 | 2 |  | 3 |  | 4
2 |  | P1 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988 |  | 1.9993 |  |  |  | 0.00066667 |  |  | 0
|  | P2 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P3 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P4 |  | 2 | 2 |  | 2 |  | 2
|  | P5 |  | 1.9999 | 1.9999 |  | 1.9999 |  | 1.9999
|  | P6 |  | 1.9997 | 1.9997 |  | 1.9997 |  | 1.9997
2 |  | P1 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988 |  | 1.9993 |  |  |  | 0.00066667 |  |  | 0
|  | P2 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P3 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P4 |  | 2 | 2 |  | 2 |  | 2
|  | P5 |  | 1.9999 | 1.9999 |  | 1.9999 |  | 1.9999
|  | P6 |  | 1.9997 | 1.9997 |  | 1.9997 |  | 1.9997
2 |  | P1 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988 |  | 1.9993 |  |  |  | 0.00066667 |  |  | 0
|  | P2 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P3 |  | 1.9988 | 1.9988 |  | 1.9988 |  | 1.9988
|  | P4 |  | 2 | 2 |  | 2 |  | 2
|  | P5 |  | 1.9999 | 1.9999 |  | 1.9999 |  | 1.9999
|  | P6 |  | 1.9997 | 1.9997 |  | 1.9997 |  | 1.9997
Calibration by

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR15120012-1 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
| Equipment Name |  |  | : | Pitch Gauge
| Manufacturer |  |  | : | Mitutoyo
| Model |  |  | : | 123
| Serial Number |  |  | : | 456
| ID. Number |  |  | : | 789
| Received Date |  |  | : | 2015-01-11
| Calibration Date |  |  | : | 2015-01-12
| Recommended Due Date |  |  | : | 2016-01-12
Environmental Conditions
| Ambient Temperature |  |  | : | 23 °C | ± 2 °C
| Relative Humidity |  |  | : | 50% | ± 15 %
| Location of Calibration |  |  | : | In-Lab
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. standards used to perform this
calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. the result reported herein apply only to the calibration of the item described above.
| all calibrations are performed manufacture's specifications, full, without the expressed written consent of | [9 empty] | 1 | Mr.Sombut Srikampa
| SP Metrology System (Thailand). | [9 empty] | 3 | Mr. Natthaphol Boonmee
[11 empty] | 8 | Mr. Vichan Ananta
[11 empty] | 9 | Mr.Kittikorn Kingmali
| Date of Issue : |  | 2015-01-13 |  |  |  | Approved by : |  |  |  | 10 | Mr.Chainarong Matchayamat
[8 empty] | 3 | ( Mr. Natthaphol Boonmee ) |  | 11 | Ms. Arunkamon Raramanus
| Calibrated by : |  | Ms. Arunkamon Raramanus | [5 empty] | Authorized Signatory
Note on O39: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120012-1 |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| 3D Vision Measuring Machine |  |  | D3020-T | 110021 | MTO 150453-1 | 2016-06-26
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  | SPR15120012-1 |  |  |  |  | Page : 3 of 3
[8 empty] | Unit : | mm
|  |  | Nominal Value |  | STD Reading | Error | Uncertainty ( ± ) µm
|  |  | 2 |  | 1.9993 | 0.00066667 | 3.0007
|  |  | 2 |  | 1.9993 | 0.00066667 | 3.0007
|  |  | 2 |  | 1.9993 | 0.00066667 | 3.0007
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Pitch Gauge
Nominal Value | Repeatability |  | Uncert of 3D Vision |  | Temperature Effect |  | Resolution of 3D Vision |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
2 | 0 | 0 | 0.003 | 0.0015 | 2.3e-05 | 1.3279e-05 | 5e-05 | 2.8868e-05 | 0.0015003 | 0 | ∞ | 2 | 3.0007
2 | 0 | 0 | 0.003 | 0.0015 | 2.3e-05 | 1.3279e-05 | 5e-05 | 2.8868e-05 | 0.0015003 | 0 | ∞ | 2 | 3.0007
2 | 0 | 0 | 0.003 | 0.0015 | 2.3e-05 | 1.3279e-05 | 5e-05 | 2.8868e-05 | 0.0015003 | 0 | ∞ | 2 | 3.0007
Note on F4: Certificate of Calibration 3D Vision
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3

## Cert of STD
Certificate of Calibration (3D) |  |  |  |  | Certificate of Calibration (3D)
SP-SD-005 |  |  |  |  | SP-SD-005
Due Date |  | 2015-06-26 |  |  | Due Date |  | 2015-06-26
X - Axis |  |  |  |  | Y - Axis
0 | 3 | µm | 0.003 | mm | 0 | 3 | µm | 0.003 | mm
5 | 3 | µm | 0.003 | mm | 5 | 3 | µm | 0.003 | mm
10 | 3 | µm | 0.003 | mm | 10 | 3 | µm | 0.003 | mm
30 | 3 | µm | 0.003 | mm | 30 | 3 | µm | 0.003 | mm
50 | 3 | µm | 0.003 | mm | 50 | 3 | µm | 0.003 | mm
80 | 3 | µm | 0.003 | mm | 80 | 3 | µm | 0.003 | mm
100 | 3 | µm | 0.003 | mm | 100 | 3 | µm | 0.003 | mm
150 | 3 | µm | 0.003 | mm | 150 | 3 | µm | 0.003 | mm
200 | 3 | µm | 0.003 | mm | 200 | 3 | µm | 0.003 | mm
250 | 3 | µm | 0.003 | mm
300 | 3 | µm | 0.003 | mm
Note on B5: Grade 0
Note on H5: Grade 0
Note on D6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on C8: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on E8: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I8: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on K8: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน