## Data Record (pitch)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [10 empty] | Page |  | 1 | of | 1
[9 empty] | Receive Date : |  |  |  | 2015-12-17 |  |  |  |  | Calibration Date : |  |  | 2015-12-20
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Taper Thread Ring | [8 empty] | Location |  |  |  |  | In Lab |  |  |  |  | On Site
Customer Name : |  |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  |  | Taper Thread Ring | [6 empty] | Manufacturer : |  |  |  | OPI |  |  |  | Model : |  | ygi
Serial No. : |  |  | 123 |  |  |  |  | ID No :
Symbol of Taper Thread Ring Gauge | [13 empty] | GO | [7 empty] | NO | GO
Overall Inspection | [6 empty] | Good |  |  |  | Not Good
Referance Standard : | [11 empty] | Due Date :
1. Measurement Results
Description : |  |  |  |  | M | 10 |  | x | P | 0.8 |  |  | GPII |  |  |  | Pitch Diameter [E] |  |  |  |  | 9.3072 |  |  | mm
Thread angle: | [5 empty] | 60 | [10 empty] | 2-Ball Stylus (dD) |  |  |  |  | 1.35 |  |  | mm |  |  |  | 2 Ball Stylus of SP
| Slope: |  | 16 | [13 empty] | Best Dimeter (do) |  |  |  |  | 0.46188 |  |  | mm |  |  |  | 0.2 | 1.1
[17 empty] | d๐ = (P/2)(1/COS(α/2) | [11 empty] | 0.25 | 1.35
[29 empty] | 0.28 | 1.65
Nominal |  | Position |  | STD Reading | [11 empty] | Average |  |  | Repeat |  | Pitch Dia [E] |  |  | Deviation |  |  |  |  | 0.455 | 2.05
|  |  |  | X1 |  |  | X2 |  |  | X3 |  |  | X4 | [15 empty] | 0.53 | 2.55
X |  | 1/5T |  | 10 |  |  | 10 |  |  | 10 |  |  | 10 |  |  | 10 |  |  | 0.87228 |  | 9.3072 |  |  | 0 |  |  |  |  | 0.62 | 3.2
|  | 1/2T |  | 11 |  |  | 11 |  |  | 11 |  |  | 11 |  |  | 11 |  |  |  |  | 10.3072 |  |  | 1 |  |  |  |  | 0.725 | 4
|  | 4/5T |  | 12 |  |  | 12 |  |  | 12 |  |  | 12 |  |  | 12 |  |  |  |  | 11.3072 |  |  | 2 |  |  |  |  | 0.895
Y |  | 1/5T |  | 13 |  |  | 13 |  |  | 13 |  |  | 13 |  |  | 13 |  |  |  |  | 12.3072 |  |  | 3
|  | 1/2T |  | 14 |  |  | 14 |  |  | 14 |  |  | 14 |  |  | 14 |  |  |  |  | 13.3072 |  |  | 4
|  | 4/5T |  | 15 |  |  | 15 |  |  | 15 |  |  | 15 |  |  | 15 |  |  |  |  | 14.3072 |  |  | 5
Description : |  |  |  |  | M | 20 |  | x | P | 2.5 |  |  | GPII |  |  |  | Pitch Diameter [E] |  |  |  |  | 17.8349 |  |  | mm
Thread angle: | [5 empty] | 60 | [10 empty] | 2-Ball Stylus (dD) |  |  |  |  | 1.65 |  |  | mm
[17 empty] | Best Dimeter (do) |  |  |  |  | 1.4434 |  |  | mm
d๐ = (P/2)(1/COS(α/2)
Nominal |  | Position |  | STD Reading | [11 empty] | Average |  |  | Repeat |  | Pitch Dia [E] |  |  | Deviation
|  |  |  | X1 |  |  | X2 |  |  | X3 |  |  | X4
X |  | 1/5T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 3.6291e-15 |  | 15.7014 |  |  | -2.1335
|  | 1/2T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  |  |  | 15.7014 |  |  | -2.1335
|  | 4/5T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  |  |  | 15.7014 |  |  | -2.1335
Y |  | 1/5T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  |  |  | 15.7014 |  |  | -2.1335
|  | 1/2T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  |  |  | 15.7014 |  |  | -2.1335
|  | 4/5T |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  | 17.8665 |  |  |  |  | 15.7014 |  |  | -2.1335
Calibrated By |  |  |  | :

## Data Record (minor)
SP METROLOGY SYSTEM THAILAND | [6 empty] | Certificate No. : |  |  |  | SPR15120012-1 | [10 empty] | Page | 1 | of | 1
[7 empty] | Receive Date : |  |  |  | 2015-12-17 |  |  |  |  | Calibration Date : |  |  | 2015-12-20
Equipment Name | [6 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Taper Thread Ring Gauge | [6 empty] | Location |  |  |  |  | In Lab |  |  |  |  | On Site
Customer Name : |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  | Taper Thread Ring Gauge | [5 empty] | Manufacturer : |  |  |  | OPI |  |  |  | Model : |  | ygi
Serial No. : |  | 123 |  |  |  | ID No :
Symbol of Taper Thread Ring Gauge | [11 empty] | GO | [7 empty] | NO | GO
Overall Inspection |  |  |  |  | Good |  |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
[26 empty] | 0.28 | 1.65
Nominal | Position |  | STD Reading | [10 empty] | Average |  |  | Repeat |  | Deviation | [6 empty] | 0.455 | 2.05
|  |  | X1 |  | X2 |  |  | X3 |  |  | X4 | [14 empty] | 0.53 | 2.55
X | 1/5T |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 10 |  |  | 0.87228 |  | 1.7446 | [6 empty] | 0.62 | 3.2
| 1/2T |  | 11 |  | 11 |  |  | 11 |  |  | 11 |  |  | 11 | [11 empty] | 0.725 | 4
| 4/5T |  | 12 |  | 12 |  |  | 12 |  |  | 12 |  |  | 12 | [11 empty] | 0.895
Y | 1/5T |  | 13 |  | 13 |  |  | 13 |  |  | 13 |  |  | 13
| 1/2T |  | 14 |  | 14 |  |  | 14 |  |  | 14 |  |  | 14
| 4/5T |  | 15 |  | 15 |  |  | 15 |  |  | 15 |  |  | 15
Nominal | Position |  | STD Reading | [10 empty] | Average |  |  | Repeat |  | Deviation
|  |  | X1 |  | X2 |  |  | X3 |  |  | X4
X | 1/5T |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 10 |  |  | 0.87228 |  | 1.7446
| 1/2T |  | 11 |  | 11 |  |  | 11 |  |  | 11 |  |  | 11
| 4/5T |  | 12 |  | 12 |  |  | 12 |  |  | 12 |  |  | 12
Y | 1/5T |  | 13 |  | 13 |  |  | 13 |  |  | 13 |  |  | 13
| 1/2T |  | 14 |  | 14 |  |  | 14 |  |  | 14 |  |  | 14
| 4/5T |  | 15 |  | 15 |  |  | 15 |  |  | 15 |  |  | 15

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR15120012-1 |  |  |  |  | Page : 1 of 3
| Customer |  |  | :
| Equipment Name |  |  | : | Taper Thread Ring
| Manufacturer |  |  | : | OPI
| Model |  |  | : | 123
| Serial Number |  |  | : | 0
| ID. Number |  |  | : | 0
| Received Date |  |  | : | 2015-12-17
| Calibration Date |  |  | : | 2015-12-20
| Recommended Due Date |  |  | : | 2015-12-20
Environmental Conditions
| Ambient Temperature |  |  | : | 20 °C | ± 1 °C
| Relative Humidity |  |  | : | 50% | ± 15 %
| Location of Calibration |  |  | : | In-Lab
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. standards used to perform this
calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. the result reported herein apply only to the calibration of the item described above.
| all calibrations are performed manufacture's specifications, full, without the expressed written consent of | [9 empty] | 1 | Mr.Sombut Srikampa
| SP Metrology System (Thailand). | [9 empty] | 3 | Mr. Natthaphol Boonmee
[11 empty] | 8 | Mr. Vichan Ananta
[11 empty] | 9 | Mr.Kittikorn Kingmali
| Date of Issue : |  | 2015-12-21 |  |  |  | Approved by : |  |  |  | 10 | Mr.Chainarong Matchayamat
[8 empty] | 3 | ( Mr. Natthaphol Boonmee ) |  | 11 | Ms. Arunkamon Raramanus
| Calibrated by : |  | Ms. Arunkamon Raramanus | [5 empty] | Authorized Signatory
Note on O39: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120012-1 |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Universal Length Measuring |  |  | LMI 01-680 PC | 050021 | 1000959-1 | 2017-04-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result (pitch)
Result of Calibration
|  | Certificate No. : |  | SPR15120012-1 | [7 empty] | Page : 3 of 4
Symbol of Taper Thread Ring |  |  | Pitch Diameter | [5 empty] | Minor Diameter
|  |  | Nominal Value (mm) |  | Permissible Deviation (μm) |  |  |  | Nominal Value (mm) |  | Permissible Deviation (μm)
[5 empty] | Upper | Lower | Tolerance |  |  |  | Upper | Lower | Tolerance
0 |  |  | 0 |  | 0 | 0 | 0 |  | 0 |  | 0 | 0 | 0
0 |  |  | 0 |  | 0 | 0 | 0 |  | □ |  | 0 | 0 | 0
Measurement Results Pitch Diameter
| Symbol of Taper Thread Ring Gague |  |  | Positions |  | Measured Values |  |  |  | Deviation |  | Uncertianty ( ± ) µm
[6 empty] | x-axis |  | y-axis
| 0 |  |  | 1/5 T |  | 9.3072 |  | 12.3072 |  | -10.9349 |  | 0.89624
|  |  |  | 1/2 T |  | 10.3072 |  | 13.3072
|  |  |  | 4/5 T |  | 11.3072 |  | 14.3072
| 0 |  |  | 1/5 T |  | 15.7014 |  | 15.7014 |  | -15.7014 |  | 0.89624
|  |  |  | 1/2 T |  | 15.7014 |  | 15.7014
|  |  |  | 4/5 T |  | 15.7014 |  | 15.7014

## Result (minor)
Result of Calibration
|  | Certificate No. : | SPR15120012-1 |  |  |  | Page : 4 of 4
Measurement Results Minor Diameter
| Symbol of Taper Thread Ring |  | Positions | Measured Values |  | Deviation | Uncertianty ( ± ) µm
|  |  |  | x-axis | y-axis
| 0 |  | 1/5 T | 10 | 13 | 1.7446 | 0.44751
|  |  | 1/2 T | 11 | 14
|  |  | 4/5 T | 12 | 15
| 0 |  | 1/5 T | 10 | 13 | 1.7446 | 0.44751
|  |  | 1/2 T | 11 | 14
|  |  | 4/5 T | 12 | 15
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget(P0.8)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 2 |  | x | P | 0.8 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 1.3072 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 0.455 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 0.46188 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2 |  |  | uA2 = | 2WVO
uB1 | Uncertainty of Ring Gauge |  |  | 0.00037033 |  | Normal |  | 2 |  | 1 | 0.00018517
uB2 | Uncertainty of ULM |  |  | 0.00021 |  | Normal |  | 2 |  | 1 | 0.000105 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00067379 |  | Rectangular |  | √3 |  | 1 | 0.00038901
uB5 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06 |  |  |  | Wo = | 1.6845e-07 |  | Wvo = | 3.3689e-07
uB7 | Uncertainty of Flank Angle |  |  | -2.3834e-05 |  | Rectangular |  | √3 |  | 1 | -6.0011e-05
uB8 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uB9 | Uncertainty of Temperature |  |  | 9.2e-06 |  | Rectangular |  | √3 |  | 1 | 5.3116e-06 |  |  | dD = | Ball diameter |  | 0.455
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.00044812 | 2268316.9061 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 0.89624 | 2.00 |  | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(Minor)
Uncertainty Budget of Thread Ring Gauge
Nominal Value | Repeatability |  | Uncertainty of GB |  | Uncertainty of ULM |  | Resolution of ULM |  | Elastic Deformation |  | Temperature Effect |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
8.9175 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00022375 | 0 | 0.022559 | 2.00 | 0.44751
8.9175 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00022375 | 0 | 0.022559 | 2.00 | 0.44751
8 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 9.2e-05 | 5.3116e-05 | 0.00022222 | 0 | 0.021947 | 2.00 | 0.44444
10 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.000115 | 6.6395e-05 | 0.00023058 | 0 | 0.025442 | 2.00 | 0.46117
12 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.000138 | 7.9674e-05 | 0.00023475 | 0 | 0.027332 | 2.00 | 0.4695
18 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.000207 | 0.00011951 | 0.00025108 | 0 | 0.03577 | 2.00 | 0.50217
20 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.00023 | 0.00013279 | 0.00025767 | 0 | 0.039673 | 2.00 | 0.51534
25 | 0 | 0 | 0 | 0 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.0002875 | 0.00016599 | 0.00028515 | 0 | 0.059505 | 2.00 | 0.57031
30 | 0 | 0 | 0 | 0 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.000345 | 0.00019919 | 0.00030567 | 0 | 0.078572 | 2.00 | 0.61134
75 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.0008625 | 0.00049796 | 0.00056704 | 0 | 0.93043 | 2.00 | 1.1341
90 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.001035 | 0.00059756 | 0.00065623 | 0 | 1.669 | 2.00 | 1.3125
100 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.00115 | 0.00066395 | 0.00071721 | 0 | 2.3814 | 2.00 | 1.4344
125 | 0 | 0 | 0.00018 | 0.00010392 | 0.00039 | 0.000195 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.0014375 | 0.00082994 | 0.0008793 | 0 | 5.38 | 2.00 | 1.7586
150 | 0 | 0 | 0.00018 | 0.00010392 | 0.00039 | 0.000195 | 5e-06 | 2.8868e-06 | 0.00032647 | 0.00018849 | 0.001725 | 0.00099593 | 0.0010374 | 0 | 10.4245 | 2.00 | 2.0748
Elastic Deformation
[9 empty] | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
[8 empty] | Wo = | 8.1617e-08 |  | Wvo = | 1.6323e-07
2 Ball Stylus of SP
[7 empty] | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2 |  |  |  | 0.2 | 1.1
[7 empty] | dD = | Ball diameter |  | 4 | [5 empty] | 0.25 | 1.35
[7 empty] | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25 |  |  | 0.28 | 1.65
[7 empty] | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2 |  | 0.455 | 2.05
[16 empty] | 0.53 | 2.55
[16 empty] | 0.62 | 3.2
[16 empty] | 0.725 | 4
0.895
Note on F4: Certificate of Calubration Gauge Block
Note on H4: Certificate of Calubration ULM
Note on N4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = √3
Note on I6: Divisor = 2
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3

## Uncertainty Budget(P1.5)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 5 |  | x | P | 1.5 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 3.701 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 0.895 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 0.86603 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2 |  |  | uA2 = | 2WVO
uB1 | Uncertainty of Ring Gauge |  |  | 0.00037033 |  | Normal |  | 2 |  | 1 | 0.00018517
uB2 | Uncertainty of ULM |  |  | 0.00021 |  | Normal |  | 2 |  | 1 | 0.000105 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00053776 |  | Rectangular |  | √3 |  | 1 | 0.00031047
uB5 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06 |  |  |  | Wo = | 1.3444e-07 |  | Wvo = | 2.6888e-07
uB7 | Uncertainty of Flank Angle |  |  | 0.00010037 |  | Rectangular |  | √3 |  | 1 | 0.00025272
uB8 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uB9 | Uncertainty of Temperature |  |  | 1.725e-05 |  | Rectangular |  | √3 |  | 1 | 9.9593e-06 |  |  | dD = | Ball diameter |  | 0.895
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.00045411 | 2392012.5958 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 0.90822 | 2.00 |  | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(P2.5)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 11 |  | x | P | 2.5 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 8.8349 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 1.35 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 1.4434 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2 |  |  | uA2 = | 2WVO
uB1 | Uncertainty of Ring Gauge |  |  | 0.00037033 |  | Normal |  | 2 |  | 1 | 0.00018517
uB2 | Uncertainty of ULM |  |  | 0.00023 |  | Normal |  | 2 |  | 1 | 0.000115 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB3 | Uncertainty of Elastic Deformation |  |  | 0.0004689 |  | Rectangular |  | √3 |  | 1 | 0.00027072
uB5 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06 |  |  |  | Wo = | 1.1723e-07 |  | Wvo = | 2.3445e-07
uB7 | Uncertainty of Flank Angle |  |  | -0.00032346 |  | Rectangular |  | √3 |  | 1 | -0.00081444
uB8 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uB9 | Uncertainty of Temperature |  |  | 2.875e-05 |  | Rectangular |  | √3 |  | 1 | 1.6599e-05 |  |  | dD = | Ball diameter |  | 1.35
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.00088597 | 34657026.0177 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 1.7719 | 2.00 |  | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2
25 | 0 | 0 | [10 empty] | 1.2 | [8 empty] | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(P4.5)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 22 |  | x | P | 4.5 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 18.1029 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 2.55 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 2.5981 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2 |  |  | uA2 = | 2WVO
uB1 | Uncertainty of Ring Gauge |  |  | 0.00037033 |  | Normal |  | 2 |  | 1 | 0.00018517
uB2 | Uncertainty of ULM |  |  | 0.00027 |  | Normal |  | 2 |  | 1 | 0.000135 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00037933 |  | Rectangular |  | √3 |  | 1 | 0.000219
uB4 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06 |  |  |  | Wo = | 9.4832e-08 |  | Wvo = | 1.8966e-07
uB5 | Uncertainty of Flank Angle |  |  | -0.00016654 |  | Rectangular |  | √3 |  | 1 | -0.00041933
uB6 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uB7 | Uncertainty of Temperature |  |  | 5.175e-05 |  | Rectangular |  | √3 |  | 1 | 2.9878e-05 |  |  | dD = | Ball diameter |  | 2.55
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.00052702 | 4339328.2599 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 1.054 | 2.00 |  | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(P6)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 45 |  | x | P | 6 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 39.8038 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 3.2 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 3.4641 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2
uB1 | Uncertainty of Ring Gauge |  |  | 0.002 |  | Normal |  | 2 |  | 1 | 0.001
uB2 | Uncertainty of ULM |  |  | 0.00027 |  | Normal |  | 2 |  | 1 | 0.000135
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00035168 |  | Rectangular |  | √3 |  | 1 | 0.00020304
uB4 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06
uB5 | Uncertainty of Flank Angle |  |  | -0.00091487 |  | Rectangular |  | √3 |  | 1 | -0.0023036
uB6 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05
uB7 | Uncertainty of Temperature |  |  | 6.9e-05 |  | Rectangular |  | √3 |  | 1 | 3.9837e-05
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.0025235 | 2280971380.9167
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 5.047 | 2.00
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
[15 empty] | uA2 = | 2WVO
[16 empty] | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
[15 empty] | Wo = | 8.7919e-08 |  | Wvo = | 1.7584e-07
[14 empty] | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
[14 empty] | dD = | Ball diameter |  | 3.2
[14 empty] | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
[14 empty] | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2

## Uncertainty Budget(P8)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 90 |  | x | P | 8 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 83.0718 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 4 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 4.6188 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2
uB1 | Uncertainty of Ring Gauge |  |  | 0.002 |  | Normal |  | 2 |  | 1 | 0.001 |  |  |  | uA2 = | 2WVO
uB2 | Uncertainty of ULM |  |  | 0.00039 |  | Normal |  | 2 |  | 1 | 0.000195
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00032647 |  | Rectangular |  | √3 |  | 1 | 0.00018849 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB4 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06
uB5 | Uncertainty of Flank Angle |  |  | -0.0021436 |  | Rectangular |  | √3 |  | 1 | -0.0053973 |  |  |  | Wo = | 8.1617e-08 |  | Wvo = | 1.6323e-07
uB6 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05
uB7 | Uncertainty of Temperature |  |  | 9.2e-05 |  | Rectangular |  | √3 |  | 1 | 5.3116e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.0054962 | 51329722644.9367 |  | dD = | Ball diameter |  | 4
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 10.9924 | 2.00 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
25 | 0 | 0 | [10 empty] | 1.2 | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2 | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(P8)2
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 180 |  | x | P | 8 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 173.0718 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 4 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 4.6188 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 2e-05 |  | Normal |  | 1 |  | 1 | 2e-05 | 2
uB1 | Uncertainty of Ring Gauge |  |  | 0.002 |  | Normal |  | 2 |  | 1 | 0.001 |  |  |  | uA2 = | 2WVO
uB2 | Uncertainty of ULM |  |  | 0.00039 |  | Normal |  | 2 |  | 1 | 0.000195
uB3 | Uncertainty of Elastic Deformation |  |  | 0.00032647 |  | Rectangular |  | √3 |  | 1 | 0.00018849 |  |  |  |  | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
uB4 | Uncertainty of Resolution |  |  | 1e-05 |  | Rectangular |  | 2√3 |  | 1 | 2.8868e-06 |  |  |  | Wo = | 8.1617e-08 |  | Wvo = | 1.6323e-07
uB5 | Uncertainty of Flank Angle |  |  | -0.0021436 |  | Rectangular |  | √3 |  | 1 | -0.0053973
uB6 | Uncertainty of Abbe' Error |  |  | 2e-05 |  | Rectangular |  | √3 |  | 1 | 1.1547e-05 |  |  | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2
uB7 | Uncertainty of Temperature |  |  | 9.2e-05 |  | Rectangular |  | √3 |  | 1 | 5.3116e-05 |  |  | dD = | Ball diameter |  | 4
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.0054962 | 51329722644.9367 |  | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 10.9924 | 2.00 |  | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.3e-05 | 1.3279e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00011501 | 6.6402e-05 | #REF! | 0 | #REF! | #REF! | #REF!

## Uncertainty Budget(Cal)
Uncertainty Budget of Taper Thread Ring Gauge
Description : |  | M | 10 |  | x | P | 0.8 |  | GPII |  |  |  |  | 2 Ball Stylus of SP | [17 empty] | 0.2
Thread angle: |  | 60 | [11 empty] | 0.2 | 1.1 | [16 empty] | 0.25
Slope: |  | 16 |  |  | Pitch Diameter [E] |  | 9.3072 | mm | [5 empty] | 0.25 | 1.35 | [16 empty] | 0.28
[5 empty] | 2-Ball Stylus (dD) |  | 1.35 | mm | [5 empty] | 0.28 | 1.65
[5 empty] | Best Dimeter (do) |  | 0.46188 | mm | [5 empty] | 0.455 | 2.05
[14 empty] | 0.53 | 2.55 | [19 empty] | Unit: µm
[12 empty] | Unit: µm |  | 0.62 | 3.2
Symbol | Source of Uncertainty |  |  | Value |  | Probability |  | Divisor |  | Ci | Ui | Vi or Veff |  | 0.725 | 4
|  |  |  | ± |  | Distrbution |  |  |  |  | ± |  |  | 0.895
uA | Repeatability |  |  | 0 |  | Normal |  | 1 |  | 1 | 0 | 2
uB1 | Uncertainty of Ring Gauge |  |  | 0.7 |  | Normal |  | 2 |  | 1 | 0.35
uB2 | Uncertainty of ULM |  |  | 0.21 |  | Normal |  | 2 |  | 1 | 0.105
uB3 | Uncertainty of Elastic Deformation |  |  | 0 |  | Rectangular |  | √3 |  | 1 | 0
uB4 | Uncertainty of Slop |  |  | 1 |  | Rectangular |  | √3 |  | 1 | 0.57735
uB5 | Uncertainty of Resolution of ULM |  |  | 0.01 |  | Rectangular |  | 2√3 |  | 1 | 0.0028868
uB6 | Uncertainty of 2 Ball Diameter |  |  | 1 |  | Rectangular |  | √3 |  | 1 | 0.57735
uB7 | Uncertainty of Flank Angle |  |  | 0.0030765 |  | Rectangular |  | √3 |  | 1 | 0.0077464
uB8 | Uncertainty of Abbe' Error |  |  | 0.2 |  | Rectangular |  | √3 |  | 1 | 0.11547
uB9 | Uncertainty of Temperature |  |  | 9.2e-06 |  | Rectangular |  | √3 |  | 1 | 5.3116e-06
uc | Combined Uncertainty |  |  |  |  | Normal |  |  |  |  | 0.90199 | 5.9574e+12
u95% | Expanded Uncertainty |  |  |  |  | Nomal k = 2 |  |  |  |  | 1.804 | 2.00
25 | 0 | 0 | [10 empty] | 1.2 | 0.6 | 0.27 | 0.135 | 0 | 0 | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.05e-06 | 4.6477e-06 | #REF! | 0 | #REF! | #REF! | #REF!
30 | 0 | 0 | [10 empty] | 1.4 | 0.7 | 0.27 | 0.135 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 8.625e-06 | 4.9796e-06 | #REF! | 0 | #REF! | #REF! | #REF!
75 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 9.2e-06 | 5.3116e-06 | #REF! | 0 | #REF! | #REF! | #REF!
90 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.15e-05 | 6.6395e-06 | #REF! | 0 | #REF! | #REF! | #REF!
100 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 1.4375e-05 | 8.2994e-06 | #REF! | 0 | #REF! | #REF! | #REF!
125 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.0001725 | 9.9593e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0.39 | 0.195 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 2.875e-05 | 1.6599e-05 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | #VALUE! | #VALUE! | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00020546 | 0.00011863 | #REF! | 0 | #REF! | #REF! | #REF!
150 | 0 | 0 | [10 empty] | 2.8 | 1.4 | 0 | 0 | #REF! | #REF! | 0 | 0 | #REF! | #REF! | 1 | 0.57735 | #REF! | #REF! | #REF! | #REF! | 0.00020546 | 0.00011863 | #REF! | 0 | #REF! | #REF! | #REF!

## Cert of STD
Certificate of Calibration (ULM) | [5 empty] | Certificate of Calibration (Ring Gauge) |  |  |  |  | Certificate of Calibration (Gauge Block)
SP-SD-001 | [5 empty] | SP-SD-007 |  |  |  |  | SP-SD-009
Due Date |  |  | 2016-04-28 |  |  | Due Date |  | 2016-10-08 |  |  | Due Date |  | 2016-05-13
1 | mm | 0.21 | µm | 0.00021 | mm | 2 | 0.59 | µm | 0.00059 | mm | 2.5 | 0.08 | µm | 8e-05 | mm
1.01 | mm | 0.21 | µm | 0.00021 | mm | 3 | 0.59 | µm | 0.00059 | mm | 5.1 | 0.09 | µm | 9e-05 | mm
1.05 | mm | 0.21 | µm | 0.00021 | mm | 4 | 0.7 | µm | 0.0007 | mm | 7.7 | 0.09 | µm | 9e-05 | mm
1.1 | mm | 0.21 | µm | 0.00021 | mm | 5 | 0.7 | µm | 0.0007 | mm | 10.3 | 0.09 | µm | 9e-05 | mm
2 | mm | 0.21 | µm | 0.00021 | mm | 6 | 0.75 | µm | 0.00075 | mm | 12.9 | 0.09 | µm | 9e-05 | mm
5 | mm | 0.21 | µm | 0.00021 | mm | 8 | 0.75 | µm | 0.00075 | mm | 15 | 0.1 | µm | 0.0001 | mm
10 | mm | 0.21 | µm | 0.00021 | mm | 10 | 0.8 | µm | 0.0008 | mm | 17.6 | 0.1 | µm | 0.0001 | mm
20 | mm | 0.23 | µm | 0.00023 | mm | 12 | 0.87 | µm | 0.00087 | mm | 20.2 | 0.1 | µm | 0.0001 | mm
30 | mm | 0.27 | µm | 0.00027 | mm | 16 | 0.99 | µm | 0.00099 | mm | 22.8 | 0.1 | µm | 0.0001 | mm
40 | mm | 0.27 | µm | 0.00027 | mm | 18 | 0.99 | µm | 0.00099 | mm | 25 | 0.11 | µm | 0.00011 | mm
50 | mm | 0.27 | µm | 0.00027 | mm | 20 | 1.1 | µm | 0.0011 | mm | 50 | 0.13 | µm | 0.00013 | mm
60 | mm | 0.32 | µm | 0.00032 | mm | 22 | 1.2 | µm | 0.0012 | mm | 75 | 0.16 | µm | 0.00016 | mm
70 | mm | 0.32 | µm | 0.00032 | mm | 25 | 1.2 | µm | 0.0012 | mm | 100 | 0.18 | µm | 0.00018 | mm
80 | mm | 0.39 | µm | 0.00039 | mm | 28 | 1.4 | µm | 0.0014 | mm
90 | mm | 0.39 | µm | 0.00039 | mm | 30 | 1.4 | µm | 0.0014 | mm
100 | mm | 0.39 | µm | 0.00039 | mm | 75 | 2.8 | µm | 0.0028 | mm
Note on B5: Universal Length Measuring
Note on O5: Grade 0
Note on E6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on K6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on Q6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B7: Normal Value (mm)
Note on D7: มาจากค่าความไม่แน่นอนของการวัด ULM
Note on F7: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on J7: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Plain Ring Gauge
Note on L7: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on P7: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on R7: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน