## Data Record
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[9 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [8 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Riser Block | [8 empty] | Location |  | In Lab
Customer Name : |  |  |  | LG
Equipment Name : |  |  |  | Riser Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : | [10 empty] | Due Date :
Referance Standard : | [10 empty] | Due Date :
1. Block Pitch Accuracy and Paralellism
Nominal Lengh |  | Position |  | Position Point |  | UUC Error | [8 empty] | Average |  |  |  | Deviation
[6 empty] | X1 |  |  | X2 | X3 |  | X4
150 |  | A |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  |  | 0.00015
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0001 |  | -0.0001 |  |  | -0.0001
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002 |  |  | -0.0002
150 |  | B |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  |  | 0.00015
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0001 |  | -0.0001 |  |  | -0.0001
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002 |  |  | -0.0002
150 |  | C |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  |  | 0.00015
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 0.0001
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0001 |  | -0.0001 |  |  | -0.0001
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002 |  |  | -0.0002
Nominal Lengh |  | Position |  | Position Point |  | UUC Error | [8 empty] | STD Reading |  |  | Error |  |  | Repeatability
[6 empty] | X1 |  |  | X2 | X3 |  | X4
150 |  | A |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 150.0001 |  |  | -0.0001 |  |  | 5.2042e-05
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | -0.0001 |  | -0.0001 |  |  | 150 |  |  | 0
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0002 |  | -0.0002 |  |  | 149.9999 |  |  | 0.00015
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | 0.0001 |  | 0.0001 |  |  | 150 |  |  | 5e-05
150 |  | B |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 150.0001 |  |  | -0.0001 |  |  | 5.2042e-05
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | -0.0001 |  | -0.0001 |  |  | 150 |  |  | 0
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0002 |  | -0.0002 |  |  | 149.9999 |  |  | 0.00015
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | 0.0001 |  | 0.0001 |  |  | 150 |  |  | 5e-05
150 |  | C |  | 1 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | 150.0001 |  |  | -0.0001 |  |  | 5.5434e-05
|  |  |  | 2 |  | 0.0001 |  |  | 0.0001 | -0.0001 |  | -0.0001 |  |  | 150 |  |  | 0
|  |  |  | 3 |  | -0.0001 |  |  | -0.0001 | -0.0002 |  | -0.0002 |  |  | 149.9999 |  |  | 0.00015
|  |  |  | 4 |  | -0.0002 |  |  | -0.0002 | 0 |  | 0 |  |  | 149.9999 |  |  | 0.0001
Calibrate By :
| 11 | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120023-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | : | LG
| Equipment Name |  |  |  | : | Riser Block
| Manufacturer |  |  |  | : | Mittutoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | sdfdthj
| ID. Number |  |  |  | : | 123
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  |  | : | 2016-05-05
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-05-06
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2017-05-06
| Calibration Procedure |  |  |  |  | : | SP-CPD-04-40
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-05-07 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120023-1 |  |  |  | Page : 2 of 3
Reference Standards
Equipment Name |  |  | Model |  | Serial No. | Certificate No. | Due. Date
Linear Height Master |  |  | TESA-HITE M+ |  | 5C062602 | MTO 150453-2 | 2016-06-26
Gauge Block Set |  |  | N/A |  | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  |  | SPR15120023-1 | [5 empty] | Page : 3 of 3
| Nominal Lengh |  | Position |  | Position Point | STD Reading | Error |  | Uncertainty ( ± ) µm
| 150 |  | A |  | 1 | 150.0001 | -0.0001 |  | 9.415
[5 empty] | 2 | 150 | 0
[5 empty] | 3 | 149.9999 | 0.00015
[5 empty] | 4 | 150 | 5e-05
| 150 |  | B |  | 1 | 150.0001 | -0.0001 |  | 9.2019
[5 empty] | 2 | 150 | 0
[5 empty] | 3 | 149.9999 | 0.00015
[5 empty] | 4 | 150 | 5e-05
| 150 |  | C |  | 1 | 150.0001 | -0.0001 |  | 9.202
[5 empty] | 2 | 150 | 0
[5 empty] | 3 | 149.9999 | 0.00015
[5 empty] | 4 | 149.9999 | 0.0001
|  | Block Pitch Accuracy and Paralellism Devation Position A | [5 empty] | 0.00015 | mm
|  | Block Pitch Accuracy and Paralellism Devation Position B | [5 empty] | 0.00015 | mm
|  | Block Pitch Accuracy and Paralellism Devation Position C | [5 empty] | 0.00015 | mm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget Riser Block
Nominal Value |  | Repeatability |  | Uncertainty of STD |  | Temperature Effect |  | Resolution of STD |  | Variation of G/B |  | Wring Gauge Block |  | Contact Deformation |  | Uc | Ui | Veff | K95 | U95%
(mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value |  | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
150 | A | 5.2042e-05 | 5.2042e-05 | 0.002 | 0.001 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0077778 | 0.0044905 | 4e-05 | 2.3094e-05 | 0 | 0 | 0.0047075 | 5.2042e-05 | 200853058.5633 | 2 | 9.415
| B | 5.2042e-05 | 5.2042e-05 | 0.002 | 0.001 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0077778 | 0.0044905 | 4e-05 | 2.3094e-05 | 0 | 0 | 0.0046009 | 5.2042e-05 | 183275640.9264 | 2 | 9.2019
| C | 5.5434e-05 | 5.5434e-05 | 0.002 | 0.001 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.0077778 | 0.0044905 | 4e-05 | 2.3094e-05 | 0 | 0 | 0.004601 | 5.5434e-05 | 142371273.6682 | 2 | 9.202
Contact Deformation
[14 empty] | STD | UUC
[14 empty] | Steel | Steel | 0 | µm
[14 empty] | Steel | Ceramic | 0.00575 | µm
[14 empty] | Steel | Tungsten | -0.2026 | µm
Note on F4: Certificate of Calubration Gauge Block
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on J4: Resolution of Lenear Hi Master
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3
Note on Q6: Divisor = √3

## Cert STD
SP-SD-006 |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
Due Date |  | 2015-06-26 |  |  | SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
0 | 2 | µm | 0.002 | mm | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09
20 | 2 | µm | 0.002 | mm | 2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
50 | 2 | µm | 0.002 | mm | 5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
100 | 2 | µm | 0.002 | mm | 7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
150 | 2 | µm | 0.002 | mm | 10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
200 | 2 | µm | 0.002 | mm | 12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
250 | 2 | µm | 0.002 | mm | 15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
300 | 2 | µm | 0.002 | mm | 17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
400 | 2 | µm | 0.002 | mm | 20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
500 | 2 | µm | 0.002 | mm | 22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
600 | 2 | µm | 0.002 | mm | 25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
700 | 2 | µm | 0.002 | mm | [5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
[5 empty] | Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
[5 empty] | Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[15 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[15 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[15 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[15 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[15 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[15 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[15 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[15 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[15 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[15 empty] | 14 | 0.07 | µm | 7e-05 | mm
[15 empty] | 15 | 0.07 | µm | 7e-05 | mm
[15 empty] | 16 | 0.07 | µm | 7e-05 | mm
[15 empty] | 17 | 0.07 | µm | 7e-05 | mm
[15 empty] | 18 | 0.07 | µm | 7e-05 | mm
[15 empty] | 19 | 0.07 | µm | 7e-05 | mm
[15 empty] | 20 | 0.07 | µm | 7e-05 | mm
[15 empty] | 21 | 0.07 | µm | 7e-05 | mm
[15 empty] | 22 | 0.07 | µm | 7e-05 | mm
[15 empty] | 23 | 0.07 | µm | 7e-05 | mm
[15 empty] | 24 | 0.07 | µm | 7e-05 | mm
[15 empty] | 25 | 0.07 | µm | 7e-05 | mm
[15 empty] | 50 | 0.09 | µm | 9e-05 | mm
[15 empty] | 75 | 0.1 | µm | 0.0001 | mm
[15 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on D3: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I3: Grade 0
Note on O3: Grade 0
Note on U3: Grade 0
Note on AA3: Grade 0
Note on AG3: Grade 0
Note on C4: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Linear Hight Gauge
Note on E4: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on K4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on Q4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on W4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AC4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AI4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on L5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on P5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on R5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on V5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on X5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AB5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AD5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AH5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AJ5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน