## Data Record (10ชิ้น)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Gauge Block | [7 empty] | Location |  | In Lab
Customer Name : |  |  | LG
Equipment Name : |  |  | Gauge Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
Referance Standard : | [9 empty] | Due Date :
[20 empty] | Unit : |  | µm
Nominal Lengh | Position Point |  | UUC Error | [7 empty] | l c |  |  | f o |  |  | f u |  |  | Repeatability
|  |  | X1 |  | X2 |  |  | X3 | X4
1 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
11 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
2 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
22 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
3 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
33 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
4 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
44 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
5 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
55 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
6 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
66 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
7 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
77 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
8 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
88 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
9 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
99 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
10 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
100 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0

## Data Record (20ชิ้น)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Gauge Block | [7 empty] | Location |  | In Lab
Customer Name : |  |  | LG
Equipment Name : |  |  | Gauge Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
Referance Standard : | [9 empty] | Due Date :
[20 empty] | Unit : |  | µm
Nominal Lengh | Position Point |  | UUC Error | [7 empty] | l c |  |  | f o |  |  | f u |  |  | Repeatability
|  |  | X1 |  | X2 |  |  | X3 | X4
11 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
110 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
12 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
120 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
13 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
130 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
14 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
140 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
15 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
150 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
16 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
160 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
17 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
170 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
18 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
180 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
19 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
190 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
20 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
200 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0

## Data Record (30ชิ้น)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Gauge Block | [7 empty] | Location |  | In Lab
Customer Name : |  |  | LG
Equipment Name : |  |  | Gauge Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
Referance Standard : | [9 empty] | Due Date :
[20 empty] | Unit : |  | µm
Nominal Lengh | Position Point |  | UUC Error | [7 empty] | l c |  |  | f o |  |  | f u |  |  | Repeatability
|  |  | X1 |  | X2 |  |  | X3 | X4
21 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
210 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
22 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
220 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
23 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
230 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
24 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
240 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
25 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
250 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
26 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
260 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
27 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
270 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
28 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
280 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
29 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
290 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
30 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
300 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0

## Data Record (40ชิ้น)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Gauge Block | [7 empty] | Location |  | In Lab
Customer Name : |  |  | LG
Equipment Name : |  |  | Gauge Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
Referance Standard : | [9 empty] | Due Date :
[20 empty] | Unit : |  | µm
Nominal Lengh | Position Point |  | UUC Error | [7 empty] | l c |  |  | f o |  |  | f u |  |  | Repeatability
|  |  | X1 |  | X2 |  |  | X3 | X4
31 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
310 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
32 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
320 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
33 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
330 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
34 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
340 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
35 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
350 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
36 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
360 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
37 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
370 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
38 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
380 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
39 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
390 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
40 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
400 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0

## Data Record (50ชิ้น)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  | SPR15120023-1 | [8 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  | 2015-12-12 |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  | 20 | °C | 50 | %RH
Gauge Block | [7 empty] | Location |  | In Lab
Customer Name : |  |  | LG
Equipment Name : |  |  | Gauge Block |  |  |  |  | Manufacturer : |  | Mittutoyo |  |  |  |  | Model : |  |  | 123
Serial No. : |  | sdfdthj |  |  |  | ID No : |  | 123 |  | Range : |  |  | 0 |  | to | 300 |  | mm
Overall Inspection |  |  |  | Good |  |  | Not Good
Referance Standard : | [9 empty] | Due Date :
Referance Standard : | [9 empty] | Due Date :
[20 empty] | Unit : |  | µm
Nominal Lengh | Position Point |  | UUC Error | [7 empty] | l c |  |  | f o |  |  | f u |  |  | Repeatability
|  |  | X1 |  | X2 |  |  | X3 | X4
41 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
410 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
42 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
420 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
43 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
430 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
44 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
440 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
45 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
450 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
46 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
460 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
47 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
470 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
48 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
480 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
49 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
490 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
50 | C |  | 0.0001 |  | 0.0001 |  |  | 0.0001 | 0.0001 |  | 0.0001 |  |  | -0.0001 |  |  | 0.0003 |  |  | 0
| 1 |  | -0.0002 |  | -0.0002 |  |  | -0.0002 | -0.0002 |  | -0.0002
| 2 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
S/N | 3 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0
500 | 4 |  | 0 |  | 0 |  |  | 0 | 0 |  | 0

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120023-1 | [5 empty] | Page : 1 of 5
| Customer |  |  |  | : | LG
| Equipment Name |  |  |  | : | Gauge Block
| Manufacturer |  |  |  | : | Mittutoyo
| Model |  |  |  | : | 1900-05-02
| Serial Number |  |  |  | : | sdfdthj
| ID. Number |  |  |  | : | 123
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  | : | 2015-12-12
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  | : | 2015-12-12
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  | : | 2016-12-11
| Calibration Procedure |  |  |  |  | : | SP-CPT-04-44
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2015-12-13 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  |  | 3 |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120023-1 |  |  | Page : 2 of 5
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Universal Length Measuring |  |  | LMI 01-680 PC | 050021 | 1000959-1 | 2017-04-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  |  | SPR15120023-1 | [5 empty] | Page : 3 of 5
[9 empty] | Unit : | mm
|  | Nominal Value | Serial No. |  | l c | f o | f u | Uncertainty ( ± ) µm
|  | 1 | 11 |  | 0.0001 | -0.0001 | 0.0003 | 0.2263
|  | 2 | 22 |  | 0.0001 | -0.0001 | 0.0003 | 0.22009
|  | 3 | 33 |  | 0.0001 | -0.0001 | 0.0003 | 0.22208
|  | 4 | 44 |  | 0.0001 | -0.0001 | 0.0003 | 0.22484
|  | 5 | 55 |  | 0.0001 | -0.0001 | 0.0003 | 0.22835
|  | 6 | 66 |  | 0.0001 | -0.0001 | 0.0003 | 0.23255
|  | 7 | 77 |  | 0.0001 | -0.0001 | 0.0003 | 0.23743
|  | 8 | 88 |  | 0.0001 | -0.0001 | 0.0003 | 0.24294
|  | 9 | 99 |  | 0.0001 | -0.0001 | 0.0003 | 0.24903
|  | 10 | 100 |  | 0.0001 | -0.0001 | 0.0003 | 0.25567
|  | 11 | 110 |  | 0.0001 | -0.0001 | 0.0003 | 0.28137
|  | 12 | 120 |  | 0.0001 | -0.0001 | 0.0003 | 0.28849
|  | 13 | 130 |  | 0.0001 | -0.0001 | 0.0003 | 0.29603
|  | 14 | 140 |  | 0.0001 | -0.0001 | 0.0003 | 0.30396 |  |  | lc - ln | Deviation in length from the norminal length
|  | 15 | 150 |  | 0.0001 | -0.0001 | 0.0003 | 0.31226 |  |  |  | of the central length
|  | 16 | 160 |  | 0.0001 | -0.0001 | 0.0003 | 0.3209 |  |  | fo = lmax - lc | Deviation of Maximum Length
|  | 17 | 170 |  | 0.0001 | -0.0001 | 0.0003 | 0.32984 |  |  |  | from the Central Length
|  | 18 | 180 |  | 0.0001 | -0.0001 | 0.0003 | 0.33907 |  |  | fu = lc - lmin | Deviation of Minimum Length
|  | 19 | 190 |  | 0.0001 | -0.0001 | 0.0003 | 0.34855 |  |  |  | from the Central Length
|  | 20 | 200 |  | 0.0001 | -0.0001 | 0.0003 | 0.35828 |  |  | V = lmax - lmin | Variation in Length
|  | 21 | 210 |  | 0.0001 | -0.0001 | 0.0003 | 0.39446 |  |  | ln | Nominal Length
|  | 22 | 220 |  | 0.0001 | -0.0001 | 0.0003 | 0.40395 |  |  | lc | Central Length
|  | 23 | 230 |  | 0.0001 | -0.0001 | 0.0003 | 0.41366 |  |  | lmax | Maximum Length
|  | 24 | 240 |  | 0.0001 | -0.0001 | 0.0003 | 0.42356 |  |  | lmin | Minimum Length
|  | 25 | 250 |  | 0.0001 | -0.0001 | 0.0003 | 0.43364

## Result (2)
Result of Calibration
| Certificate No. : |  |  | SPR15120023-1 | [5 empty] | Page : 4 of 5
[9 empty] | Unit : | mm
|  | Nominal Value | Serial No. |  | l c | f o | f u | Uncertainty ( ± ) µm
|  | 26 | 11 |  | 0.0001 | -0.0001 | 0.0003 | 0.45192
|  | 27 | 22 |  | 0.0001 | -0.0001 | 0.0003 | 0.46731
|  | 28 | 33 |  | -0.0002 | -0.0001 | 0.0003 | 0.47758
|  | 29 | 44 |  | 0.0001 | -0.0001 | 0.0003 | 0.48799
|  | 30 | 55 |  | 0.0001 | -0.0001 | 0.0003 | 0.49369
|  | 31 | 310 |  | 0.0001 | -0.0001 | 0.0003 | 0.52085
|  | 32 | 320 |  | 0.0001 | -0.0001 | 0.0003 | 0.53141
|  | 33 | 330 |  | 0.0001 | -0.0001 | 0.0003 | 0.54209
|  | 34 | 340 |  | 0.0001 | -0.0001 | 0.0003 | 0.55288
|  | 35 | 350 |  | 0.0001 | -0.0001 | 0.0003 | 0.56377
|  | 36 | 360 |  | 0.0001 | -0.0001 | 0.0003 | 0.57477
|  | 37 | 370 |  | 0.0001 | -0.0001 | 0.0003 | 0.58586
|  | 38 | 380 |  | 0.0001 | -0.0001 | 0.0003 | 0.59704
|  | 39 | 390 |  | 0.0001 | -0.0001 | 0.0003 | 0.60831 |  |  | lc - ln | Deviation in length from the norminal length
|  | 40 | 400 |  | 0.0001 | -0.0001 | 0.0003 | 0.61965 |  |  |  | of the central length
|  | 41 | 410 |  | 0.0001 | -0.0001 | 0.0003 | 0.64981 |  |  | fo = lmax - lc | Deviation of Maximum Length
|  | 42 | 420 |  | 0.0001 | -0.0001 | 0.0003 | 0.66097 |  |  |  | from the Central Length
|  | 43 | 430 |  | 0.0001 | -0.0001 | 0.0003 | 0.67222 |  |  | fu = lc - lmin | Deviation of Minimum Length
|  | 44 | 440 |  | 0.0001 | -0.0001 | 0.0003 | 0.68353 |  |  |  | from the Central Length
|  | 45 | 450 |  | 0.0001 | -0.0001 | 0.0003 | 0.69492 |  |  | V = lmax - lmin | Variation in Length
|  | 46 | 460 |  | 0.0001 | -0.0001 | 0.0003 | 0.70637 |  |  | ln | Nominal Length
|  | 47 | 470 |  | 0.0001 | -0.0001 | 0.0003 | 0.71788 |  |  | lc | Central Length
|  | 48 | 480 |  | 0.0001 | -0.0001 | 0.0003 | 0.72946 |  |  | lmax | Maximum Length
|  | 49 | 490 |  | 0.0001 | -0.0001 | 0.0003 | 0.74109 |  |  | lmin | Minimum Length
|  | 50 | 500 |  | 0.0001 | -0.0001 | 0.0003 | 0.7224
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %.
- End of Certificate -

## Result (3)
Result of Calibration
| Certificate No. : | SPR15120023-1 |  |  | Page : 5 of 5
|  |  | lc - ln | Deviation in length from the norminal length
of the central length
|  |  | fo = lmax - lc | Deviation of Maximum Length
from the Central Length
|  |  | fu = lc - lmin | Deviation of Minimum Length
from the Central Length
|  |  | V = lmax - lmin | Variation in Length
|  |  | ln | Nominal Length
|  |  | lc | Central Length
|  |  | lmax | Maximum Length
|  |  | lmin | Minimum Length
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Result (ไม่เกิน12ชิ้น)
Result of Calibration
| Certificate No. : |  |  | SPR15120023-1 | [5 empty] | Page : 3 of 3
[9 empty] | Unit : | mm
|  | Nominal Value | Serial No. |  | l c | f o | f u | Uncertainty ( ± ) µm
|  | 1 | 11 |  | 0.0001 | -0.0001 | 0.0003 | 0.2263
|  | 2 | 22 |  | 0.0001 | -0.0001 | 0.0003 | 0.22009
|  | 3 | 33 |  | 0.0001 | -0.0001 | 0.0003 | 0.22208
|  | 4 | 44 |  | 0.0001 | -0.0001 | 0.0003 | 0.22484
|  | 5 | 55 |  | 0.0001 | -0.0001 | 0.0003 | 0.22835
|  | 6 | 66 |  | 0.0001 | -0.0001 | 0.0003 | 0.23255
|  | 7 | 77 |  | 0.0001 | -0.0001 | 0.0003 | 0.23743
|  | 8 | 88 |  | 0.0001 | -0.0001 | 0.0003 | 0.24294
|  | 9 | 99 |  | 0.0001 | -0.0001 | 0.0003 | 0.24903
|  | 10 | 100 |  | 0.0001 | -0.0001 | 0.0003 | 0.25567
|  | 11 | 110 |  | 0.0001 | -0.0001 | 0.0003 | 0.28137
|  | 12 | 120 |  | 0.0001 | -0.0001 | 0.0003 | 0.28849
[6 empty] | lc - ln | Deviation in length from the norminal length
of the central length
[6 empty] | fo = lmax - lc | Deviation of Maximum Length
from the Central Length
[6 empty] | fu = lc - lmin | Deviation of Minimum Length
from the Central Length
[6 empty] | V = lmax - lmin | Variation in Length
[6 empty] | ln | Nominal Length
[6 empty] | lc | Central Length
[6 empty] | lmax | Maximum Length
[6 empty] | lmin | Minimum Length
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Gauge Block
Nominal Value | Repeatability |  | Uncertainty of GB |  | Contact Deformation |  | Temperature Effect |  | Uncertainty of ULM |  | Resolution of ULM |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
1 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 1.15e-05 | 6.6395e-06 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 0.00011315 | 0 | ∞ | 2 | 0.2263
2 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 2.3e-05 | 1.3279e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011004 | 0 | ∞ | 2 | 0.22009
3 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 3.45e-05 | 1.9919e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011104 | 0 | ∞ | 2 | 0.22208
4 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 4.6e-05 | 2.6558e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011242 | 0 | ∞ | 2 | 0.22484
5 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 5.75e-05 | 3.3198e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011417 | 0 | ∞ | 2 | 0.22835
6 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 6.9e-05 | 3.9837e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011628 | 0 | ∞ | 2 | 0.23255
7 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 8.05e-05 | 4.6477e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00011872 | 0 | ∞ | 2 | 0.23743
8 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 9.2e-05 | 5.3116e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00012147 | 0 | ∞ | 2 | 0.24294
9 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 0.0001035 | 5.9756e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00012452 | 0 | ∞ | 2 | 0.24903
10 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 0.000115 | 6.6395e-05 | 0.00021 | 0.000105 | 5e-06 | 2.8868e-06 | 0.00012783 | 0 | ∞ | 2 | 0.25567
11 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0001265 | 7.3035e-05 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00014069 | 0 | ∞ | 2 | 0.28137
12 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000138 | 7.9674e-05 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00014424 | 0 | ∞ | 2 | 0.28849
13 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0001495 | 8.6314e-05 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00014801 | 0 | ∞ | 2 | 0.29603
14 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000161 | 9.2953e-05 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00015198 | 0 | ∞ | 2 | 0.30396
15 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0001725 | 9.9593e-05 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00015613 | 0 | ∞ | 2 | 0.31226
16 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000184 | 0.00010623 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00016045 | 0 | ∞ | 2 | 0.3209
17 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0001955 | 0.00011287 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00016492 | 0 | ∞ | 2 | 0.32984
18 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000207 | 0.00011951 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00016953 | 0 | ∞ | 2 | 0.33907
19 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0002185 | 0.00012615 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00017428 | 0 | ∞ | 2 | 0.34855
20 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.00023 | 0.00013279 | 0.00023 | 0.000115 | 5e-06 | 2.8868e-06 | 0.00017914 | 0 | ∞ | 2 | 0.35828
21 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0002415 | 0.00013943 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00019723 | 0 | ∞ | 2 | 0.39446
22 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000253 | 0.00014607 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00020198 | 0 | ∞ | 2 | 0.40395
23 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0002645 | 0.00015271 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00020683 | 0 | ∞ | 2 | 0.41366
24 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.000276 | 0.00015935 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00021178 | 0 | ∞ | 2 | 0.42356
25 | 0 | 0 | 7e-05 | 3.5e-05 | 0 | 0 | 0.0002875 | 0.00016599 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00021682 | 0 | ∞ | 2 | 0.43364
26 | 0 | 0 | 0.00011 | 5.5e-05 | 0 | 0 | 0.000299 | 0.00017263 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00022596 | 0 | ∞ | 2 | 0.45192
27 | 0 | 0 | 0.00013 | 6.5e-05 | 0 | 0 | 0.0003105 | 0.00017927 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00023366 | 0 | ∞ | 2 | 0.46731
28 | 0 | 0 | 0.00013 | 6.5e-05 | 0 | 0 | 0.000322 | 0.00018591 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00023879 | 0 | ∞ | 2 | 0.47758
29 | 0 | 0 | 0.00013 | 6.5e-05 | 0 | 0 | 0.0003335 | 0.00019255 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00024399 | 0 | ∞ | 2 | 0.48799
30 | 0 | 0 | 0.00011 | 5.5e-05 | 0 | 0 | 0.000345 | 0.00019919 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00024685 | 0 | ∞ | 2 | 0.49369
31 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.0003565 | 0.00020583 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00026043 | 0 | ∞ | 2 | 0.52085
32 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.000368 | 0.00021246 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00026571 | 0 | ∞ | 2 | 0.53141
33 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.0003795 | 0.0002191 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00027104 | 0 | ∞ | 2 | 0.54209
34 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.000391 | 0.00022574 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00027644 | 0 | ∞ | 2 | 0.55288
35 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.0004025 | 0.00023238 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00028189 | 0 | ∞ | 2 | 0.56377
36 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.000414 | 0.00023902 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00028739 | 0 | ∞ | 2 | 0.57477
37 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.0004255 | 0.00024566 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00029293 | 0 | ∞ | 2 | 0.58586
38 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.000437 | 0.0002523 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00029852 | 0 | ∞ | 2 | 0.59704
39 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.0004485 | 0.00025894 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00030415 | 0 | ∞ | 2 | 0.60831
40 | 0 | 0 | 0.00017 | 8.5e-05 | 0 | 0 | 0.00046 | 0.00026558 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00030983 | 0 | ∞ | 2 | 0.61965
41 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.0004715 | 0.00027222 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.0003249 | 0 | ∞ | 2 | 0.64981
42 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.000483 | 0.00027886 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00033049 | 0 | ∞ | 2 | 0.66097
43 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.0004945 | 0.0002855 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00033611 | 0 | ∞ | 2 | 0.67222
44 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.000506 | 0.00029214 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00034177 | 0 | ∞ | 2 | 0.68353
45 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.0005175 | 0.00029878 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00034746 | 0 | ∞ | 2 | 0.69492
46 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.000529 | 0.00030542 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00035318 | 0 | ∞ | 2 | 0.70637
47 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.0005405 | 0.00031206 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00035894 | 0 | ∞ | 2 | 0.71788
48 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.000552 | 0.0003187 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00036473 | 0 | ∞ | 2 | 0.72946
49 | 0 | 0 | 0.00023 | 0.000115 | 0 | 0 | 0.0005635 | 0.00032534 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.00037054 | 0 | ∞ | 2 | 0.74109
50 | 0 | 0 | 9e-05 | 4.5e-05 | 0 | 0 | 0.000575 | 0.00033198 | 0.00027 | 0.000135 | 5e-06 | 2.8868e-06 | 0.0003612 | 0 | ∞ | 2 | 0.7224
Contact Deformation
[5 empty] | STD | UUC
[5 empty] | Steel | Steel | 0 | µm
[5 empty] | Steel | Ceramic | 0.00575 | µm
[5 empty] | Steel | Tungsten | -0.2026 | µm
Note on F4: Certificate of Calubration Gauge Block
Note on J4: (Normal Value * 11.5*10^-6 * 1)
Note on L4: Certificate of Calubration Gauge Block
Note on E6: Divisor = 1
Note on G6: Divisor = √3
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3

## Cert of STD
Certificate of Calibration | [5 empty] | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
SP-SD-001 | [5 empty] | SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
1 | mm | 0.21 | µm | 0.00021 | mm | 2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
1.01 | mm | 0.21 | µm | 0.00021 | mm | 5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
1.05 | mm | 0.21 | µm | 0.00021 | mm | 7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
1.1 | mm | 0.21 | µm | 0.00021 | mm | 10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
2 | mm | 0.21 | µm | 0.00021 | mm | 12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
5 | mm | 0.21 | µm | 0.00021 | mm | 15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
10 | mm | 0.21 | µm | 0.00021 | mm | 17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20 | mm | 0.23 | µm | 0.00023 | mm | 20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
30 | mm | 0.27 | µm | 0.00027 | mm | 22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
40 | mm | 0.27 | µm | 0.00027 | mm | 25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
50 | mm | 0.27 | µm | 0.00027 | mm | [5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
60 | mm | 0.32 | µm | 0.00032 | mm | Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
70 | mm | 0.32 | µm | 0.00032 | mm | Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
80 | mm | 0.39 | µm | 0.00039 | mm | [10 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
90 | mm | 0.39 | µm | 0.00039 | mm | [10 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
100 | mm | 0.39 | µm | 0.00039 | mm | [10 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[16 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[16 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[16 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[16 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[16 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[16 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[16 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[16 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[16 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[16 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[16 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[16 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[16 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[16 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[16 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[16 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[16 empty] | 14 | 0.07 | µm | 7e-05 | mm
[16 empty] | 15 | 0.07 | µm | 7e-05 | mm
[16 empty] | 16 | 0.07 | µm | 7e-05 | mm
[16 empty] | 17 | 0.07 | µm | 7e-05 | mm
[16 empty] | 18 | 0.07 | µm | 7e-05 | mm
[16 empty] | 19 | 0.07 | µm | 7e-05 | mm
[16 empty] | 20 | 0.07 | µm | 7e-05 | mm
[16 empty] | 21 | 0.07 | µm | 7e-05 | mm
[16 empty] | 22 | 0.07 | µm | 7e-05 | mm
[16 empty] | 23 | 0.07 | µm | 7e-05 | mm
[16 empty] | 24 | 0.07 | µm | 7e-05 | mm
[16 empty] | 25 | 0.07 | µm | 7e-05 | mm
[16 empty] | 50 | 0.09 | µm | 9e-05 | mm
[16 empty] | 75 | 0.1 | µm | 0.0001 | mm
[16 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on B4: Normal Value (mm)
Note on D4: มาจากค่าความไม่แน่นอนของการวัด ULM
Note on F4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on J4: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on L4: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on P4: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on R4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on V4: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on X4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AB4: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AD4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AH4: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AJ4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน