## Data Record
SP METROLOGY SYSTEM THAILAND | [10 empty] | Certificate No. : |  |  |  | SPR16010012-1 | [10 empty] | Page 1 of 1
[11 empty] | Receive Date : |  |  |  | 2016-01-08 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-13
Equipment Name | [10 empty] | Temp & Humiduty : |  |  |  |  | 24.3 |  | oC | 64.5 |  | %RH
Surface Plate | [10 empty] | Location |  |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | LG Electronics (Thailand)
Equipment Name : |  |  |  |  | Granite Surface Plate | [9 empty] | Manufacturer : |  |  |  | Choten
Model : |  | 1000x2000 | [8 empty] | Serial No. : |  |  | N21012065 | [7 empty] | ID No : |  | AC-IQC-166
Range : |  | Wide |  | 1000 |  |  | mm |  |  | Long |  | 2000 |  |  | mm |  |  | Thick |  | 100 |  |  | mm
Overall Inspection | [6 empty] | Good |  |  |  | Not Good
Referance Standard : | [5 empty] | SP-SD-029 | [10 empty] | Due Date :
Referance Standard : | [5 empty] | SP-ST-004 | [10 empty] | Due Date :
Measurement Results
| Before Temperature : | [6 empty] | Top : | 21.3 |  | °C |  |  |  |  | After Temperature : | [5 empty] | Top : | 21.3 |  | °C
[8 empty] | Bottom : | 21.7 |  | °C | [10 empty] | Bottom : | 21.7 |  | °C
Range : |  |  |  | Wide = |  | 1000 |  |  |  | Long = |  | 2000 |  |  |  | Thick = |  | 100 |  |  |  | Ratio = |  | 2
| a | 5.2e-06 |  |  |  | °k-1 |  |  |  | E | 77000000000 |  |  |  | N/m2
| g | 9.8 |  |  |  | m/s2 |  |  |  | n | 0.26 | [8 empty] | ∆T |  | = | -0.4
| k | 0.103 | [7 empty] | r | 6370000 |  |  |  | m |  |  |  |  | u (a) |  | = | 5.7735e-07
| r | 2750 |  |  |  | kg/m3 |  |  |  | Ld | 2236.068 | [8 empty] | u (∆T) |  | = | -0.2
Longitudinal | [25 empty] | Unit: mm
Line |  |  | A |  | B |  | C |  | D |  | E |  | F |  | G |  | H |  | I |  | J |  | MAX |  | Repeat
Line 1 |  |  | 0.5 |  | 1 |  | 0.001 |  | 0.0015 |  | 0.001 |  | -0.5 |  | -0.5 |  | 0.0015 |  | 0.5 |  | 0.001 |  | 1 |  | 0.067735
Line 2 |  |  | 0.5 |  | 1 |  | -1 |  | -0.5 |  | 0.001 |  | 0.5 |  | -1 |  | 0 |  | 1 |  | 0.001 |  | 1
Line 3 |  |  | 0.5 |  | 1 |  | 0.001 |  | 0.5 |  | 1 |  | 1 |  | -0.5 |  | 0.0015 |  | 1 |  | 0.001 |  | 1
Line 4 |  |  | -0.5 |  | 0.5 |  | 0.001 |  | 0.0015 |  | 0.5 |  | -0.5 |  | -0.5 |  | 1 |  | 0 |  | -0.5 |  | 1
Line 5 |  |  | 0.001 |  | 1 |  | 0 |  | -0.5 |  | 0 |  | 0 |  | -0.5 |  | 0 |  | -0.5 |  | -0.5 |  | 1
Line 6 |  |  | 0.5 |  | 0.0015 |  | -0.5 |  | 0.0015 |  | 0.001 |  | 0.5 |  | 0.001 |  | 0.0015 |  | 0.001 |  | 0.001 |  | 0.5
| Calibrated By : |  |  |  |  | Ms. Arunkamon Raramanus
|  |  |  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16010012-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Granite Surface Plate
| Manufacturer |  |  |  | : | Choten
| Model |  |  |  | : | 1000x2000
| Serial Number |  |  |  | : | N21012065
| ID. Number |  |  |  | : | AC-IQC-166
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-01-08
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2016-01-13
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2017-01-12
| Calibration Procedure |  |  |  | : | SP-CPD-04-17
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-14 |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16010012-1 |  |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  |  | Model | Serial No. | Certificate No. | Due. Date
2016-12-02
2016-10-05
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  |  | SPR16010012-1 | [7 empty] | Page : 3 of 3
Measurement Result
[15 empty] | Unit | mm
| Line |  | A | B | C |  | D |  | E | F | G | H |  | I | J
| 1 |  | 0.5 | 1 | 0.001 |  | 0.0015 |  | 0.001 | -0.5 | -0.5 | 0.0015 |  | 0.5 | 0.001
| 2 |  | 0.5 | 1 | -1 |  | -0.5 |  | 0.001 | 0.5 | -1 | 0 |  | 1 | 0.001
| 3 |  | 0.5 | 1 | 0.001 |  | 0.5 |  | 1 | 1 | -0.5 | 0.0015 |  | 1 | 0.001
| 4 |  | -0.5 | 0.5 | 0.001 |  | 0.0015 |  | 0.5 | -0.5 | -0.5 | 1 |  | 0 | -0.5
| 5 |  | 0.5 | 1 | 0 |  | -0.5 |  | 0 | 0 | -0.5 | 0 |  | -0.5 | -0.5
| 6 |  | 0.001 | 0.0015 | -0.5 |  | 0.0015 |  | 0.001 | 0.5 | 0.001 | 0.0015 |  | 0.001 | 0.001
| Maximun Error |  |  | 1
| The Uncertainty of Measurement |  |  |  |  | ± | 1.0228 | mm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%
- End of Certificate -

## Uncertainty Budget(CMC)
Uncertainty Budget of Surface Plate
| Range : | Wide = | 1000 | Long = | 2000 | Thick = | 100 | Ratio = | 2
| a | 5.2e-06 | °k-1 | E | 77000000000 | N/m2 | ∆T | 0.4
| g | 9.8 | m/s2 | n | 0.26 |  | u (a) | 5.7735e-07
| k | 0.103 |  | rr | 6370000 | m | u (∆T) | 0.2
| r | 2750 | kg/m3 | Ld | =SQRT(W^2+L^2)
Unit: µm
| Symbol | Source of Uncertainty |  |  | Value | Probability Distrbution |  | Divisor | Ci | Ui | Vi or Veff | 300x300 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.0035322 | 0.0020393 | 1.1631e-08 | 6.715e-09 | 0.5114 | 0.067735 | 29244.2286 | 2.00 | 1.0228
| uA | Repeatability |  |  | 0.067735 | Normal |  | 1 | 1 | 0.067735 | 2 | 400x400 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.0062794 | 0.0036254 | 3.6759e-12 | 2.1223e-12 | 0.51141 | 0.067735 | 29246.238 | 2.00 | 1.0228
| uB1 | Uncertainty of Dial Indicator |  |  | 1 | Normal |  | 2 | 1 | 0.5 | ∞ | 500x500 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.0098116 | 0.0056647 | 1.4277e-07 | 8.243e-08 | 0.51143 | 0.067735 | 29250.4753 | 2.00 | 1.0229
| uB2 | Resolution of STD |  |  | 0.1 | Rectangular |  | √3 | 1 | 0.057735 | ∞ | 600x600 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.014129 | 0.0081572 | 1.8609e-07 | 1.0744e-07 | 0.51146 | 0.067735 | 29258.1812 | 2.00 | 1.0229
| uB3 | Difference Temp |  |  | 0.104 | Rectangular |  | √3 | 1 | 0.060044 | ∞ | 750x750 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.022076 | 0.012746 | 4.5433e-07 | 2.6231e-07 | 0.51156 | 0.067735 | 29279.6398 | 2.00 | 1.0231
| uB4 | Reference Point |  |  | 0.00088305 | Rectangular |  | √3 | 1 | 0.00050983 | ∞ | 900x900 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.03179 | 0.018354 | 9.4209e-07 | 5.4392e-07 | 0.51173 | 0.067735 | 29318.681 | 2.00 | 1.0235
| uB5 | Supporting Error |  |  | 7.2692e-10 | Rectangular |  | √3 | 1 | 4.1969e-10 | ∞ | 1000x1000 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.039246 | 0.022659 | 1.4359e-09 | 8.2901e-10 | 0.5119 | 0.067735 | 29358.232 | 2.00 | 1.0238
| Uc | Combined Uncertainty |  |  |  | Normal |  |  |  | 0.5114 | 29243.3567 | 1100x1100 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.047488 | 0.027417 | 2.1023e-06 | 1.2138e-06 | 0.51213 | 0.067735 | 29411.6492 | 2.00 | 1.0243
| U95% | Expanded Uncertainty |  |  |  | Normal k = | 2.00 |  |  | 1.0228 | 2.00 | 1200x1200 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.056515 | 0.032629 | 2.9775e-06 | 1.719e-06 | 0.51244 | 0.067735 | 29481.8754 | 2.00 | 1.0249
[12 empty] | 1400x1400 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.076923 | 0.044412 | 5.5161e-06 | 3.1847e-06 | 0.51332 | 0.067735 | 29686.0572 | 2.00 | 1.0266
[12 empty] | 1500x1500 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.088305 | 0.050983 | 7.2692e-06 | 4.1969e-06 | 0.51393 | 0.067735 | 29827.467 | 2.00 | 1.0279
[12 empty] | 1600x1600 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.10047 | 0.058007 | 9.4103e-06 | 5.433e-06 | 0.51468 | 0.067735 | 30000.6284 | 2.00 | 1.0294
[12 empty] | 1800x1800 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.12716 | 0.073415 | 1.5073e-05 | 8.7027e-06 | 0.51664 | 0.067735 | 30461.0609 | 2.00 | 1.0333
[12 empty] | 2000x2000 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.15699 | 0.090636 | 2.2974e-05 | 1.3264e-05 | 0.51937 | 0.067735 | 31109.2833 | 2.00 | 1.0387
[12 empty] | 2300x2300 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.20761 | 0.11987 | 4.0182e-05 | 2.3199e-05 | 0.52526 | 0.067735 | 32544.7158 | 2.00 | 1.0505
[12 empty] | 2500x2500 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.24529 | 0.14162 | 5.609e-05 | 3.2383e-05 | 0.53064 | 0.067735 | 33900.4617 | 2.00 | 1.0613
[12 empty] | 2600x2600 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.26531 | 0.15317 | 6.5617e-05 | 3.7884e-05 | 0.53384 | 0.067735 | 34725.6947 | 2.00 | 1.0677
[12 empty] | 2800x2800 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.30769 | 0.17765 | 8.8258e-05 | 5.0956e-05 | 0.54137 | 0.067735 | 36726.6361 | 2.00 | 1.0827
[12 empty] | 3000x3000 | 0.067735 | 0.067735 | 1 | 0.5 | 0.1 | 0.057735 | 0.104 | 0.060044 | 0.35322 | 0.20393 | 0.00011631 | 6.715e-05 | 0.55056 | 0.067735 | 39283.2542 | 2.00 | 1.1011

## Cert of STD
Certificate of Calibration (Dial Indicator)
SP-SD-029-1
Due Date | 2016-06-30
Span of Error (ft) | 1 | µm | 0.001 | mm
Span of Error (fe) | 1 | µm | 0.001 | mm
Total Span of Error (fegs) | 1 | µm | 0.001 | mm
Hysterresis Error (fu) | 1 | µm | 0.001 | mm
Repeatability (fw) | 1 | µm | 0.001 | mm
Note on C5: มาจากค่าความไม่แน่นอนของการวัด Dial Indicator