## Data Record
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR15120012-1 |  |  |  |  | Page | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 2015-12-11 |  |  |  | Calibration Date : |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Universal Length Machine | [7 empty] | Location |  |  |  |  | In Lab | [5 empty] | On Site
Customer Name : |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  | ULM | [7 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  | 123 |  |  |  | Serial No. : |  |  | 456 | [6 empty] | ID No : | 789
Range : |  | 0 | to | 600 | mm |  | Resolution : |  |  | 0.0001 |  |  | mm
Overall Inspection |  |  |  |  | Good |  |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Referance Standard : | [13 empty] | Due Date :
| Nominal Value |  | UUC Reading | [10 empty] | Average |  |  | Error |  | Repertability
|  |  | X1 |  | X2 |  |  | X3 |  |  | X4
| 0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  | 0
| 1 |  | 1 |  | 1 |  |  | 1 |  |  | 1 |  |  | 1 |  |  | 0 |  | 0
| 1.005 |  | 1.005 |  | 1.005 |  |  | 1.005 |  |  | 1.005 |  |  | 1.005 |  |  | 0 |  | 0
| 1.01 |  | 1.01 |  | 1.01 |  |  | 1.01 |  |  | 1.01 |  |  | 1.01 |  |  | 0 |  | 0
| 1.05 |  | 1.05 |  | 1.05 |  |  | 1.05 |  |  | 1.05 |  |  | 1.05 |  |  | 0 |  | 0
| 1.1 |  | 1.1 |  | 1.1 |  |  | 1.1 |  |  | 1.1 |  |  | 1.1 |  |  | 0 |  | 0
| 2 |  | 2 |  | 2 |  |  | 2 |  |  | 2 |  |  | 2 |  |  | 0 |  | 0
| 5 |  | 5 |  | 5 |  |  | 5 |  |  | 5 |  |  | 5 |  |  | 0 |  | 0
| 10 |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 10 |  |  | 0 |  | 0
| 20 |  | 20 |  | 20 |  |  | 20 |  |  | 20 |  |  | 20 |  |  | 0 |  | 0
| 50 |  | 50 |  | 50 |  |  | 50 |  |  | 50 |  |  | 50 |  |  | 0 |  | 0
| 75 |  | 75 |  | 75 |  |  | 75 |  |  | 75 |  |  | 75 |  |  | 0 |  | 0
| 100 |  | 100 |  | 100 |  |  | 100 |  |  | 100 |  |  | 100 |  |  | 0 |  | 0
| 200 |  | 200 |  | 200 |  |  | 200 |  |  | 200 |  |  | 200 |  |  | 0 |  | 0
| 300 |  | 300 |  | 300 |  |  | 300 |  |  | 300 |  |  | 300 |  |  | 0 |  | 0
| 400 |  | 400 |  | 400 |  |  | 400 |  |  | 400 |  |  | 400 |  |  | 0 |  | 0
| 500 |  | 500 |  | 500 |  |  | 500 |  |  | 500 |  |  | 500 |  |  | 0 |  | 0
| 600 |  | 600 |  | 600 |  |  | 600 |  |  | 600 |  |  | 600 |  |  | 0 |  | 0
Calibrated By :

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR15120012-1 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
| Equipment Name |  |  | : | ULM
| Manufacturer |  |  | : | Mitutoyo
| Model |  |  | : | 123
| Serial Number |  |  | : | 456
| ID. Number |  |  | : | 789
| Received Date |  |  | : | 2015-12-11
| Calibration Date |  |  | : | 2015-12-12
| Recommended Due Date |  |  | : | 2016-12-12
Environmental Conditions
| Ambient Temperature |  |  | : | 20 °C | ± 1 °C
| Relative Humidity |  |  | : | 50% | ± 15 %
| Location of Calibration |  |  | : | In-Lab
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. standards used to perform this
calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. the result reported herein apply only to the calibration of the item described above.
| all calibrations are performed manufacture's specifications, full, without the expressed written consent of | [9 empty] | 1 | Mr.Sombut Srikampa
| SP Metrology System (Thailand). | [9 empty] | 3 | Mr. Natthaphol Boonmee
[11 empty] | 8 | Mr. Vichan Ananta
[11 empty] | 9 | Mr.Kittikorn Kingmali
| Date of Issue : |  | 2015-12-13 |  |  |  | Approved by : |  |  |  | 10 | Mr.Chainarong Matchayamat
[8 empty] | 3 | ( Mr. Natthaphol Boonmee ) |  | 11 | Ms. Arunkamon Raramanus
| Calibrated by : |  | Ms. Arunkamon Raramanus | [5 empty] | Authorized Signatory
Note on O39: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number : |  | SPR15120012-1 |  |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  | SPR15120012-1 | [5 empty] | Page : 3 of 3
[7 empty] | Unit : | mm
|  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  | 0 |  | 0 | 0 | 0.062458
|  | 1 |  | 1 | 0 | 0.063854
|  | 1.005 |  | 1.005 | 0 | 0.063868
|  | 1.01 |  | 1.01 | 0 | 0.063882
|  | 1.05 |  | 1.05 | 0 | 0.063996
|  | 1.1 |  | 1.1 | 0 | 0.064144
|  | 2 |  | 2 | 0 | 0.06787
|  | 5 |  | 5 | 0 | 0.091156
|  | 10 |  | 10 | 0 | 0.14675
|  | 20 |  | 20 | 0 | 0.2752
|  | 50 |  | 50 | 0 | 0.67025
|  | 75 |  | 75 | 0 | 1.0011
|  | 100 |  | 100 | 0 | 1.3334
|  | 200 |  | 200 | 0 | 2.7122
|  | 300 |  | 300 | 0 | 4.0465
|  | 400 |  | 400 | 0 | 5.3857
|  | 500 |  | 500 | 0 | 6.7301
|  | 600 |  | 600 | 0 | 8.0603
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Universal Length Measuring
Nominal Value | Uncertainty of GB |  | Contact Deformation |  | Temperature Effect |  | Resolution of UUC |  | Repeatability |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 0 | 0 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.1229e-05 | 0 | ∞ | 2 | 0.062458
1 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 1.15e-05 | 6.6395e-06 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.1927e-05 | 0 | ∞ | 2 | 0.063854
1.005 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 1.1557e-05 | 6.6727e-06 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.1934e-05 | 0 | ∞ | 2 | 0.063868
1.01 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 1.1615e-05 | 6.7059e-06 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.1941e-05 | 0 | ∞ | 2 | 0.063882
1.05 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 1.2075e-05 | 6.9715e-06 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.1998e-05 | 0 | ∞ | 2 | 0.063996
1.1 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 1.265e-05 | 7.3035e-06 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.2072e-05 | 0 | ∞ | 2 | 0.064144
2 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 2.3e-05 | 1.3279e-05 | 5e-06 | 2.8868e-06 | 0 | 0 | 3.3935e-05 | 0 | ∞ | 2 | 0.06787
5 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 5.75e-05 | 3.3198e-05 | 5e-06 | 2.8868e-06 | 0 | 0 | 4.5578e-05 | 0 | ∞ | 2 | 0.091156
10 | 6e-05 | 3e-05 | 1.417e-05 | 8.1808e-06 | 0.000115 | 6.6395e-05 | 5e-06 | 2.8868e-06 | 0 | 0 | 7.3373e-05 | 0 | ∞ | 2 | 0.14675
20 | 7e-05 | 3.5e-05 | 1.417e-05 | 8.1808e-06 | 0.00023 | 0.00013279 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.0001376 | 0 | ∞ | 2 | 0.2752
50 | 9e-05 | 4.5e-05 | 1.417e-05 | 8.1808e-06 | 0.000575 | 0.00033198 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.00033512 | 0 | ∞ | 2 | 0.67025
75 | 0.0001 | 5e-05 | 1.417e-05 | 8.1808e-06 | 0.0008625 | 0.00049796 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.00050054 | 0 | ∞ | 2 | 1.0011
100 | 0.00012 | 6e-05 | 1.417e-05 | 8.1808e-06 | 0.00115 | 0.00066395 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.00066671 | 0 | ∞ | 2 | 1.3334
200 | 0.00055 | 0.000275 | 1.417e-05 | 8.1808e-06 | 0.0023 | 0.0013279 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.0013561 | 0 | ∞ | 2 | 2.7122
300 | 0.00071 | 0.000355 | 1.417e-05 | 8.1808e-06 | 0.00345 | 0.0019919 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.0020233 | 0 | ∞ | 2 | 4.0465
400 | 0.00089 | 0.000445 | 1.417e-05 | 8.1808e-06 | 0.0046 | 0.0026558 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.0026928 | 0 | ∞ | 2 | 5.3857
500 | 0.0011 | 0.00055 | 1.417e-05 | 8.1808e-06 | 0.00575 | 0.0033198 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.003365 | 0 | ∞ | 2 | 6.7301
600 | 0.00122 | 0.00061 | 1.417e-05 | 8.1808e-06 | 0.0069 | 0.0039837 | 5e-06 | 2.8868e-06 | 0 | 0 | 0.0040302 | 0 | ∞ | 2 | 8.0603
Contact Deformation
|  |  | l1 | 1.1459e-12
|  |  | l2 | 1.1459e-12
|  |  | d | 2.8339e-08 |  | m | A2 | 2.8339e-05 | mm
Note on D4: Certificate of Calubration Gauge Block
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = √3
Note on G6: Divisor = √3
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = 1

## Cert of STD
Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block)
SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[5 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[5 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[5 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[5 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[5 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[5 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[5 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[5 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[5 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[5 empty] | 14 | 0.07 | µm | 7e-05 | mm
[5 empty] | 15 | 0.07 | µm | 7e-05 | mm
[5 empty] | 16 | 0.07 | µm | 7e-05 | mm
[5 empty] | 17 | 0.07 | µm | 7e-05 | mm
[5 empty] | 18 | 0.07 | µm | 7e-05 | mm
[5 empty] | 19 | 0.07 | µm | 7e-05 | mm
[5 empty] | 20 | 0.07 | µm | 7e-05 | mm
[5 empty] | 21 | 0.07 | µm | 7e-05 | mm
[5 empty] | 22 | 0.07 | µm | 7e-05 | mm
[5 empty] | 23 | 0.07 | µm | 7e-05 | mm
[5 empty] | 24 | 0.07 | µm | 7e-05 | mm
[5 empty] | 25 | 0.07 | µm | 7e-05 | mm
[5 empty] | 50 | 0.09 | µm | 9e-05 | mm
[5 empty] | 75 | 0.1 | µm | 0.0001 | mm
[5 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on B4: Grade 0
Note on H4: Grade 0
Note on N4: Grade 0
Note on T4: Grade 0
Note on C5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on E5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on K5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน