## Data
SP METROLOGY SYSTEM THAILAND | [10 empty] | Certificate No. : |  |  |  |  | SPR16010011-1 | [9 empty] | Page |  | 1 | of | 2
[11 empty] | Receive Date : |  |  |  | 2016-01-01 |  |  |  | Calibration Date : |  |  |  |  | 2016-01-02
Equipment Name | [10 empty] | Temp & Humidity : | [5 empty] | 20 |  | oC | 50 |  | %RH
External Micrometer | [10 empty] | Location |  |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SCG
Equipment Name : |  |  |  |  | External Micrometer | [7 empty] | Manufacturer : |  |  |  | Mitutoyo |  |  |  |  | Model :
Serial No. : |  |  | 123-456 |  |  |  |  | ID No:
Range : |  |  | 0 |  | to | 25 |  | mm |  |  | Resolution : |  |  | 0.001 |  | mm
Overall Inspection | [6 empty] | Good |  |  |  | Not Good
Referance Standard : | [16 empty] | Due Date :
1. Flatness Measurement
| Measuring Positions | [6 empty] | UUC Reading
| Anvil | [6 empty] | 0.3
| Spindle | [6 empty] | 0
2. PARALLELISM OF MEASURING FACES
2.1 External Micrometer of length 0-50 mm.
| Size of Optical |  |  |  |  | Number of | [5 empty] | Paralllelism |  |  |  |  | Max |  |  |  | P = (l / 2) x (nAnvil + nSpindle)
| Paralllelism |  |  |  |  | Fringes (n)
| (mm) |  |  |  |  | Anvil |  |  | Spindle |  |  | (mm) | [8 empty] | P = Parallelism of measuring faces
| 12 |  |  |  |  | 1 |  |  | 2 |  |  |  | ≤ | 0.9 | [6 empty] | l = Wavelength
| 12.12 |  |  |  |  | 1 |  |  | 2 |  |  |  | ≤ | 0.9 |  |  | ≤ | 0.9 |  |  | nAnvil = Number of fringes (Anvil)
| 12.25 |  |  |  |  | 1 |  |  | 2 |  |  |  | ≤ | 0.9 | [6 empty] | nSpindle = Number of fringes (Spindle)
| 12.37 |  |  |  |  | 1 |  |  | 2 |  |  |  | ≤ | 0.9
2.2 External Micrometer of length 50-175 mm.
| Size of Optical |  |  |  |  | Size of Gauge |  |  |  |  | Number of | [5 empty] | Paralllelism |  |  |  |  | Max
| Paralllelism |  |  |  |  | Block |  |  |  |  | Fringes (n)
| (mm) |  |  |  |  | (mm) |  |  |  |  | Anvil |  |  | Spindle |  |  | (mm)
| 12.00 - 12.12 | [9 empty] | 1 |  |  | 2 |  |  | 0.9
| 12.00 - 12.25 | [9 empty] | 1 |  |  | 2 |  |  | 0.9 | [5 empty] | ≤ | 0.9
| 12.00 - 12.37 | [9 empty] | 1 |  |  | 2 |  |  | 0.9
[23 empty] | Page |  | 2 | of | 2
3. Instrument Error
| Normal Value |  |  | UUC Reading | [11 empty] | Average |  |  | Repeatability |  |  |  | Error
|  |  |  | X1 |  |  | X2 |  |  | X3 |  |  | X4
| 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0
| 2.5 |  |  | 2.5 |  |  | 2.5 |  |  | 2.5001 |  |  | 2.5 |  |  | 2.5 |  |  | 2.5e-05 |  |  |  | 2.5e-05
| 5.1 |  |  | 5.1 |  |  | 5.1 |  |  | 5.1 |  |  | 5.1 |  |  | 5.1 |  |  | 0 |  |  |  | 0
| 7.7 |  |  | 7.7 |  |  | 7.7 |  |  | 7.7 |  |  | 7.7 |  |  | 7.7 |  |  | 0 |  |  |  | 0
| 10.3 |  |  | 10.3 |  |  | 10.3 |  |  | 10.3 |  |  | 10.3 |  |  | 10.3 |  |  | 0 |  |  |  | 0
| 12.9 |  |  | 12.9 |  |  | 12.9 |  |  | 12.9 |  |  | 12.9 |  |  | 12.9 |  |  | 0 |  |  |  | 0
| 15 |  |  | 15 |  |  | 15 |  |  | 15 |  |  | 15 |  |  | 15 |  |  | 0 |  |  |  | 0
| 17.6 |  |  | 17.6 |  |  | 17.6 |  |  | 17.6 |  |  | 17.6 |  |  | 17.6 |  |  | 0 |  |  |  | 0
| 20.2 |  |  | 20.2 |  |  | 20.2 |  |  | 20.2 |  |  | 20.2 |  |  | 20.2 |  |  | 0 |  |  |  | 0
| 22.8 |  |  | 22.8 |  |  | 22.8 |  |  | 22.8 |  |  | 22.8 |  |  | 22.8 |  |  | 0 |  |  |  | 0
| 25 |  |  | 25 |  |  | 25 |  |  | 25 |  |  | 25 |  |  | 25 |  |  | 0 |  |  |  | 0
Calibrated By : |  |  |  |  | Ms. Arunkamon Raramanus
| 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16010011-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | External Micrometer
| Manufacturer |  |  |  | : | Mitutoyo
| Model |  |  |  | : | 123-456
| Serial Number |  |  |  | : | 0
| ID. Number |  |  |  | : | 0
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  |  | : | 2016-01-01
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-01-02
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2017-01-01
| Calibration Procedure |  |  |  |  | : | SP-CPD-04-01
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-03 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
[10 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16010011-1 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  | SPR16010011-1 | [10 empty] | Page : 3 of 3
| Scale Range |  | : | 0 | to | 25 | mm. |  | Resolution |  | : | 0.001 |  | mm.
Unit : mm
|  | Normal Value |  |  |  | UUC Reading |  |  | Error |  | Uncertainty (± ) µm
|  | 0 |  |  |  | 0 |  |  | 0 |  | 0.57735
|  | 2.5 |  |  |  | 2.5 |  |  | 2.5e-05 |  | 0.58357
|  | 5.1 |  |  |  | 5.1 |  |  | 0 |  | 0.5844
|  | 7.7 |  |  |  | 7.7 |  |  | 0 |  | 0.5894
|  | 10.3 |  |  |  | 10.3 |  |  | 0 |  | 0.59745
|  | 12.9 |  |  |  | 12.9 |  |  | 0 |  | 0.60628
|  | 15 |  |  |  | 15 |  |  | 0 |  | 0.61474
|  | 17.6 |  |  |  | 17.6 |  |  | 0 |  | 0.62678
|  | 20.2 |  |  |  | 20.2 |  |  | 0 |  | 0.64046
|  | 22.8 |  |  |  | 22.8 |  |  | 0 |  | 0.65567
|  | 25 |  |  |  | 25 |  |  | 0 |  | 0.66966
|  | Flatness on measuring faces | [5 empty] | Parallelism between measuring faces
|  | Anvil |  | : | ≤ | 0.3 | µm |  |  | ≤ | 0.9 | mm
|  | Spindle |  | : | ≤ | 0 | µm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget 0 to 25mm
Uncertainty Budget External Micrometer
Nominal Value | Repeatability |  | Uncertainty of STD |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
mm. | mm. |  | mm. |  | mm. |  | mm. | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0005 | 0.00028868 | 0.00028868 | 0.00028868 | ∞ | 2 | 0.57735
2.5 | 2.5e-05 | 2.5e-05 | 6e-05 | 3e-05 | 2.875e-05 | 1.6599e-05 | 0.0005 | 0.00028868 | 0.00029178 | 0.00029178 | 55662.8976 | 2 | 0.58357
5.1 | 0 | 0 | 6e-05 | 3e-05 | 5.865e-05 | 3.3862e-05 | 0.0005 | 0.00028868 | 0.0002922 | 0.0002922 | ∞ | 2 | 0.5844
7.7 | 0 | 0 | 6e-05 | 3e-05 | 8.855e-05 | 5.1124e-05 | 0.0005 | 0.00028868 | 0.0002947 | 0.0002947 | ∞ | 2 | 0.5894
10.3 | 0 | 0 | 7e-05 | 3.5e-05 | 0.00011845 | 6.8387e-05 | 0.0005 | 0.00028868 | 0.00029872 | 0.00029872 | ∞ | 2 | 0.59745
12.9 | 0 | 0 | 7e-05 | 3.5e-05 | 0.00014835 | 8.565e-05 | 0.0005 | 0.00028868 | 0.00030314 | 0.00030314 | ∞ | 2 | 0.60628
15 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0001725 | 9.9593e-05 | 0.0005 | 0.00028868 | 0.00030737 | 0.00030737 | ∞ | 2 | 0.61474
17.6 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0002024 | 0.00011686 | 0.0005 | 0.00028868 | 0.00031339 | 0.00031339 | ∞ | 2 | 0.62678
20.2 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0002323 | 0.00013412 | 0.0005 | 0.00028868 | 0.00032023 | 0.00032023 | ∞ | 2 | 0.64046
22.8 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0002622 | 0.00015138 | 0.0005 | 0.00028868 | 0.00032783 | 0.00032783 | ∞ | 2 | 0.65567
25 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0002875 | 0.00016599 | 0.0005 | 0.00028868 | 0.00033483 | 0.00033483 | ∞ | 2 | 0.66966
Note on E4: Certificate of Calubration ULM
Note on G4: (Normal Value * 11.5*10^-6 * 1)
Note on D6: Divisor = 1
Note on F6: Divisor = 2
Note on H6: Divisor = √3
Note on J6: Divisor = √3

## Uncert of STD
Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09
2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
[5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[10 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[10 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[10 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[10 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[10 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[10 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[10 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[10 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[10 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[10 empty] | 14 | 0.07 | µm | 7e-05 | mm
[10 empty] | 15 | 0.07 | µm | 7e-05 | mm
[10 empty] | 16 | 0.07 | µm | 7e-05 | mm
[10 empty] | 17 | 0.07 | µm | 7e-05 | mm
[10 empty] | 18 | 0.07 | µm | 7e-05 | mm
[10 empty] | 19 | 0.07 | µm | 7e-05 | mm
[10 empty] | 20 | 0.07 | µm | 7e-05 | mm
[10 empty] | 21 | 0.07 | µm | 7e-05 | mm
[10 empty] | 22 | 0.07 | µm | 7e-05 | mm
[10 empty] | 23 | 0.07 | µm | 7e-05 | mm
[10 empty] | 24 | 0.07 | µm | 7e-05 | mm
[10 empty] | 25 | 0.07 | µm | 7e-05 | mm
[10 empty] | 50 | 0.09 | µm | 9e-05 | mm
[10 empty] | 75 | 0.1 | µm | 0.0001 | mm
[10 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on A3: Grade 0
Note on G3: Grade 0
Note on M3: Grade 0
Note on S3: Grade 0
Note on Y3: Grade 0
Note on C4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on O4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on U4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AA4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on D5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on H5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on J5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on N5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on P5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on T5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on V5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on Z5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AB5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน

## Configuration
<?xml version="1.0" encoding="UTF-8"?> <sheets> <sheet name="Data"> <visible>true</visible> <writeable-ranges> <writeable-range>Q2:T2</writeable-range> <writeable-range>Z2:AC2</writeable-range> <writeable-range>S3:T3</writeable-range> <writeable-range>V3:W3</writeable-range> <writeable-range>P8:Q8</writeable-range> <writeable-range>H11:P11</writeable-range> <writeable-range>U11:AD11</writeable-range> <writeable-range>H23:M26</writeable-range> <writeable-range>F42:Q52</writeable-range> <writeable-range>H32:R34</writeable-range> <writeable-range>G55</writeable-range> </writeable-ranges> </sheet> <sheet name="Certificate"> <visible>true</visible> </sheet> <sheet name="Report"> <visible>true</visible> </sheet> <sheet name="Result"> <visible>true</visible> </sheet> <sheet name="Uncertainty Budget 0 to 25mm"> <visible>false</visible> </sheet> <sheet name="Uncert of STD"> <visible>false</visible> </sheet> <sheet name="Configuration"> <visible>false</visible> </sheet> <sheet name="$$^^$$"> <visible>true</visible> </sheet> </sheets>

## $$^^$$
Measurement Result Units: |  |  |  |  | Measurement Range and Parameter:
Lists of Constants
Uncertainty Component Description | Symbol | Estimated Uncertainty | Units | d.f. | Estimated Unc in Measurement Units () | Type (A, B) | Probability Distribution | Divisor | Std Unc () | Relative Contribution (%) | Explanation/Source/Notes | For effective degrees of freedom calculation. |  |  |  | Acceptable Units | Description
[13 empty] | Type | Probability | Divisor
[13 empty] | A | Normal, 1s | 1 |  | blank
[13 empty] | B | Normal, 2s | 2 | µg | microgram
[14 empty] | Normal, 3s | 3 | mg | milligram
[14 empty] | Norm@99% | 2.58 | g | gram
[14 empty] | Rectangular | 1.7321 | kg | kilogram
[14 empty] | Triangular | 2.4495 | µlb | micropound
[14 empty] | U-Shaped | 1.4142 | lb | pound
[14 empty] | Rect 1/2 | 3.4641 | in | inch
[16 empty] | in² | square inch
[16 empty] | in³ | cubic inch
[16 empty] | pt | pint
[16 empty] | qt | quart
[16 empty] | gal | gallon
[16 empty] | mL | milliliter
Instructions: Finish selections or assess resulting values. | [9 empty] | Instructions: Complete data entry. | [5 empty] | L | liter
Min Degrees of Freedom | ν |  |  | Entries incomplete. | [11 empty] | nm | nanometer
Effective Degrees of Freedom | νeff |  |  | TBD | [11 empty] | µm | micrometer
Combined Uncertainty, uc | [8 empty] | 0 | Instructions: Assess data entry and values before reporting rounded result. | [5 empty] | mm | millimeter
Coverage factor, k, uses effective degrees of freedom | [8 empty] | TBD | [6 empty] | cm | centimeter
Expanded Uncertainty, U | [8 empty] | TBD | [6 empty] | cm² | square centimeter
Expanded Uncertainty, U, Rounded to 2 Significant Digits | [8 empty] | TBD | No units selected. | [5 empty] | cm³ | cubic centimeter
[16 empty] | m | meter
[16 empty] | m² | square meter
[16 empty] | m³ | cubic meter
[16 empty] | °F | degrees Fahrenheit
[16 empty] | °C | degrees Celcius
[16 empty] | K | Kelvin
[16 empty] | fl dr | fluid drams
[16 empty] | fl oz | Fluid ounce(s)
[16 empty] | ft | foot
[16 empty] | Pa | Pascal
[16 empty] | kPa | Kilopascal
[16 empty] | s | second
[16 empty] | min | minute
[16 empty] | h | hour
[16 empty] | A | Ampere
[16 empty] | C | coulomb
[16 empty] | V | volt
[16 empty] | J | joule
[16 empty] | W | watt
[16 empty] | Hz | hertz
[16 empty] | H | henry
[16 empty] | F | farad
[16 empty] | Ω | ohm
[16 empty] | S | siemens
[16 empty] | Wb | weber
[16 empty] | T | tesla
[16 empty] | A/m | Ampere per meter (magnetic field strength)
[16 empty] | cd | candela
[16 empty] | lm | lumen
[16 empty] | lx | lux
[16 empty] | cd/m² | candela per square meter
[16 empty] | rad | radian
[16 empty] | kg/m³ | kilogram per cubic meter
[16 empty] | g/cm³ | grams per cubic centimeter
[16 empty] | kg/L | kilograms per liter
[16 empty] | m/s | meters per second
[16 empty] | km/h | kilometers per hour
[16 empty] | m/s² | meter per second squared
[16 empty] | m²/s | meter squared per second
Note on G3: Select Type A or B.
Note on O12: Use if you need to divide the value by 1/2 and the square root of 3.