## Data Record
SP METROLOGY SYSTEM THAILAND | [9 empty] | Certificate No. : |  |  |  | SPR16010011-1 | [7 empty] | Page |  | 1 | of | 2
[10 empty] | Receive Date : |  |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  | 2016-01-02
Equipment Name | [9 empty] | Temp & Humiduty : |  |  |  |  | 20 |  | oC | 50 |  | %RH
External Micrometer | [9 empty] | Location |  |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SCG
Equipment Name : |  |  |  |  | External Micrometer | [8 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  |  | 123-456 | [6 empty] | Serial No. : |  |  | 98778 | [7 empty] | ID No : |  | SP01
Range : |  |  | 125 |  | to | 150 |  | mm |  | Resolution : |  |  | 0.001 |  | mm |  |  |  | Readability :
Overall Inspection | [6 empty] | Good |  |  | Not Good
Referance Standard : | [15 empty] | Due Date :
1. Flatness Measurement
| Measuring Positions | [6 empty] | Reading
| Anvil | [6 empty] | 0.3
| Spindle | [6 empty] | 0
2. PARALLELISM OF MEASURING FACES
2.1 External Micrometer of length 0-50 mm.
| SIZE OF OPTICAL |  |  |  |  | NUMBER OF |  |  | PARALLELISM |  |  |  |  | Max |  |  | P = (l / 2) x (nAnvil + nSpindle)
| PARALLELISM |  |  |  |  | FRINGES (n)
| (mm) |  |  |  |  | ANVIL |  | SPINDLE | (mm) | [7 empty] | P = Parallelism of measuring faces
| 12 |  |  |  |  | 1 |  | 2 |  |  | ≤ | 0.9 |  | ≤ | 0.9 |  | l = Wavelength
| 12.12 |  |  |  |  | 1 |  | 2 |  |  | ≤ | 0.9 |  |  |  |  | nAnvil = Number of fringes (Anvil)
| 12.25 |  |  |  |  | 1 |  | 2 |  |  | ≤ | 0.9 |  |  |  |  | nSpindle = Number of fringes (Spindle)
| 12.37 |  |  |  |  | 1 |  | 2 |  |  | ≤ | 0.9
2.2 External Micrometer of length 50-175 mm.
| SIZE OF OPTICAL |  |  |  |  | SIZE OF GAUGE |  |  | NUMBER OF |  |  |  | PARALLELISM |  |  |  |  | Max
| PARALLELISM |  |  |  |  | BLOCK |  |  | FRINGES (n)
| (mm) |  |  |  |  | (mm) |  |  | ANVIL |  | SPINDLE |  | (mm)
| 12.00 - 12.12 | [7 empty] | 1 |  | 2 |  | 0.9
| 12.00 - 12.25 | [7 empty] | 1 |  | 2 |  | 0.9 | [5 empty] | ≤ | 0.9
| 12.00 - 12.37 | [7 empty] | 1 |  | 2 |  | 0.9
[22 empty] | Page |  | 2 | of | 2
3. Instrument Error
| Norminal Value |  |  | UUC Reading | [10 empty] | Average |  |  | Repeatability |  |  |  | Error
|  |  |  | X1 |  |  | X2 |  | X3 |  |  | X4
| 125 |  |  | 75 |  |  | 75 |  | 75 |  |  | 75 |  |  | 75 |  |  | 0 |  |  |  | -50
| 127.5 |  |  | 77.5 |  |  | 77.5 |  | 77.5 |  |  | 77.5 |  |  | 77.5 |  |  | 0 |  |  |  | -50 |  |  |  |  | 2.5
| 130.1 |  |  | 80.1 |  |  | 80.1 |  | 80.1 |  |  | 80.1 |  |  | 80.1 |  |  | 0 |  |  |  | -50 |  |  |  |  | 5.1
| 132.7 |  |  | 82.7 |  |  | 82.7 |  | 82.7 |  |  | 82.7 |  |  | 82.7 |  |  | 0 |  |  |  | -50 |  |  |  |  | 7.7
| 135.3 |  |  | 85.3 |  |  | 85.3 |  | 85.3 |  |  | 85.3 |  |  | 85.3 |  |  | 0 |  |  |  | -50 |  |  |  |  | 10.3
| 137.9 |  |  | 87.9 |  |  | 87.9 |  | 87.9 |  |  | 87.9 |  |  | 87.9 |  |  | 0 |  |  |  | -50 |  |  |  |  | 12.9
| 140 |  |  | 90 |  |  | 90 |  | 90 |  |  | 90 |  |  | 90 |  |  | 0 |  |  |  | -50 |  |  |  |  | 15
| 142.6 |  |  | 92.6 |  |  | 92.6 |  | 92.6 |  |  | 92.6 |  |  | 92.6 |  |  | 0 |  |  |  | -50 |  |  |  |  | 17.6
| 145.2 |  |  | 95.2 |  |  | 95.2 |  | 95.2 |  |  | 95.2 |  |  | 95.2 |  |  | 0 |  |  |  | -50 |  |  |  |  | 20.2
| 147.8 |  |  | 97.8 |  |  | 97.8 |  | 97.8 |  |  | 97.8 |  |  | 97.8 |  |  | 0 |  |  |  | -50 |  |  |  |  | 22.8
| 150 |  |  | 100 |  |  | 100 |  | 100 |  |  | 100 |  |  | 100 |  |  | 0 |  |  |  | -50 |  |  |  |  | 25
Calibrated By : |  |  |  | Ms. Arunkamon Raramanus
| 10 |  | Mr.Chainarong Matchayamat
| 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16010011-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | External Micrometer
| Manufacturer |  |  |  | : | Mitutoyo
| Model |  |  |  | : | 123-456
| Serial Number |  |  |  | : | 98778
| ID. Number |  |  |  | : | SP01
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-01-01
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2016-01-02
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2017-01-01
| Calibration Procedure |  |  |  | : | SP-CPD-04-01
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-03 |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 10 | Mr.Chainarong Matchayamat | [7 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16010011-1 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  | SPR16010011-1 | [10 empty] | Page : 3 of 3
| Scale Range |  | : | 125 | to | 150 | mm. | Resolution |  | : | 0.001 |  | mm.
Unit : mm
|  | Normal Value |  |  |  | UUC Reading |  |  | Error |  | Uncertainty (± ) µm
|  | 125 |  |  |  | 75 |  |  | -50 |  | 1.8069
|  | 127.5 |  |  |  | 77.5 |  |  | -50 |  | 1.8574
|  | 130.1 |  |  |  | 80.1 |  |  | -50 |  | 1.8916
|  | 132.7 |  |  |  | 82.7 |  |  | -50 |  | 1.9232
|  | 135.3 |  |  |  | 85.3 |  |  | -50 |  | 1.9548
|  | 137.9 |  |  |  | 87.9 |  |  | -50 |  | 1.9866
|  | 140 |  |  |  | 90 |  |  | -50 |  | 2.0149
|  | 142.6 |  |  |  | 92.6 |  |  | -50 |  | 2.0468
|  | 145.2 |  |  |  | 95.2 |  |  | -50 |  | 2.0788
|  | 147.8 |  |  |  | 97.8 |  |  | -50 |  | 2.1109
|  | 150 |  |  |  | 100 |  |  | -50 |  | 2.1264
|  | Flatness on measuring faces | [5 empty] | Parallelism between measuring faces
|  | Anvil : |  | ≤ | 0.3 | µm |  |  |  | ≤ | 0.9 | mm
|  | Spindle : |  | ≤ | 0 | µm
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget 125 to 150mm
Uncertainty Budget External Micrometer
Nominal Value | Repeatability |  | Uncertainty of STD |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
125 | 0 | 0 | 0.00042 | 0.00021 | 0.0014375 | 0.00082994 | 0.0005 | 0.00028868 | 0.00090346 | 0.00090346 | ∞ | 2 | 1.8069
127.5 | 0 | 0 | 0.0005 | 0.00025 | 0.0014663 | 0.00084654 | 0.0005 | 0.00028868 | 0.00092869 | 0.00092869 | ∞ | 2 | 1.8574
130.1 | 0 | 0 | 0.00051 | 0.000255 | 0.0014961 | 0.0008638 | 0.0005 | 0.00028868 | 0.00094579 | 0.00094579 | ∞ | 2 | 1.8916
132.7 | 0 | 0 | 0.00051 | 0.000255 | 0.001526 | 0.00088107 | 0.0005 | 0.00028868 | 0.00096158 | 0.00096158 | ∞ | 2 | 1.9232
135.3 | 0 | 0 | 0.00051 | 0.000255 | 0.001556 | 0.00089833 | 0.0005 | 0.00028868 | 0.00097742 | 0.00097742 | ∞ | 2 | 1.9548
137.9 | 0 | 0 | 0.00051 | 0.000255 | 0.0015858 | 0.00091559 | 0.0005 | 0.00028868 | 0.00099331 | 0.00099331 | ∞ | 2 | 1.9866
140 | 0 | 0 | 0.00052 | 0.00026 | 0.00161 | 0.00092953 | 0.0005 | 0.00028868 | 0.0010075 | 0.0010075 | ∞ | 2 | 2.0149
142.6 | 0 | 0 | 0.00052 | 0.00026 | 0.0016399 | 0.0009468 | 0.0005 | 0.00028868 | 0.0010234 | 0.0010234 | ∞ | 2 | 2.0468
145.2 | 0 | 0 | 0.00052 | 0.00026 | 0.0016698 | 0.00096406 | 0.0005 | 0.00028868 | 0.0010394 | 0.0010394 | ∞ | 2 | 2.0788
147.8 | 0 | 0 | 0.00052 | 0.00026 | 0.0016997 | 0.00098132 | 0.0005 | 0.00028868 | 0.0010554 | 0.0010554 | ∞ | 2 | 2.1109
150 | 0 | 0 | 0.00047 | 0.000235 | 0.001725 | 0.00099593 | 0.0005 | 0.00028868 | 0.0010632 | 0.0010632 | ∞ | 2 | 2.1264
Note on F4: Certificate of Calibration Gauge Block
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on P6: Unit of Uncertainty is µm

## Uncert of STD
Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09
2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
22.8 | 0.07 | µm | 7e-05 | 2.07 | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
[5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[10 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[10 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[10 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[10 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[10 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[10 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[10 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[10 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[10 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[10 empty] | 14 | 0.07 | µm | 7e-05 | mm
[10 empty] | 15 | 0.07 | µm | 7e-05 | mm
[10 empty] | 16 | 0.07 | µm | 7e-05 | mm
[10 empty] | 17 | 0.07 | µm | 7e-05 | mm
[10 empty] | 18 | 0.07 | µm | 7e-05 | mm
[10 empty] | 19 | 0.07 | µm | 7e-05 | mm
[10 empty] | 20 | 0.07 | µm | 7e-05 | mm
[10 empty] | 21 | 0.07 | µm | 7e-05 | mm
[10 empty] | 22 | 0.07 | µm | 7e-05 | mm
[10 empty] | 23 | 0.07 | µm | 7e-05 | mm
[10 empty] | 24 | 0.07 | µm | 7e-05 | mm
[10 empty] | 25 | 0.07 | µm | 7e-05 | mm
[10 empty] | 50 | 0.09 | µm | 9e-05 | mm
[10 empty] | 75 | 0.1 | µm | 0.0001 | mm
[10 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on A3: Grade 0
Note on G3: Grade 0
Note on M3: Grade 0
Note on S3: Grade 0
Note on Y3: Grade 0
Note on C4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on O4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on U4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AA4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on D5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on H5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on J5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on N5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on P5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on T5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on V5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on Z5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AB5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน