## Data
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  | SPR17010147-1 | [5 empty] | Page |  | 1 | of | 1
[9 empty] | Receive Date : |  |  | 2017-01-27 |  |  |  | Calibration Date : |  |  | 2017-01-27
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  | 20 |  | oC | 50 | %RH
Feeler Gauge | [8 empty] | Location |  |  |  | In Lab | [5 empty] | On Site
Customer Name : |  |  |  | CHAROENRUNGRUANG KRITTAPAT AUTOPRESSING CO.,LTD
Equipment Name : |  |  |  | Feeler Gauge | [7 empty] | Manufacturer : |  |  |  | INSIZE
Model : |  | 4605-201 | [5 empty] | Serial No. : |  |  | 0303110531 | [5 empty] | ID No : | QA-QFG-002
Range : | 0.05 |  | to | 1 |  | mm
Overall Inspection | [5 empty] | Good |  |  | Not Good
Referance Standard : |  |  |  |  | SP-SD-OO1 | [8 empty] | Due Date :
Referance Standard : | [13 empty] | Due Date :
Nominal Value |  | Side | Position |  | STD Reading | [9 empty] | Average |  | 𝜎n-1 |  | ¦A-B¦ |  |  | Repeatability | Error | Lateral Wrap (mm) | 𝝈n-1 (max) | uA
[5 empty] | X1 |  | X2 |  |  | X3 |  | X4
0.05 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | -0.0021 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | -0.0023
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | -0.0022
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | -0.0024
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | -0.0024
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | -0.0022
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | -0.002
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | -0.0532
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | -0.0533
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | -0.0531
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | -0.0531
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | -0.0533
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | -0.0532
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | -0.0531
0.1 |  | A | 1 |  | 0.1512 |  | 0.1512 |  |  | 0.1512 |  | 0.1512 |  |  | 0.1512 |  | 0 |  | 0.0522 |  |  | 0.00012051 | -0.0518 | 0.0522 | 0 | 0
|  |  | 2 |  | 0.1518 |  | 0.1518 |  |  | 0.1518 |  | 0.1518 |  |  | 0.1518 |  | 0 | [5 empty] | -0.0515
|  |  | 3 |  | 0.1515 |  | 0.1515 |  |  | 0.1515 |  | 0.1515 |  |  | 0.1515 |  | 0 |  | 0.0514 |  |  |  | -0.0511
|  |  | 4 |  | 0.1511 |  | 0.1511 |  |  | 0.1511 |  | 0.1511 |  |  | 0.1511 |  | 0 | [5 empty] | -0.0514
|  |  | 5 |  | 0.1514 |  | 0.1514 |  |  | 0.1514 |  | 0.1514 |  |  | 0.1514 |  | 0 |  | 0.052 |  |  |  | -0.0516
|  |  | 6 |  | 0.1516 |  | 0.1516 |  |  | 0.1516 |  | 0.1516 |  |  | 0.1516 |  | 0 | [5 empty] | -0.0513
|  |  | 7 |  | 0.1513 |  | 0.1513 |  |  | 0.1513 |  | 0.1513 |  |  | 0.1513 |  | 0 |  | 0.052 |  |  |  | -0.1034
|  | B | 1 |  | 0.2034 |  | 0.2034 |  |  | 0.2034 |  | 0.2034 |  |  | 0.2034 |  | 0 |  |  |  |  | 7.5593e-05 | -0.1032
|  |  | 2 |  | 0.2032 |  | 0.2032 |  |  | 0.2032 |  | 0.2032 |  |  | 0.2032 |  | 0 |  | 0.0519 |  |  |  | -0.1035
|  |  | 3 |  | 0.2035 |  | 0.2035 |  |  | 0.2035 |  | 0.2035 |  |  | 0.2035 |  | 0 | [5 empty] | -0.1031
|  |  | 4 |  | 0.2031 |  | 0.2031 |  |  | 0.2031 |  | 0.2031 |  |  | 0.2031 |  | 0 |  | 0.0519 |  |  |  | -0.1033
|  |  | 5 |  | 0.2033 |  | 0.2033 |  |  | 0.2033 |  | 0.2033 |  |  | 0.2033 |  | 0 | [5 empty] | -0.1035
|  |  | 6 |  | 0.2035 |  | 0.2035 |  |  | 0.2035 |  | 0.2035 |  |  | 0.2035 |  | 0 |  | 0.0521 |  |  |  | -0.1034
|  |  | 7 |  | 0.2034 |  | 0.2034 |  |  | 0.2034 |  | 0.2034 |  |  | 0.2034 |  | 0 | [5 empty] | -0.1522
0.15 |  | A | 1 |  | 0.2522 |  | 0.2522 |  |  | 0.2522 |  | 0.2522 |  |  | 0.2522 |  | 0 |  | 0.0499 |  |  | 8.5912e-05 | -0.1022 | 0.0503 | 0 | 0
|  |  | 2 |  | 0.2523 |  | 0.2523 |  |  | 0.2523 |  | 0.2523 |  |  | 0.2523 |  | 0 | [5 empty] | -0.1023
|  |  | 3 |  | 0.2525 |  | 0.2525 |  |  | 0.2525 |  | 0.2525 |  |  | 0.2525 |  | 0 |  | 0.05 |  |  |  | -0.1025
|  |  | 4 |  | 0.2521 |  | 0.2521 |  |  | 0.2521 |  | 0.2521 |  |  | 0.2521 |  | 0 | [5 empty] | -0.1021
|  |  | 5 |  | 0.2522 |  | 0.2522 |  |  | 0.2522 |  | 0.2522 |  |  | 0.2522 |  | 0 |  | 0.05 |  |  |  | -0.1022
|  |  | 6 |  | 0.252 |  | 0.252 |  |  | 0.252 |  | 0.252 |  |  | 0.252 |  | 0 | [5 empty] | -0.102
|  |  | 7 |  | 0.2524 |  | 0.2524 |  |  | 0.2524 |  | 0.2524 |  |  | 0.2524 |  | 0 |  | 0.0501 |  |  |  | -0.1024
|  | B | 1 |  | 0.3021 |  | 0.3021 |  |  | 0.3021 |  | 0.3021 |  |  | 0.3021 |  | 0 |  |  |  |  | 6.7259e-05 | -0.1521
|  |  | 2 |  | 0.3023 |  | 0.3023 |  |  | 0.3023 |  | 0.3023 |  |  | 0.3023 |  | 0 |  | 0.0502 |  |  |  | -0.1523
|  |  | 3 |  | 0.3025 |  | 0.3025 |  |  | 0.3025 |  | 0.3025 |  |  | 0.3025 |  | 0 | [5 empty] | -0.1525
|  |  | 4 |  | 0.3022 |  | 0.3022 |  |  | 0.3022 |  | 0.3022 |  |  | 0.3022 |  | 0 |  | 0.0503 |  |  |  | -0.1522
|  |  | 5 |  | 0.3024 |  | 0.3024 |  |  | 0.3024 |  | 0.3024 |  |  | 0.3024 |  | 0 | [5 empty] | -0.1524
|  |  | 6 |  | 0.3023 |  | 0.3023 |  |  | 0.3023 |  | 0.3023 |  |  | 0.3023 |  | 0 |  | 0.0498 |  |  |  | -0.1523
|  |  | 7 |  | 0.3022 |  | 0.3022 |  |  | 0.3022 |  | 0.3022 |  |  | 0.3022 |  | 0 | [5 empty] | -0.1522
0.2 |  | A | 1 |  | 0.3513 |  | 0.3513 |  |  | 0.3513 |  | 0.3513 |  |  | 0.3513 |  | 0 |  | 0.0511 |  |  | 7.5593e-05 | -0.1513 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.3514 |  | 0.3514 |  |  | 0.3514 |  | 0.3514 |  |  | 0.3514 |  | 0 | [5 empty] | -0.1514
|  |  | 3 |  | 0.3512 |  | 0.3512 |  |  | 0.3512 |  | 0.3512 |  |  | 0.3512 |  | 0 |  | 0.051 |  |  |  | -0.1512
|  |  | 4 |  | 0.3515 |  | 0.3515 |  |  | 0.3515 |  | 0.3515 |  |  | 0.3515 |  | 0 | [5 empty] | -0.1515
|  |  | 5 |  | 0.3515 |  | 0.3515 |  |  | 0.3515 |  | 0.3515 |  |  | 0.3515 |  | 0 |  | 0.0511 |  |  |  | -0.1515
|  |  | 6 |  | 0.3514 |  | 0.3514 |  |  | 0.3514 |  | 0.3514 |  |  | 0.3514 |  | 0 | [5 empty] | -0.1514
|  |  | 7 |  | 0.3511 |  | 0.3511 |  |  | 0.3511 |  | 0.3511 |  |  | 0.3511 |  | 0 |  | 0.0508 |  |  |  | -0.1511
|  | B | 1 |  | 0.4024 |  | 0.4024 |  |  | 0.4024 |  | 0.4024 |  |  | 0.4024 |  | 0 |  |  |  |  | 6.2678e-05 | -0.2024
|  |  | 2 |  | 0.4024 |  | 0.4024 |  |  | 0.4024 |  | 0.4024 |  |  | 0.4024 |  | 0 |  | 0.051 |  |  |  | -0.2024
|  |  | 3 |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0 | [5 empty] | -0.2023
|  |  | 4 |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0 |  | 0.0509 |  |  |  | -0.2023
|  |  | 5 |  | 0.4025 |  | 0.4025 |  |  | 0.4025 |  | 0.4025 |  |  | 0.4025 |  | 0 | [5 empty] | -0.2025
|  |  | 6 |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 |  |  | 0.4023 |  | 0 |  | 0.051 |  |  |  | -0.2023
|  |  | 7 |  | 0.4021 |  | 0.4021 |  |  | 0.4021 |  | 0.4021 |  |  | 0.4021 |  | 0 | [5 empty] | -0.2021
0.25 |  | A | 1 |  | 0.4514 |  | 0.4514 |  |  | 0.4514 |  | 0.4514 |  |  | 0.4514 |  | 0 |  | 0.0517 |  |  | 8.9974e-05 | -0.2014 | 0.0523 | 0 | 0
|  |  | 2 |  | 0.4516 |  | 0.4516 |  |  | 0.4516 |  | 0.4516 |  |  | 0.4516 |  | 0 | [5 empty] | -0.2016
|  |  | 3 |  | 0.4513 |  | 0.4513 |  |  | 0.4513 |  | 0.4513 |  |  | 0.4513 |  | 0 |  | 0.0516 |  |  |  | -0.2013
|  |  | 4 |  | 0.4515 |  | 0.4515 |  |  | 0.4515 |  | 0.4515 |  |  | 0.4515 |  | 0 | [5 empty] | -0.2015
|  |  | 5 |  | 0.4515 |  | 0.4515 |  |  | 0.4515 |  | 0.4515 |  |  | 0.4515 |  | 0 |  | 0.0523 |  |  |  | -0.2015
|  |  | 6 |  | 0.4511 |  | 0.4511 |  |  | 0.4511 |  | 0.4511 |  |  | 0.4511 |  | 0 | [5 empty] | -0.2011
|  |  | 7 |  | 0.4512 |  | 0.4512 |  |  | 0.4512 |  | 0.4512 |  |  | 0.4512 |  | 0 |  | 0.0522 |  |  |  | -0.2012
|  | B | 1 |  | 0.5031 |  | 0.5031 |  |  | 0.5031 |  | 0.5031 |  |  | 0.5031 |  | 0 |  |  |  |  | 0.0001215 | -0.2531
|  |  | 2 |  | 0.5032 |  | 0.5032 |  |  | 0.5032 |  | 0.5032 |  |  | 0.5032 |  | 0 |  | 0.0516 |  |  |  | -0.2532
|  |  | 3 |  | 0.5036 |  | 0.5036 |  |  | 0.5036 |  | 0.5036 |  |  | 0.5036 |  | 0 | [5 empty] | -0.2536
|  |  | 4 |  | 0.5037 |  | 0.5037 |  |  | 0.5037 |  | 0.5037 |  |  | 0.5037 |  | 0 |  | 0.0521 |  |  |  | -0.2537
|  |  | 5 |  | 0.5031 |  | 0.5031 |  |  | 0.5031 |  | 0.5031 |  |  | 0.5031 |  | 0 | [5 empty] | -0.2531
|  |  | 6 |  | 0.5032 |  | 0.5032 |  |  | 0.5032 |  | 0.5032 |  |  | 0.5032 |  | 0 |  | 0.0522 |  |  |  | -0.2532
|  |  | 7 |  | 0.5034 |  | 0.5034 |  |  | 0.5034 |  | 0.5034 |  |  | 0.5034 |  | 0 | [5 empty] | -0.2534
0.3 |  | A | 1 |  | 0.5522 |  | 0.5522 |  |  | 0.5522 |  | 0.5522 |  |  | 0.5522 |  | 0 |  | 0.0502 |  |  | 6.7259e-05 | -0.2522 | 0.0504 | 0 | 0
|  |  | 2 |  | 0.5524 |  | 0.5524 |  |  | 0.5524 |  | 0.5524 |  |  | 0.5524 |  | 0 | [5 empty] | -0.2524
|  |  | 3 |  | 0.5523 |  | 0.5523 |  |  | 0.5523 |  | 0.5523 |  |  | 0.5523 |  | 0 |  | 0.0499 |  |  |  | -0.2523
|  |  | 4 |  | 0.5521 |  | 0.5521 |  |  | 0.5521 |  | 0.5521 |  |  | 0.5521 |  | 0 | [5 empty] | -0.2521
|  |  | 5 |  | 0.5525 |  | 0.5525 |  |  | 0.5525 |  | 0.5525 |  |  | 0.5525 |  | 0 |  | 0.05 |  |  |  | -0.2525
|  |  | 6 |  | 0.5523 |  | 0.5523 |  |  | 0.5523 |  | 0.5523 |  |  | 0.5523 |  | 0 | [5 empty] | -0.2523
|  |  | 7 |  | 0.5522 |  | 0.5522 |  |  | 0.5522 |  | 0.5522 |  |  | 0.5522 |  | 0 |  | 0.0504 |  |  |  | -0.2522
|  | B | 1 |  | 0.6024 |  | 0.6024 |  |  | 0.6024 |  | 0.6024 |  |  | 0.6024 |  | 0 |  |  |  |  | 6.455e-05 | -0.3024
|  |  | 2 |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0 |  | 0.0497 |  |  |  | -0.3023
|  |  | 3 |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0 | [5 empty] | -0.3023
|  |  | 4 |  | 0.6025 |  | 0.6025 |  |  | 0.6025 |  | 0.6025 |  |  | 0.6025 |  | 0 |  | 0.0498 |  |  |  | -0.3025
|  |  | 5 |  | 0.6022 |  | 0.6022 |  |  | 0.6022 |  | 0.6022 |  |  | 0.6022 |  | 0 | [5 empty] | -0.3022
|  |  | 6 |  | 0.6021 |  | 0.6021 |  |  | 0.6021 |  | 0.6021 |  |  | 0.6021 |  | 0 |  | 0.0501 |  |  |  | -0.3021
|  |  | 7 |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 |  |  | 0.6023 |  | 0 | [5 empty] | -0.3023
0.35 |  | A | 1 |  | 0.6522 |  | 0.6522 |  |  | 0.6522 |  | 0.6522 |  |  | 0.6522 |  | 0 |  | 0.0493 |  |  | 0.00010408 | -0.3022 | 0.0498 | 0 | 0
|  |  | 2 |  | 0.6523 |  | 0.6523 |  |  | 0.6523 |  | 0.6523 |  |  | 0.6523 |  | 0 | [5 empty] | -0.3023
|  |  | 3 |  | 0.6524 |  | 0.6524 |  |  | 0.6524 |  | 0.6524 |  |  | 0.6524 |  | 0 |  | 0.0498 |  |  |  | -0.3024
|  |  | 4 |  | 0.6528 |  | 0.6528 |  |  | 0.6528 |  | 0.6528 |  |  | 0.6528 |  | 0 | [5 empty] | -0.3028
|  |  | 5 |  | 0.6524 |  | 0.6524 |  |  | 0.6524 |  | 0.6524 |  |  | 0.6524 |  | 0 |  | 0.0492 |  |  |  | -0.3024
|  |  | 6 |  | 0.6522 |  | 0.6522 |  |  | 0.6522 |  | 0.6522 |  |  | 0.6522 |  | 0 | [5 empty] | -0.3022
|  |  | 7 |  | 0.6525 |  | 0.6525 |  |  | 0.6525 |  | 0.6525 |  |  | 0.6525 |  | 0 |  | 0.0485 |  |  |  | -0.3025
|  | B | 1 |  | 0.7015 |  | 0.7015 |  |  | 0.7015 |  | 0.7015 |  |  | 0.7015 |  | 0 |  |  |  |  | 0.00014142 | -0.3515
|  |  | 2 |  | 0.7021 |  | 0.7021 |  |  | 0.7021 |  | 0.7021 |  |  | 0.7021 |  | 0 |  | 0.0489 |  |  |  | -0.3521
|  |  | 3 |  | 0.7016 |  | 0.7016 |  |  | 0.7016 |  | 0.7016 |  |  | 0.7016 |  | 0 | [5 empty] | -0.3516
|  |  | 4 |  | 0.7013 |  | 0.7013 |  |  | 0.7013 |  | 0.7013 |  |  | 0.7013 |  | 0 |  | 0.0494 |  |  |  | -0.3513
|  |  | 5 |  | 0.7013 |  | 0.7013 |  |  | 0.7013 |  | 0.7013 |  |  | 0.7013 |  | 0 | [5 empty] | -0.3513
|  |  | 6 |  | 0.7016 |  | 0.7016 |  |  | 0.7016 |  | 0.7016 |  |  | 0.7016 |  | 0 |  | 0.0493 |  |  |  | -0.3516
|  |  | 7 |  | 0.7018 |  | 0.7018 |  |  | 0.7018 |  | 0.7018 |  |  | 0.7018 |  | 0 | [5 empty] | -0.3518
0.4 |  | A | 1 |  | 0.7513 |  | 0.7513 |  |  | 0.7513 |  | 0.7513 |  |  | 0.7513 |  | 0 |  | 0.051 |  |  | 9.759e-05 | -0.3513 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.7515 |  | 0.7515 |  |  | 0.7515 |  | 0.7515 |  |  | 0.7515 |  | 0 | [5 empty] | -0.3515
|  |  | 3 |  | 0.7512 |  | 0.7512 |  |  | 0.7512 |  | 0.7512 |  |  | 0.7512 |  | 0 |  | 0.0506 |  |  |  | -0.3512
|  |  | 4 |  | 0.7516 |  | 0.7516 |  |  | 0.7516 |  | 0.7516 |  |  | 0.7516 |  | 0 | [5 empty] | -0.3516
|  |  | 5 |  | 0.7511 |  | 0.7511 |  |  | 0.7511 |  | 0.7511 |  |  | 0.7511 |  | 0 |  | 0.051 |  |  |  | -0.3511
|  |  | 6 |  | 0.7516 |  | 0.7516 |  |  | 0.7516 |  | 0.7516 |  |  | 0.7516 |  | 0 | [5 empty] | -0.3516
|  |  | 7 |  | 0.7514 |  | 0.7514 |  |  | 0.7514 |  | 0.7514 |  |  | 0.7514 |  | 0 |  | 0.0504 |  |  |  | -0.3514
|  | B | 1 |  | 0.8023 |  | 0.8023 |  |  | 0.8023 |  | 0.8023 |  |  | 0.8023 |  | 0 |  |  |  |  | 8.0178e-05 | -0.4023
|  |  | 2 |  | 0.8021 |  | 0.8021 |  |  | 0.8021 |  | 0.8021 |  |  | 0.8021 |  | 0 |  | 0.0511 |  |  |  | -0.4021
|  |  | 3 |  | 0.8022 |  | 0.8022 |  |  | 0.8022 |  | 0.8022 |  |  | 0.8022 |  | 0 | [5 empty] | -0.4022
|  |  | 4 |  | 0.802 |  | 0.802 |  |  | 0.802 |  | 0.802 |  |  | 0.802 |  | 0 |  | 0.0509 |  |  |  | -0.402
|  |  | 5 |  | 0.8022 |  | 0.8022 |  |  | 0.8022 |  | 0.8022 |  |  | 0.8022 |  | 0 | [5 empty] | -0.4022
|  |  | 6 |  | 0.8025 |  | 0.8025 |  |  | 0.8025 |  | 0.8025 |  |  | 0.8025 |  | 0 |  | 0.0509 |  |  |  | -0.4025
|  |  | 7 |  | 0.8023 |  | 0.8023 |  |  | 0.8023 |  | 0.8023 |  |  | 0.8023 |  | 0 | [5 empty] | -0.4023
0.45 |  | A | 1 |  | 0.8532 |  | 0.8532 |  |  | 0.8532 |  | 0.8532 |  |  | 0.8532 |  | 0 |  | 0.0481 |  |  | 8.5912e-05 | -0.4032 | 0.0487 | 0 | 0
|  |  | 2 |  | 0.8534 |  | 0.8534 |  |  | 0.8534 |  | 0.8534 |  |  | 0.8534 |  | 0 | [5 empty] | -0.4034
|  |  | 3 |  | 0.8532 |  | 0.8532 |  |  | 0.8532 |  | 0.8532 |  |  | 0.8532 |  | 0 |  | 0.048 |  |  |  | -0.4032
|  |  | 4 |  | 0.8533 |  | 0.8533 |  |  | 0.8533 |  | 0.8533 |  |  | 0.8533 |  | 0 | [5 empty] | -0.4033
|  |  | 5 |  | 0.853 |  | 0.853 |  |  | 0.853 |  | 0.853 |  |  | 0.853 |  | 0 |  | 0.0479 |  |  |  | -0.403
|  |  | 6 |  | 0.8535 |  | 0.8535 |  |  | 0.8535 |  | 0.8535 |  |  | 0.8535 |  | 0 | [5 empty] | -0.4035
|  |  | 7 |  | 0.8531 |  | 0.8531 |  |  | 0.8531 |  | 0.8531 |  |  | 0.8531 |  | 0 |  | 0.0482 |  |  |  | -0.4031
|  | B | 1 |  | 0.9013 |  | 0.9013 |  |  | 0.9013 |  | 0.9013 |  |  | 0.9013 |  | 0 |  |  |  |  | 0.00011852 | -0.4513
|  |  | 2 |  | 0.9014 |  | 0.9014 |  |  | 0.9014 |  | 0.9014 |  |  | 0.9014 |  | 0 |  | 0.0487 |  |  |  | -0.4514
|  |  | 3 |  | 0.9011 |  | 0.9011 |  |  | 0.9011 |  | 0.9011 |  |  | 0.9011 |  | 0 | [5 empty] | -0.4511
|  |  | 4 |  | 0.9015 |  | 0.9015 |  |  | 0.9015 |  | 0.9015 |  |  | 0.9015 |  | 0 |  | 0.0475 |  |  |  | -0.4515
|  |  | 5 |  | 0.9017 |  | 0.9017 |  |  | 0.9017 |  | 0.9017 |  |  | 0.9017 |  | 0 | [5 empty] | -0.4517
|  |  | 6 |  | 0.901 |  | 0.901 |  |  | 0.901 |  | 0.901 |  |  | 0.901 |  | 0 |  | 0.0483 |  |  |  | -0.451
|  |  | 7 |  | 0.9014 |  | 0.9014 |  |  | 0.9014 |  | 0.9014 |  |  | 0.9014 |  | 0 | [5 empty] | -0.4514
0.5 |  | A | 1 |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0 |  | 0.0511 |  |  | 8.0917e-05 | -0.4522 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.952 |  | 0.952 |  |  | 0.952 |  | 0.952 |  |  | 0.952 |  | 0 | [5 empty] | -0.452
|  |  | 3 |  | 0.9524 |  | 0.9524 |  |  | 0.9524 |  | 0.9524 |  |  | 0.9524 |  | 0 |  | 0.0508 |  |  |  | -0.4524
|  |  | 4 |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0 | [5 empty] | -0.4522
|  |  | 5 |  | 0.9525 |  | 0.9525 |  |  | 0.9525 |  | 0.9525 |  |  | 0.9525 |  | 0 |  | 0.0505 |  |  |  | -0.4525
|  |  | 6 |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 |  |  | 0.9522 |  | 0 | [5 empty] | -0.4522
|  |  | 7 |  | 0.9523 |  | 0.9523 |  |  | 0.9523 |  | 0.9523 |  |  | 0.9523 |  | 0 |  | 0.0504 |  |  |  | -0.4523
|  | B | 1 |  | 1.0033 |  | 1.0033 |  |  | 1.0033 |  | 1.0033 |  |  | 1.0033 |  | 0 |  |  |  |  | 0.00015924 | -0.5033
|  |  | 2 |  | 1.0028 |  | 1.0028 |  |  | 1.0028 |  | 1.0028 |  |  | 1.0028 |  | 0 |  | 0.0499 |  |  |  | -0.5028
|  |  | 3 |  | 1.0029 |  | 1.0029 |  |  | 1.0029 |  | 1.0029 |  |  | 1.0029 |  | 0 | [5 empty] | -0.5029
|  |  | 4 |  | 1.0026 |  | 1.0026 |  |  | 1.0026 |  | 1.0026 |  |  | 1.0026 |  | 0 |  | 0.0508 |  |  |  | -0.5026
|  |  | 5 |  | 1.0024 |  | 1.0024 |  |  | 1.0024 |  | 1.0024 |  |  | 1.0024 |  | 0 | [5 empty] | -0.5024
|  |  | 6 |  | 1.003 |  | 1.003 |  |  | 1.003 |  | 1.003 |  |  | 1.003 |  | 0 |  | 0.0509 |  |  |  | -0.503
|  |  | 7 |  | 1.0032 |  | 1.0032 |  |  | 1.0032 |  | 1.0032 |  |  | 1.0032 |  | 0 | [5 empty] | -0.5032
0.55 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.4979 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.4977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.4978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.4976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.4976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.4978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.498
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.4468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.4467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.4469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.4469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.4467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.4468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.4469
0.6 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.5479 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.5477
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.5478
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.5476
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.5476
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.5478
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.548
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.4968
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.4967
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.4969
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.4969
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.4967
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.4968
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.4969
0.65 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.5979 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.5977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.5978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.5976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.5976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.5978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.598
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.5468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.5467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.5469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.5469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.5467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.5468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.5469
0.7 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.6479 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.6477
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.6478
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.6476
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.6476
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.6478
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.648
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.5968
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.5967
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.5969
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.5969
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.5967
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.5968
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.5969
0.75 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.6979 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.6977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.6978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.6976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.6976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.6978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.698
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.6468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.6467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.6469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.6469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.6467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.6468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.6469
0.8 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.7479 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.7477
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.7478
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.7476
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.7476
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.7478
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.748
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.6968
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.6967
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.6969
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.6969
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.6967
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.6968
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.6969
0.85 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.7979 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.7977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.7978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.7976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.7976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.7978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.798
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.7468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.7467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.7469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.7469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.7467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.7468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.7469
0.9 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.8479 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.8477
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.8478
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.8476
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.8476
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.8478
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.848
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.7968
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.7967
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.7969
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.7969
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.7967
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.7968
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.7969
0.95 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.8979 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.8977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.8978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.8976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.8976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.8978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.898
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.8468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.8467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.8469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.8469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.8467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.8468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.8469
1 |  | A | 1 |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 |  |  | 0.0521 |  | 0 |  | 0.0511 |  |  | 7.4801e-05 | 0.9479 | 0.0511 | 0 | 0
|  |  | 2 |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 |  |  | 0.0523 |  | 0 | [5 empty] | 0.0977
|  |  | 3 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 |  | 0.051 |  |  |  | 0.0978
|  |  | 4 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 | [5 empty] | 0.0976
|  |  | 5 |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 |  |  | 0.0524 |  | 0 |  | 0.0509 |  |  |  | 0.0976
|  |  | 6 |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 |  |  | 0.0522 |  | 0 | [5 empty] | 0.0978
|  |  | 7 |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0.052 |  |  | 0.052 |  | 0 |  | 0.0507 |  |  |  | 0.098
|  | B | 1 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  |  |  |  | 4.4987e-05 | 0.0468
|  |  | 2 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 |  | 0.0509 |  |  |  | 0.0467
|  |  | 3 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.0469
|  |  | 4 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 |  | 0.051 |  |  |  | 0.0469
|  |  | 5 |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 |  |  | 0.1033 |  | 0 | [5 empty] | 0.0467
|  |  | 6 |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 |  |  | 0.1032 |  | 0 |  | 0.0511 |  |  |  | 0.0468
|  |  | 7 |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 |  |  | 0.1031 |  | 0 | [5 empty] | 0.0469

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR17010147-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Feeler Gauge
| Manufacturer |  |  |  | : | INSIZE
| Model |  |  |  | : | 4605-201
| Serial Number |  |  |  | : | 0303110531
| ID. Number |  |  |  | : | QA-QFG-002
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2017-01-27
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2017-01-27
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2018-01-27
| Calibration Procedure |  |  |  | : | SP-CPT-04-09
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2017-01-28 |  |  | Approved by :
| Calibrated by |  | : | 0 |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | #REF! |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Universal Length Measuring |  |  | LMI 01-680 PC | 050021 | 1000959-1 | 2017-04-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  |  | #REF! | [5 empty] | Page : 3 of 4
[12 empty] | Unit : | mm
Normal Value |  |  | STD Reading (Position) | [7 empty] | Error | Uncertainty ( ± ) µm
|  |  | 1 | 2 |  | 3 | 4 | 5 | 6 | 7
0.05 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | -0.0277 | 0.21
0.1 |  |  | 0.1773 | 0.1775 |  | 0.1775 | 0.1771 | 0.17735 | 0.17755 | 0.17735 | -0.077325 | 0.21
0.15 |  |  | 0.27715 | 0.2773 |  | 0.2775 | 0.27715 | 0.2773 | 0.27715 | 0.2773 | -0.12732 | 0.21
0.2 |  |  | 0.37685 | 0.3769 |  | 0.37675 | 0.3769 | 0.377 | 0.37685 | 0.3766 | -0.1768 | 0.21
0.25 |  |  | 0.47725 | 0.4774 |  | 0.47745 | 0.4776 | 0.4773 | 0.47715 | 0.4773 | -0.22737 | 0.21
0.3 |  |  | 0.5773 | 0.57735 |  | 0.5773 | 0.5773 | 0.57735 | 0.5772 | 0.57725 | -0.27727 | 0.21
0.35 |  |  | 0.67685 | 0.6772 |  | 0.677 | 0.67705 | 0.67685 | 0.6769 | 0.67715 | -0.32703 | 0.21
0.4 |  |  | 0.7768 | 0.7768 |  | 0.7767 | 0.7768 | 0.77665 | 0.77705 | 0.77685 | -0.37685 | 0.21
0.45 |  |  | 0.87725 | 0.8774 |  | 0.87715 | 0.8774 | 0.87735 | 0.87725 | 0.87725 | -0.42728 | 0.21
0.5 |  |  | 0.97775 | 0.9774 |  | 0.97765 | 0.9774 | 0.97745 | 0.9776 | 0.97775 | -0.47758 | 0.21
0.55 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.4723 | 0.21
0.6 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.5223 | 0.21
0.65 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.5723 | 0.21
0.7 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.6223 | 0.21
0.75 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.6723 | 0.21
0.8 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.7223 | 0.21
0.85 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.7723 | 0.21
0.9 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.8223 | 0.21
0.95 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.8723 | 0.21
1 |  |  | 0.07765 | 0.0778 |  | 0.07765 | 0.07775 | 0.07785 | 0.0777 | 0.07755 | 0.9223 | 0.21
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Feeler Gauge
Nominal Value |  | Repeatability |  | Uncertainty of ULM |  | Force |  | Temperature Effect |  | Resolution of ULM |  | Uc | Ui | Veff | K95 | U95%
(mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (mm)
Value |  | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0.05 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 5.75e-07 | 3.3198e-07 | 5e-05 | 2.8868e-05 | 0.00020141 | 0 | ∞ | 2 | 0.00040282
0.1 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 1.15e-06 | 6.6395e-07 | 5e-05 | 2.8868e-05 | 0.00020141 | 0 | ∞ | 2 | 0.00040282
0.15 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 1.725e-06 | 9.9593e-07 | 5e-05 | 2.8868e-05 | 0.00020141 | 0 | ∞ | 2 | 0.00040282
0.2 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 2.3e-06 | 1.3279e-06 | 5e-05 | 2.8868e-05 | 0.00020141 | 0 | ∞ | 2 | 0.00040283
0.25 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 2.875e-06 | 1.6599e-06 | 5e-05 | 2.8868e-05 | 0.00020142 | 0 | ∞ | 2 | 0.00040283
0.3 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 3.45e-06 | 1.9919e-06 | 5e-05 | 2.8868e-05 | 0.00020142 | 0 | ∞ | 2 | 0.00040284
0.35 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 4.025e-06 | 2.3238e-06 | 5e-05 | 2.8868e-05 | 0.00020142 | 0 | ∞ | 2 | 0.00040284
0.4 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 4.6e-06 | 2.6558e-06 | 5e-05 | 2.8868e-05 | 0.00020143 | 0 | ∞ | 2 | 0.00040285
0.45 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 5.175e-06 | 2.9878e-06 | 5e-05 | 2.8868e-05 | 0.00020143 | 0 | ∞ | 2 | 0.00040286
0.5 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 5.75e-06 | 3.3198e-06 | 5e-05 | 2.8868e-05 | 0.00020144 | 0 | ∞ | 2 | 0.00040287
0.55 |  | 0 | 0 | 0.00021 | 0.000105 | 0.00029346 | 0.00016943 | 6.325e-06 | 3.6517e-06 | 5e-05 | 2.8868e-05 | 0.00020144 | 0 | ∞ | 2 | 0.00040288
0.6 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 6.9e-06 | 3.9837e-06 | 5e-05 | 2.8868e-05 | 0.00021859 | 0 | ∞ | 2 | 0.00043718
0.65 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 7.475e-06 | 4.3157e-06 | 5e-05 | 2.8868e-05 | 0.0002186 | 0 | ∞ | 2 | 0.00043719
0.7 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 8.05e-06 | 4.6477e-06 | 5e-05 | 2.8868e-05 | 0.0002186 | 0 | ∞ | 2 | 0.0004372 | Hertz's Formula
0.75 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 8.625e-06 | 4.9796e-06 | 5e-05 | 2.8868e-05 | 0.00021861 | 0 | ∞ | 2 | 0.00043722 | Parameter | Value
0.8 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 9.2e-06 | 5.3116e-06 | 5e-05 | 2.8868e-05 | 0.00021862 | 0 | ∞ | 2 | 0.00043723 | F (force) | 2.3
0.85 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 9.775e-06 | 5.6436e-06 | 5e-05 | 2.8868e-05 | 0.00021863 | 0 | ∞ | 2 | 0.00043725 | D (ball diameter) | 0.02
0.9 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 1.035e-05 | 5.9756e-06 | 5e-05 | 2.8868e-05 | 0.00021863 | 0 | ∞ | 2 | 0.00043727 | 𝜈A | 0.28
0.95 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 1.0925e-05 | 6.3076e-06 | 5e-05 | 2.8868e-05 | 0.00021864 | 0 | ∞ | 2 | 0.00043729 | 𝜈B | 0.28
1 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 1.15e-05 | 6.6395e-06 | 5e-05 | 2.8868e-05 | 0.00021865 | 0 | ∞ | 2 | 0.00043731 | E1 | 200000000000
0 |  | 0 | 0 | 0.00027 | 0.000135 | 0.00029346 | 0.00016943 | 0 | 0 | 5e-05 | 2.8868e-05 | 0.00021855 | 0 | ∞ | 2 | 0.00043711 | E2 | 200000000000
Force
[6 empty] | l1 | 1.1459e-12 | [9 empty] | (A2)3 | 2.5273e-20
[6 empty] | l2 | 1.1459e-12 | [9 empty] | A2 | 2.9346e-07
[6 empty] | d | 2.8339e-08 |  | m | A2 | 2.8339e-05 | mm |  |  |  |  | u(A2) | 1.6943e-07 | m
[18 empty] | 0.00016943 | mm
Norminal Value | Lateral Wrap (mm) |  | Sample Standard Deviation (max) (mm) |  | Uncertainty Type A (mm)
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
0
0
0
0
0
C:\Users\pakaw\oxsoft\metrikos\proecho\01_Dimension 1-10\
'/Users/pakawadee/oxsoft/job/metrology/thoth-project/metrikos/proecho/01_Dimension 1-10/02_Caliper/[02_Caliper_ External.xls]Data Record'!F48
Note on F4: Certificate of Calubration ULM
Note on J4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = √3

## Cert of STD
SP-SD-001
Due Date |  |  | 2016-04-28
1 | mm | 0.21 | µm | 0.00021 | mm
1.01 | mm | 0.21 | µm | 0.00021 | mm
1.05 | mm | 0.21 | µm | 0.00021 | mm
1.1 | mm | 0.21 | µm | 0.00021 | mm
2 | mm | 0.21 | µm | 0.00021 | mm
5 | mm | 0.21 | µm | 0.00021 | mm
10 | mm | 0.21 | µm | 0.00021 | mm
20 | mm | 0.23 | µm | 0.00023 | mm
30 | mm | 0.27 | µm | 0.00027 | mm
40 | mm | 0.27 | µm | 0.00027 | mm
50 | mm | 0.27 | µm | 0.00027 | mm
60 | mm | 0.32 | µm | 0.00032 | mm
70 | mm | 0.32 | µm | 0.00032 | mm
80 | mm | 0.39 | µm | 0.00039 | mm
90 | mm | 0.39 | µm | 0.00039 | mm
100 | mm | 0.39 | µm | 0.00039 | mm
Note on B2: Universal Length Measuring
Note on E3: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B4: Normal Value (mm)
Note on D4: มาจากค่าความไม่แน่นอนของการวัด ULM
Note on F4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน

## Data Side-A
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  | SPR17010147-1 | [6 empty] | Page |  | 1 | of | 1
[9 empty] | Receive Date : |  | 2017-01-27 |  |  |  | Calibration Date : |  |  |  | 2017-01-27
Equipment Name | [8 empty] | Temp & Humiduty : |  |  | 20 |  | oC | 50 |  | %RH
Feeler Gauge | [8 empty] | Location |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  |  | CHAROENRUNGRUANG KRITTAPAT AUTOPRESSING CO.,LTD
Equipment Name : |  |  |  | Feeler Gauge | [6 empty] | Manufacturer : |  |  |  | INSIZE
Model : |  | 4605-201 | [5 empty] | Serial No. : |  | 0303110531 | [6 empty] | ID No : | QA-QFG-002
Range : | 0.05 |  | to | 1 |  | mm
Overall Inspection | [5 empty] | Good |  |  | Not Good
Referance Standard : |  |  |  |  | SP-SD-OO1 | [7 empty] | Due Date :
Referance Standard : | [12 empty] | Due Date :
Nominal Value |  | Position |  | STD Reading | [8 empty] | Average |  |  | Repeatability |  |  | Error
|  |  |  | X1 |  |  | X2 |  | X3 | X4
0.05 |  | 1 |  | 0.0521 |  |  | 0.0521 |  | 0.0521 | 0.0521 |  |  | 0.0521 |  |  | 7.4801e-05 |  |  | -0.0021
|  | 2 |  | 0.0523 |  |  | 0.0523 |  | 0.0523 | 0.0523 |  |  | 0.0523 | [5 empty] | -0.0023
|  | 3 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 | 0.0522 |  |  | 0.0522 | [5 empty] | -0.0022
|  | 4 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 | 0.0524 |  |  | 0.0524 | [5 empty] | -0.0024
|  | 5 |  | 0.0524 |  |  | 0.0524 |  | 0.0524 | 0.0524 |  |  | 0.0524 | [5 empty] | -0.0024
|  | 6 |  | 0.0522 |  |  | 0.0522 |  | 0.0522 | 0.0522 |  |  | 0.0522 | [5 empty] | -0.0022
|  | 7 |  | 0.052 |  |  | 0.052 |  | 0.052 | 0.052 |  |  | 0.052 | [5 empty] | -0.002
0.1 |  | 1 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 | 0.1032 |  |  | 0.1032 |  |  | 4.4987e-05 |  |  | -0.0032
|  | 2 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 | 0.1033 |  |  | 0.1033 | [5 empty] | -0.0033
|  | 3 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 | 0.1031 |  |  | 0.1031 | [5 empty] | -0.0031
|  | 4 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 | 0.1031 |  |  | 0.1031 | [5 empty] | -0.0031
|  | 5 |  | 0.1033 |  |  | 0.1033 |  | 0.1033 | 0.1033 |  |  | 0.1033 | [5 empty] | -0.0033
|  | 6 |  | 0.1032 |  |  | 0.1032 |  | 0.1032 | 0.1032 |  |  | 0.1032 | [5 empty] | -0.0032
|  | 7 |  | 0.1031 |  |  | 0.1031 |  | 0.1031 | 0.1031 |  |  | 0.1031 | [5 empty] | -0.0031
0.15 |  | 1 |  | 0.1512 |  |  | 0.1512 |  | 0.1512 | 0.1512 |  |  | 0.1512 |  |  | 0.00012051 |  |  | -0.0012
|  | 2 |  | 0.1518 |  |  | 0.1518 |  | 0.1518 | 0.1518 |  |  | 0.1518 | [5 empty] | -0.0018
|  | 3 |  | 0.1515 |  |  | 0.1515 |  | 0.1515 | 0.1515 |  |  | 0.1515 | [5 empty] | -0.0015
|  | 4 |  | 0.1511 |  |  | 0.1511 |  | 0.1511 | 0.1511 |  |  | 0.1511 | [5 empty] | -0.0011
|  | 5 |  | 0.1514 |  |  | 0.1514 |  | 0.1514 | 0.1514 |  |  | 0.1514 | [5 empty] | -0.0014
|  | 6 |  | 0.1516 |  |  | 0.1516 |  | 0.1516 | 0.1516 |  |  | 0.1516 | [5 empty] | -0.0016
|  | 7 |  | 0.1513 |  |  | 0.1513 |  | 0.1513 | 0.1513 |  |  | 0.1513 | [5 empty] | -0.0013
0.2 |  | 1 |  | 0.2034 |  |  | 0.2034 |  | 0.2034 | 0.2034 |  |  | 0.2034 |  |  | 7.5593e-05 |  |  | -0.0034
|  | 2 |  | 0.2032 |  |  | 0.2032 |  | 0.2032 | 0.2032 |  |  | 0.2032 | [5 empty] | -0.0032
|  | 3 |  | 0.2035 |  |  | 0.2035 |  | 0.2035 | 0.2035 |  |  | 0.2035 | [5 empty] | -0.0035
|  | 4 |  | 0.2031 |  |  | 0.2031 |  | 0.2031 | 0.2031 |  |  | 0.2031 | [5 empty] | -0.0031
|  | 5 |  | 0.2033 |  |  | 0.2033 |  | 0.2033 | 0.2033 |  |  | 0.2033 | [5 empty] | -0.0033
|  | 6 |  | 0.2035 |  |  | 0.2035 |  | 0.2035 | 0.2035 |  |  | 0.2035 | [5 empty] | -0.0035
|  | 7 |  | 0.2034 |  |  | 0.2034 |  | 0.2034 | 0.2034 |  |  | 0.2034 | [5 empty] | -0.0034
0.25 |  | 1 |  | 0.2522 |  |  | 0.2522 |  | 0.2522 | 0.2522 |  |  | 0.2522 |  |  | 8.5912e-05 |  |  | -0.0022
|  | 2 |  | 0.2523 |  |  | 0.2523 |  | 0.2523 | 0.2523 |  |  | 0.2523 | [5 empty] | -0.0023
|  | 3 |  | 0.2525 |  |  | 0.2525 |  | 0.2525 | 0.2525 |  |  | 0.2525 | [5 empty] | -0.0025
|  | 4 |  | 0.2521 |  |  | 0.2521 |  | 0.2521 | 0.2521 |  |  | 0.2521 | [5 empty] | -0.0021
|  | 5 |  | 0.2522 |  |  | 0.2522 |  | 0.2522 | 0.2522 |  |  | 0.2522 | [5 empty] | -0.0022
|  | 6 |  | 0.252 |  |  | 0.252 |  | 0.252 | 0.252 |  |  | 0.252 | [5 empty] | -0.002
|  | 7 |  | 0.2524 |  |  | 0.2524 |  | 0.2524 | 0.2524 |  |  | 0.2524 | [5 empty] | -0.0024
0.3 |  | 1 |  | 0.3021 |  |  | 0.3021 |  | 0.3021 | 0.3021 |  |  | 0.3021 |  |  | 6.7259e-05 |  |  | -0.0021
|  | 2 |  | 0.3023 |  |  | 0.3023 |  | 0.3023 | 0.3023 |  |  | 0.3023 | [5 empty] | -0.0023
|  | 3 |  | 0.3025 |  |  | 0.3025 |  | 0.3025 | 0.3025 |  |  | 0.3025 | [5 empty] | -0.0025
|  | 4 |  | 0.3022 |  |  | 0.3022 |  | 0.3022 | 0.3022 |  |  | 0.3022 | [5 empty] | -0.0022
|  | 5 |  | 0.3024 |  |  | 0.3024 |  | 0.3024 | 0.3024 |  |  | 0.3024 | [5 empty] | -0.0024
|  | 6 |  | 0.3023 |  |  | 0.3023 |  | 0.3023 | 0.3023 |  |  | 0.3023 | [5 empty] | -0.0023
|  | 7 |  | 0.3022 |  |  | 0.3022 |  | 0.3022 | 0.3022 |  |  | 0.3022 | [5 empty] | -0.0022
0.35 |  | 1 |  | 0.3513 |  |  | 0.3513 |  | 0.3513 | 0.3513 |  |  | 0.3513 |  |  | 7.5593e-05 |  |  | -0.0013
|  | 2 |  | 0.3514 |  |  | 0.3514 |  | 0.3514 | 0.3514 |  |  | 0.3514 | [5 empty] | -0.0014
|  | 3 |  | 0.3512 |  |  | 0.3512 |  | 0.3512 | 0.3512 |  |  | 0.3512 | [5 empty] | -0.0012
|  | 4 |  | 0.3515 |  |  | 0.3515 |  | 0.3515 | 0.3515 |  |  | 0.3515 | [5 empty] | -0.0015
|  | 5 |  | 0.3515 |  |  | 0.3515 |  | 0.3515 | 0.3515 |  |  | 0.3515 | [5 empty] | -0.0015
|  | 6 |  | 0.3514 |  |  | 0.3514 |  | 0.3514 | 0.3514 |  |  | 0.3514 | [5 empty] | -0.0014
|  | 7 |  | 0.3511 |  |  | 0.3511 |  | 0.3511 | 0.3511 |  |  | 0.3511 | [5 empty] | -0.0011
0.4 |  | 1 |  | 0.4024 |  |  | 0.4024 |  | 0.4024 | 0.4024 |  |  | 0.4024 |  |  | 6.2678e-05 |  |  | -0.0024
|  | 2 |  | 0.4024 |  |  | 0.4024 |  | 0.4024 | 0.4024 |  |  | 0.4024 | [5 empty] | -0.0024
|  | 3 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 | 0.4023 |  |  | 0.4023 | [5 empty] | -0.0023
|  | 4 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 | 0.4023 |  |  | 0.4023 | [5 empty] | -0.0023
|  | 5 |  | 0.4025 |  |  | 0.4025 |  | 0.4025 | 0.4025 |  |  | 0.4025 | [5 empty] | -0.0025
|  | 6 |  | 0.4023 |  |  | 0.4023 |  | 0.4023 | 0.4023 |  |  | 0.4023 | [5 empty] | -0.0023
|  | 7 |  | 0.4021 |  |  | 0.4021 |  | 0.4021 | 0.4021 |  |  | 0.4021 | [5 empty] | -0.0021
0.45 |  | 1 |  | 0.4514 |  |  | 0.4514 |  | 0.4514 | 0.4514 |  |  | 0.4514 |  |  | 8.9974e-05 |  |  | -0.0014
|  | 2 |  | 0.4516 |  |  | 0.4516 |  | 0.4516 | 0.4516 |  |  | 0.4516 | [5 empty] | -0.0016
|  | 3 |  | 0.4513 |  |  | 0.4513 |  | 0.4513 | 0.4513 |  |  | 0.4513 | [5 empty] | -0.0013
|  | 4 |  | 0.4515 |  |  | 0.4515 |  | 0.4515 | 0.4515 |  |  | 0.4515 | [5 empty] | -0.0015
|  | 5 |  | 0.4515 |  |  | 0.4515 |  | 0.4515 | 0.4515 |  |  | 0.4515 | [5 empty] | -0.0015
|  | 6 |  | 0.4511 |  |  | 0.4511 |  | 0.4511 | 0.4511 |  |  | 0.4511 | [5 empty] | -0.0011
|  | 7 |  | 0.4512 |  |  | 0.4512 |  | 0.4512 | 0.4512 |  |  | 0.4512 | [5 empty] | -0.0012
0.5 |  | 1 |  | 0.5031 |  |  | 0.5031 |  | 0.5031 | 0.5031 |  |  | 0.5031 |  |  | 0.0001215 |  |  | -0.0031
|  | 2 |  | 0.5032 |  |  | 0.5032 |  | 0.5032 | 0.5032 |  |  | 0.5032 | [5 empty] | -0.0032
|  | 3 |  | 0.5036 |  |  | 0.5036 |  | 0.5036 | 0.5036 |  |  | 0.5036 | [5 empty] | -0.0036
|  | 4 |  | 0.5037 |  |  | 0.5037 |  | 0.5037 | 0.5037 |  |  | 0.5037 | [5 empty] | -0.0037
|  | 5 |  | 0.5031 |  |  | 0.5031 |  | 0.5031 | 0.5031 |  |  | 0.5031 | [5 empty] | -0.0031
|  | 6 |  | 0.5032 |  |  | 0.5032 |  | 0.5032 | 0.5032 |  |  | 0.5032 | [5 empty] | -0.0032
|  | 7 |  | 0.5034 |  |  | 0.5034 |  | 0.5034 | 0.5034 |  |  | 0.5034 | [5 empty] | -0.0034
0.55 |  | 1 |  | 0.5522 |  |  | 0.5522 |  | 0.5522 | 0.5522 |  |  | 0.5522 |  |  | 6.7259e-05 |  |  | -0.0022
|  | 2 |  | 0.5524 |  |  | 0.5524 |  | 0.5524 | 0.5524 |  |  | 0.5524 | [5 empty] | -0.0024
|  | 3 |  | 0.5523 |  |  | 0.5523 |  | 0.5523 | 0.5523 |  |  | 0.5523 | [5 empty] | -0.0023
|  | 4 |  | 0.5521 |  |  | 0.5521 |  | 0.5521 | 0.5521 |  |  | 0.5521 | [5 empty] | -0.0021
|  | 5 |  | 0.5525 |  |  | 0.5525 |  | 0.5525 | 0.5525 |  |  | 0.5525 | [5 empty] | -0.0025
|  | 6 |  | 0.5523 |  |  | 0.5523 |  | 0.5523 | 0.5523 |  |  | 0.5523 | [5 empty] | -0.0023
|  | 7 |  | 0.5522 |  |  | 0.5522 |  | 0.5522 | 0.5522 |  |  | 0.5522 | [5 empty] | -0.0022
0.6 |  | 1 |  | 0.6024 |  |  | 0.6024 |  | 0.6024 | 0.6024 |  |  | 0.6024 |  |  | 6.455e-05 |  |  | -0.0024
|  | 2 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 | 0.6023 |  |  | 0.6023 | [5 empty] | -0.0023
|  | 3 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 | 0.6023 |  |  | 0.6023 | [5 empty] | -0.0023
|  | 4 |  | 0.6025 |  |  | 0.6025 |  | 0.6025 | 0.6025 |  |  | 0.6025 | [5 empty] | -0.0025
|  | 5 |  | 0.6022 |  |  | 0.6022 |  | 0.6022 | 0.6022 |  |  | 0.6022 | [5 empty] | -0.0022
|  | 6 |  | 0.6021 |  |  | 0.6021 |  | 0.6021 | 0.6021 |  |  | 0.6021 | [5 empty] | -0.0021
|  | 7 |  | 0.6023 |  |  | 0.6023 |  | 0.6023 | 0.6023 |  |  | 0.6023 | [5 empty] | -0.0023
0.65 |  | 1 |  | 0.6522 |  |  | 0.6522 |  | 0.6522 | 0.6522 |  |  | 0.6522 |  |  | 0.00010408 |  |  | -0.0022
|  | 2 |  | 0.6523 |  |  | 0.6523 |  | 0.6523 | 0.6523 |  |  | 0.6523 | [5 empty] | -0.0023
|  | 3 |  | 0.6524 |  |  | 0.6524 |  | 0.6524 | 0.6524 |  |  | 0.6524 | [5 empty] | -0.0024
|  | 4 |  | 0.6528 |  |  | 0.6528 |  | 0.6528 | 0.6528 |  |  | 0.6528 | [5 empty] | -0.0028
|  | 5 |  | 0.6524 |  |  | 0.6524 |  | 0.6524 | 0.6524 |  |  | 0.6524 | [5 empty] | -0.0024
|  | 6 |  | 0.6522 |  |  | 0.6522 |  | 0.6522 | 0.6522 |  |  | 0.6522 | [5 empty] | -0.0022
|  | 7 |  | 0.6525 |  |  | 0.6525 |  | 0.6525 | 0.6525 |  |  | 0.6525 | [5 empty] | -0.0025
0.7 |  | 1 |  | 0.7015 |  |  | 0.7015 |  | 0.7015 | 0.7015 |  |  | 0.7015 |  |  | 0.00014142 |  |  | -0.0015
|  | 2 |  | 0.7021 |  |  | 0.7021 |  | 0.7021 | 0.7021 |  |  | 0.7021 | [5 empty] | -0.0021
|  | 3 |  | 0.7016 |  |  | 0.7016 |  | 0.7016 | 0.7016 |  |  | 0.7016 | [5 empty] | -0.0016
|  | 4 |  | 0.7013 |  |  | 0.7013 |  | 0.7013 | 0.7013 |  |  | 0.7013 | [5 empty] | -0.0013
|  | 5 |  | 0.7013 |  |  | 0.7013 |  | 0.7013 | 0.7013 |  |  | 0.7013 | [5 empty] | -0.0013
|  | 6 |  | 0.7016 |  |  | 0.7016 |  | 0.7016 | 0.7016 |  |  | 0.7016 | [5 empty] | -0.0016
|  | 7 |  | 0.7018 |  |  | 0.7018 |  | 0.7018 | 0.7018 |  |  | 0.7018 | [5 empty] | -0.0018
0.75 |  | 1 |  | 0.7513 |  |  | 0.7513 |  | 0.7513 | 0.7513 |  |  | 0.7513 |  |  | 9.759e-05 |  |  | -0.0013
|  | 2 |  | 0.7515 |  |  | 0.7515 |  | 0.7515 | 0.7515 |  |  | 0.7515 | [5 empty] | -0.0015
|  | 3 |  | 0.7512 |  |  | 0.7512 |  | 0.7512 | 0.7512 |  |  | 0.7512 | [5 empty] | -0.0012
|  | 4 |  | 0.7516 |  |  | 0.7516 |  | 0.7516 | 0.7516 |  |  | 0.7516 | [5 empty] | -0.0016
|  | 5 |  | 0.7511 |  |  | 0.7511 |  | 0.7511 | 0.7511 |  |  | 0.7511 | [5 empty] | -0.0011
|  | 6 |  | 0.7516 |  |  | 0.7516 |  | 0.7516 | 0.7516 |  |  | 0.7516 | [5 empty] | -0.0016
|  | 7 |  | 0.7514 |  |  | 0.7514 |  | 0.7514 | 0.7514 |  |  | 0.7514 | [5 empty] | -0.0014
0.8 |  | 1 |  | 0.8023 |  |  | 0.8023 |  | 0.8023 | 0.8023 |  |  | 0.8023 |  |  | 8.0178e-05 |  |  | -0.0023
|  | 2 |  | 0.8021 |  |  | 0.8021 |  | 0.8021 | 0.8021 |  |  | 0.8021 | [5 empty] | -0.0021
|  | 3 |  | 0.8022 |  |  | 0.8022 |  | 0.8022 | 0.8022 |  |  | 0.8022 | [5 empty] | -0.0022
|  | 4 |  | 0.802 |  |  | 0.802 |  | 0.802 | 0.802 |  |  | 0.802 | [5 empty] | -0.002
|  | 5 |  | 0.8022 |  |  | 0.8022 |  | 0.8022 | 0.8022 |  |  | 0.8022 | [5 empty] | -0.0022
|  | 6 |  | 0.8025 |  |  | 0.8025 |  | 0.8025 | 0.8025 |  |  | 0.8025 | [5 empty] | -0.0025
|  | 7 |  | 0.8023 |  |  | 0.8023 |  | 0.8023 | 0.8023 |  |  | 0.8023 | [5 empty] | -0.0023
0.85 |  | 1 |  | 0.8532 |  |  | 0.8532 |  | 0.8532 | 0.8532 |  |  | 0.8532 |  |  | 8.5912e-05 |  |  | -0.0032
|  | 2 |  | 0.8534 |  |  | 0.8534 |  | 0.8534 | 0.8534 |  |  | 0.8534 | [5 empty] | -0.0034
|  | 3 |  | 0.8532 |  |  | 0.8532 |  | 0.8532 | 0.8532 |  |  | 0.8532 | [5 empty] | -0.0032
|  | 4 |  | 0.8533 |  |  | 0.8533 |  | 0.8533 | 0.8533 |  |  | 0.8533 | [5 empty] | -0.0033
|  | 5 |  | 0.853 |  |  | 0.853 |  | 0.853 | 0.853 |  |  | 0.853 | [5 empty] | -0.003
|  | 6 |  | 0.8535 |  |  | 0.8535 |  | 0.8535 | 0.8535 |  |  | 0.8535 | [5 empty] | -0.0035
|  | 7 |  | 0.8531 |  |  | 0.8531 |  | 0.8531 | 0.8531 |  |  | 0.8531 | [5 empty] | -0.0031
0.9 |  | 1 |  | 0.9013 |  |  | 0.9013 |  | 0.9013 | 0.9013 |  |  | 0.9013 |  |  | 0.00011852 |  |  | -0.0013
|  | 2 |  | 0.9014 |  |  | 0.9014 |  | 0.9014 | 0.9014 |  |  | 0.9014 | [5 empty] | -0.0014
|  | 3 |  | 0.9011 |  |  | 0.9011 |  | 0.9011 | 0.9011 |  |  | 0.9011 | [5 empty] | -0.0011
|  | 4 |  | 0.9015 |  |  | 0.9015 |  | 0.9015 | 0.9015 |  |  | 0.9015 | [5 empty] | -0.0015
|  | 5 |  | 0.9017 |  |  | 0.9017 |  | 0.9017 | 0.9017 |  |  | 0.9017 | [5 empty] | -0.0017
|  | 6 |  | 0.901 |  |  | 0.901 |  | 0.901 | 0.901 |  |  | 0.901 | [5 empty] | -0.001
|  | 7 |  | 0.9014 |  |  | 0.9014 |  | 0.9014 | 0.9014 |  |  | 0.9014 | [5 empty] | -0.0014
0.95 |  | 1 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 | 0.9522 |  |  | 0.9522 |  |  | 8.0917e-05 |  |  | -0.0022
|  | 2 |  | 0.952 |  |  | 0.952 |  | 0.952 | 0.952 |  |  | 0.952 | [5 empty] | -0.002
|  | 3 |  | 0.9524 |  |  | 0.9524 |  | 0.9524 | 0.9524 |  |  | 0.9524 | [5 empty] | -0.0024
|  | 4 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 | 0.9522 |  |  | 0.9522 | [5 empty] | -0.0022
|  | 5 |  | 0.9525 |  |  | 0.9525 |  | 0.9525 | 0.9525 |  |  | 0.9525 | [5 empty] | -0.0025
|  | 6 |  | 0.9522 |  |  | 0.9522 |  | 0.9522 | 0.9522 |  |  | 0.9522 | [5 empty] | -0.0022
|  | 7 |  | 0.9523 |  |  | 0.9523 |  | 0.9523 | 0.9523 |  |  | 0.9523 | [5 empty] | -0.0023
1 |  | 1 |  | 1.0033 |  |  | 1.0033 |  | 1.0033 | 1.0033 |  |  | 1.0033 |  |  | 0.00015924 |  |  | -0.0033
|  | 2 |  | 1.0028 |  |  | 1.0028 |  | 1.0028 | 1.0028 |  |  | 1.0028 | [5 empty] | -0.0028
|  | 3 |  | 1.0029 |  |  | 1.0029 |  | 1.0029 | 1.0029 |  |  | 1.0029 | [5 empty] | -0.0029
|  | 4 |  | 1.0026 |  |  | 1.0026 |  | 1.0026 | 1.0026 |  |  | 1.0026 | [5 empty] | -0.0026
|  | 5 |  | 1.0024 |  |  | 1.0024 |  | 1.0024 | 1.0024 |  |  | 1.0024 | [5 empty] | -0.0024
|  | 6 |  | 1.003 |  |  | 1.003 |  | 1.003 | 1.003 |  |  | 1.003 | [5 empty] | -0.003
|  | 7 |  | 1.0032 |  |  | 1.0032 |  | 1.0032 | 1.0032 |  |  | 1.0032 | [5 empty] | -0.0032
Calibrated By : |  |  |  | Mr.Santi Thonglor