## Data Record
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR16070122-5 | [10 empty] | Page 1 of 1
[8 empty] | Receive Date : |  |  |  | 2015-12-12 |  |  |  |  | Calibration Date : |  |  |  |  | 2015-12-12
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 25 |  | oC | 50 |  | %RH
Condutivity Meter | [7 empty] | Location |  |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  | SP
Adress |  |  | : | 88/115
Equipment Name : |  |  |  | Conductivity Meter | [7 empty] | Manufacturer : |  |  |  | Digicon
Model : |  | CD-01 | [5 empty] | Serial No. : |  |  | 12345678 | [7 empty] | ID No : |  | C-01
Range : |  |  |  | to | [5 empty] | Resolution : | 0.01 | [6 empty] | Unit |  | µS/cm
Overall Inspection | [5 empty] | Good |  | Not Good
Reference Standard : |  |  |  |  | SP-SC-004 | [8 empty] | Due Date : |  | 2018-01-27
Reference Standard : |  |  |  |  | SP-SC-005 | [8 empty] | Due Date : |  | 2019-12-11
Reference Standard : |  |  |  |  | SP-SC-006 | [8 empty] | Due Date : |  | 2019-03-03
Calibration Procedure : |  |  |  |  | SP-CPC-04-02
Measurement Result
STD Buffer Solution |  | UUC Reading
Before Adjustment
|  | X1 |  |  | X2 |  | X3 |  | X4 |  |  | Average |  |  | Error |  |  | STDV |  |  |  | Repeatability
84 |  | 84.01 |  |  | 84.01 |  | 84.01 |  | 84.01 |  |  | 84.01 |  |  | 0.01 |  |  | 0 |  |  |  | 0
1413 |  | 1413 |  |  | 1413 |  | 1413 |  | 1413 |  |  | 1413 |  |  | 0 |  |  | 0 |  |  |  | 0
12880 |  | 12880 |  |  | 12880 |  | 12880 |  | 12880 |  |  | 12880 |  |  | 0 |  |  | 0 |  |  |  | 0
STD Buffer Solution |  | UUC Reading
Affter Adjustment
|  | X1 |  |  | X2 |  | X3 |  | X4 |  |  | Average |  |  | Error |  |  | STDV |  |  |  | Repeatability
84
1413
12880
| Calibrated By |  |  |  | : | Mr.Nirut Loha
Mr.Nirut Loha
Mr.Pakapon Nammontree
Mr.Prayoon Topart
Mr.Santi Thonghlor
Mr.Werayut Jampol
Note on Z21: Uncert budget Type A Repeatability
Note on Z28: Uncert budget Type A Repeatability

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR16070122-5 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | SP
88/115
| Equipment Name |  |  | : | Conductivity Meter
| Manufacturer |  |  | : | Digicon
| Model |  |  | : | CD-01
| Serial Number |  |  | : | 12345678
| ID. Number |  |  | : | C-01
Environmental Conditions
| Ambient Temperature |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  | : | 2015-12-12
| Relative Humidity |  |  |  | : | 50% ± 15 % |  | Calibration Date |  | : | 2015-12-12
| Location of Calibration |  |  |  | : | In-Lab |  | Recommended Due Date |  | : | 2016-12-11
| Calibration Procedure |  |  |  | : | SP-CPC-04-02 |  | Date of Issue |  | : | 2015-12-13
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Calibrated by | : | Mr.Nirut Loha |  |  |  | Approved by :
|  |  | Calibration Officer |  |  |  | 3 |  | (Mr.Santi Hankitudomsuk)
Authorized Signatory
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16070122-5 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Conductivity Solution |  |  | HI7033L | Lot. 9983 | 30C64 | 30 Mar 2019
| Conductivity Solution |  |  | HI7031L | Lot. 9986 | 31C63 | 31 Mar 2021
| Conductivity Solution |  |  | HI7030L | Lot. 9974 | 25C62 | 25 Mar-2021
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Standards and Technology, NIST.

## Result
Result of Calibration
| Certificate No. : |  | SPR16070122-5 |  |  |  | Page : 3 of 3
[6 empty] | Unit : | µS/cm
|  | Standard Solution |  | UUC Reading | Error | Uncertainty ( ± )
|  | 84 |  | 84.01 | 0.01 | 1
|  | 1413 |  | 1413 | 0 | 5
|  | 12880 |  | 12880 | 0 | 50
*กรณีการปรับแต่งค่าการวัดของเครื่องมือ ให้ขยับ Measyrement Uncertainty ลงด้านล่าง
| Measurement Uncertainty | [6 empty] | แล้ว Copy ตารางรายงานผลการสอบเทียบมาวางต่อ พร้อมทั้งระบุข้อความสำหรับการปรับแต่งเครื่องมือ
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 % | [7 empty] | Measurement Result (Adjust Calibration)
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Conductivity Meter
[10 empty] | Unit : | µS/cm
Nominal | Uncertainty |  | Resolution of |  | Repeatability |  | Uc | Ui | Veff | K95 | U95%
| Solution |  | UUC |  | of UUC
µS/cm | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
84 | 1 | 0.5 | 0.005 | 0.0028868 | 0 | 0 | 0.50001 | 0 | ∞ | 2 | 1
1413 | 5 | 2.5 | 0.005 | 0.0028868 | 0 | 0 | 2.5 | 0 | ∞ | 2 | 5
12880 | 50 | 25 | 0.005 | 0.0028868 | 0 | 0 | 25 | 0 | ∞ | 2 | 50
Note on C6: ค่าจากใบรายงานผลการสอบเทียบ STD Conductivity Solution
Note on E6: ค่าความละเอียดในการอ่านของ Conductivity Meter
Note on G6: ได้จากการวัดซ้ำ

## Cert of STD
Certificate of Calibration |  | Certificate of Calibration |  | Certificate of Calibration
SP-SC-004 |  | SP-SC-005 |  | SP-SC-006
Conductivity |  | Conductivity |  | Conductivity
Due Date | Jan' 2018 | Due Date | Dec' 2019 | Due Date | Mar' 2019
84 | 1 | 1413 | 5 | 12880 | 50
Note on D8: มาจากค่าความไม่แน่นอนของการวัด Conductivity 84 µS/cm
Note on H8: มาจากค่าความไม่แน่นอนของการวัด Conductivity 1413 µS/cm
Note on L8: มาจากค่าความไม่แน่นอนของการวัด Conductivity 12880 µS/cm