## Data
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR17xxxx | [5 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 2016-03-30 |  |  |  | Calibration Date : |  |  |  | 2016-03-30
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 | %RH
Balance | [7 empty] | Location |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  |  | ABC Ltd.
Adress |  | :
Equipment Name : |  |  |  | Electronic Balance | [6 empty] | Manufacturer : |  |  |  |  | AA
Model : | AB | [6 empty] | Serial No. : |  |  | 1234567 | [5 empty] | ID No : |  | N/A
Range : | 0 |  | to | 400 |  | g |  |  | Resolution : | 0.001 |  |  |  | g
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Reference Standard : |  |  |  |  | SP-SM-026 | [7 empty] | Due Date : |  |  | 2017-11-22
Referaence Standard : | [12 empty] | Due Date :
Calibration Procedure No.: |  |  |  |  | SP-CPM-04-01
Measurement Result
Repeatability ( n = 10 number of measurement ) | [8 empty] | (Max Capacity)
Value | 1 |  |  |  | 2 |  |  | 3 |  |  | 4 |  |  |  | 5 |  |  | STDEV
400 | 400 |  |  |  | 400 |  |  | 400 |  |  | 400 |  |  |  | 400
| 3 |  |  |  | 7 |  |  | 8 |  |  | 9 |  |  |  | 10 |  |  | 0
| 400 |  |  |  | 400 |  |  | 400 |  |  | 400 |  |  |  | 400
Departure of indication from nominal Value | [8 empty] | (Step up to 10%)
Nominal | UUC Reading | [13 empty] | Average |  |  | Repeatability |  |  | Error
g | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
0 | 0 |  |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0
40 | 40 |  |  |  | 40 |  |  | 40 |  |  | 40 |  |  |  | 40 |  |  | 0 |  |  | 0
80 | 80 |  |  |  | 80 |  |  | 80 |  |  | 80 |  |  |  | 80 |  |  | 0 |  |  | 0
120 | 120 |  |  |  | 120 |  |  | 120 |  |  | 120 |  |  |  | 120 |  |  | 0 |  |  | 0
160 | 160 |  |  |  | 160 |  |  | 160 |  |  | 160 |  |  |  | 160 |  |  | 0 |  |  | 0
200 | 200 |  |  |  | 200 |  |  | 200 |  |  | 200 |  |  |  | 200 |  |  | 0 |  |  | 0
240 | 240 |  |  |  | 240 |  |  | 240 |  |  | 240 |  |  |  | 240 |  |  | 0 |  |  | 0
280 | 280 |  |  |  | 280 |  |  | 280 |  |  | 280 |  |  |  | 280 |  |  | 0 |  |  | 0
320 | 320 |  |  |  | 320 |  |  | 320 |  |  | 320 |  |  |  | 320 |  |  | 0 |  |  | 0
360 | 360 |  |  |  | 360 |  |  | 360 |  |  | 360 |  |  |  | 360 |  |  | 0 |  |  | 0
400 | 400 |  |  |  | 400 |  |  | 400 |  |  | 400 |  |  |  | 400 |  |  | 0 |  |  | 0
Off - Center Loading | [6 empty] | (1/4 of Capacity) | [13 empty] | Page 2 of 2
Load test | lndicator
100 | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
Center | 100 |  |  |  | 100 |  |  | 100 |  |  | 100
Front | 100 |  |  |  | 100 |  |  | 100 |  |  | 100
Back | 100 |  |  |  | 100 |  |  | 100 |  |  | 100
Left | 100 |  |  |  | 100 |  |  | 100 |  |  | 100
Right | 100 |  |  |  | 100 |  |  | 100 |  |  | 100
Diff. Error Max. |  | 0
Calibrated By : |  |  | Mr.Nirut Loha
Mr.Nirut Loha
Mr.Pakapon Nammontree
Mr.Prayoon Topart
Mr.Santi Thonghlor
Mr.Werayut Jampol

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR17xxxx |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | ABC Ltd.
0
| Equipment Name |  |  | : | Electronic Balance
| Manufacturer |  |  | : | AA
| Model |  |  | : | AB
| Serial Number |  |  | : | 1234567
| ID. Number |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  | : | 42459
| Relative Humidity |  |  |  | : | 50% ± 15 % |  | Calibration Date |  | : | 42459
| Location of Calibration |  |  |  | : | In-Lab |  | Recommended Due Date |  | : | 42824
| Calibration Procedure |  |  |  | : | SP-CPM-04-01 |  | Date of Issue |  | : | 42460
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
Approved by :
| Calibrated by | : | Mr.Nirut Loha |  |  |  | 3 |  | (Mr. Santi Hankitudomsuk)
Authorized Signatory
Note on Q38: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR17xxxx |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Standard Weight Set |  |  | Class F1 | SS6K3324-16 | CM160770 | 22 Nov 2017
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR17xxxx | [7 empty] | Page : 3 of 3
Repeatability ( n = 10 number of measurement )
|  | Standard Weight |  | Standard Deviation
|  | 400 | g | 0 |  | g
| Departure of indication from nominal Value | [7 empty] | Unit : | g
|  | Nominal Value |  |  | UUC Reading |  | Error | Uncertainty ( ± ) mg |  |  |  | mg
[12 empty] | UUC | Range | CMC
|  | 0 |  |  | 0 |  | 0 | 0.96 |  |  |  | 0.96 | 0 - 400 g | 0.04
|  | 40 |  |  | 40 |  | 0 | 0.96 |  |  |  | 0.96 |  | 0.08 | 3.31
|  | 80 |  |  | 80 |  | 0 | 0.97 |  |  |  | 0.97 |  | 0.16 | 3.31
|  | 120 |  |  | 120 |  | 0 | 0.99 |  |  |  | 0.99 |  | 0.2 | 3.31
|  | 160 |  |  | 160 |  | 0 | 1.01 |  |  |  | 1.01 |  | 0.3 | 3.31
|  | 200 |  |  | 200 |  | 0 | 0.9 |  |  |  | 0.9 |  | 0.3 | 3.31
|  | 240 |  |  | 240 |  | 0 | 0.96 |  |  |  | 0.96 |  | 0.4 | 3.31
|  | 280 |  |  | 280 |  | 0 | 0.97 |  |  |  | 0.97 |  | 0.4 | 3.31
|  | 320 |  |  | 320 |  | 0 | 1 |  |  |  | 0.99 |  | 1 | 3.31
|  | 360 |  |  | 360 |  | 0 | 1.01 |  |  |  | 1.01 |  | 1 | 3.31
|  | 400 |  |  | 400 |  | 0 | 1 |  |  |  | 0.9 |  | 1 | 3.31
Off - Center Loading
|  | Center |  | 100 |  | g
|  | Front |  | 100 |  | g
|  | Back |  | 100 |  | g
|  | Left |  | 100 |  | g
|  | Right |  | 100 |  | g
|  | Maximum difference |  | 0 |  | g
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%
- End of Certificate -

## Uncert Budget
Uncertainty Budget of Electronic Balance
Nominal | Uncertainty |  | Air Buoyancy effect |  | Resolution of |  | Repeatability |  | u(δIecc) |  | u(δItemp) |  | Uc | Ui | Veff | K95 | U95%
| Standard Weight. |  | ± 1 ppm 10^6 |  | UUC |  | of UUC
g | g |  | g |  | g |  | g
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
40 |  | 0 | 7.0131e-07 | 4.049e-07 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00016 | 4.6188e-05 | 0.00029235 | 0 | ∞ | 2 | 0.00058469
80 |  | 0 | 1.4026e-06 | 8.098e-07 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00032 | 9.2376e-05 | 0.0003031 | 0 | ∞ | 2 | 0.00060619
120 |  | 0 | 2.1039e-06 | 1.2147e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00048 | 0.00013856 | 0.00032021 | 0 | ∞ | 2 | 0.00064042
160 |  | 0 | 2.8052e-06 | 1.6196e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00064 | 0.00018475 | 0.00034274 | 0 | ∞ | 2 | 0.00068548
200 |  | 0 | 3.5065e-06 | 2.0245e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.0008 | 0.00023094 | 0.00036969 | 0 | ∞ | 2 | 0.00073938
240 |  | 0 | 4.2078e-06 | 2.4294e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00096 | 0.00027713 | 0.00040017 | 0 | ∞ | 2 | 0.00080035
280 |  | 0 | 4.9091e-06 | 2.8343e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00112 | 0.00032332 | 0.00043345 | 0 | ∞ | 2 | 0.00086689
320 |  | 0 | 5.6104e-06 | 3.2392e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00128 | 0.0003695 | 0.00046891 | 0 | ∞ | 2 | 0.00093782
360 |  | 0 | 6.3118e-06 | 3.6441e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.00144 | 0.00041569 | 0.00050611 | 0 | ∞ | 2 | 0.0010122
400 |  | 0 | 7.0131e-06 | 4.049e-06 | 0.0005 | 0.00028868 | 0 | 0 | 0 | 0 | 0.0016 | 0.00046188 | 0.00054469 | 0 | ∞ | 2 | 0.0010894
Temperature | 23
Humidity | 50
Range | 0-400 g.
ρa | 1.1777
Resolution | 0.001
ปี 2012
Nominal |  | Variation
Length |  | of GB.
( mm ) |  | ( mm )
1 |  | 0.03
1.005 |  | 0.05
1.01 |  | 0.04
1.02 |  | 0.07
1.03 |  | 0.04
1.04 |  | 0.04
1.05 |  | 0.06
1.06 |  | 0.07
1.07 |  | 0.08
1.08 |  | 0.04
1.09 |  | 0.04
1.1 |  | 0.04
1.2 |  | 0.14
1.3 |  | 0.03
1.4 |  | 0.07
1.5 |  | 0.02
1.6 |  | 0.07
1.7 |  | 0.07
1.8 |  | 0.06
1.9 |  | 0.04
2 |  | 0.03
3 |  | 0.08
4 |  | 0.06
5 |  | 0.07
6 |  | 0.05
7 |  | 0.07
8 |  | 0.01
9 |  | 0.07
10 |  | 0.06
11 |  | 0.06
12 |  | 0.02
13 |  | 0.04
14 |  | 0.05
15 |  | 0.05
16 |  | 0.07
17 |  | 0.04
18 |  | 0.05
19 |  | 0.09
20 |  | 0.08
21 |  | 0.06
22 |  | 0.04
23 |  | 0.04
24 |  | 0.13
25 |  | 0.05
50 |  | 0.14
75 |  | 0.14
100 |  | 0.08
125 |  | 0.03
150 |  | 0
175 |  | 0
200 |  | 0
250 |  | 0
300 |  | 0
400 |  | 0
500 |  | 0
600 |  | 0
700 |  | 0
800 |  | 0
900 |  | 0
1000 |  | 0
Note on N3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การวางน้ำหนักไม่ตรงกลาง
Note on P3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การเปลี่ยนแปลงอุณภูมิ
Note on B6: (1ppm*W)/SQRT3
Note on C6: ค่าจากใบรายงานผลการสอบเทียบ STD Weight
Note on H6: (1ppm*W)/SQRT3
Note on J6: ค่าความละเอียดในการอ่านของเครื่องชั่ง
Note on L6: ได้จากการวัดซ้ำ
Note on D7: ใช้ตุ้มน้ำหนัก2ค่ารวมกัน
Note on E7: ใช้ตุ้มน้ำหนัก3ค่ารวมกัน
Note on F7: ใช้ตุ้มน้ำหนัก4ค่ารวมกัน
Note on B21: SP Metrology: The current density of an Air.

## Uncert Budget '2017
Uncertainty Budget of Electronic Balance
Function : | Mass |  |  | Testpoint | 0 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00028868 | 3
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00096 | 3
[5 empty] | 3.31 |  |  |  |  | 3
Function : | Mass |  |  | Testpoint | 40 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 4e-05 |  | Rectangular |  | 1.7321 | 1 |  | 2.3094e-05 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0002896 | 3.0385
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00096 | 3.0385
[5 empty] | 3.31 |  |  |  |  | 3.0385
Function : | Mass |  |  | Testpoint | 80 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 8e-05 |  | Rectangular |  | 1.7321 | 1 |  | 4.6188e-05 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00029235 | 3.1556
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00097 | 3.1556
[5 empty] | 3.31 |  |  |  |  | 3.1556
Function : | Mass |  |  | Testpoint | 120 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00012 |  | Rectangular |  | 1.7321 | 1 |  | 6.9282e-05 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00029687 | 3.3556
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00099 | 3.3556
[5 empty] | 3.31 |  |  |  |  | 3.3556
Function : | Mass |  |  | Testpoint | 160 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00016 |  | Rectangular |  | 1.7321 | 1 |  | 9.2376e-05 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.0003031 | 3.6459
U | Expaned uncertainty |  |  |  | k = | 3.31 |  |  |  | 0.00101 | 3.6459
[5 empty] | 3.31 |  |  |  |  | 3.6459
Function : | Mass |  |  | Testpoint | 200 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.0002 |  | Rectangular |  | 1.7321 | 1 |  | 0.00011547 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00031091 | 4.0368
U | Expaned uncertainty |  |  |  | k = | 2.87 |  |  |  | 0.0009 | 4.0368
[5 empty] | 2.87 |  |  |  |  | 4.0368
Function : | Mass |  |  | Testpoint | 240 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00024 |  | Rectangular |  | 1.7321 | 1 |  | 0.00013856 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00032021 | 4.5417
U | Expaned uncertainty |  |  |  | k = | 2.87 |  |  |  | 0.00092 | 4.5417
[5 empty] | 2.87 |  |  |  |  | 4.5417
Function : | Mass |  |  | Testpoint | 280 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00028 |  | Rectangular |  | 1.7321 | 1 |  | 0.00016166 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00033086 | 5.1766
U | Expaned uncertainty |  |  |  | k = | 2.65 |  |  |  | 0.00088 | 5.1766
[5 empty] | 2.65 |  |  |  |  | 5.1766
Function : | Mass |  |  | Testpoint | 320 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00032 |  | Rectangular |  | 1.7321 | 1 |  | 0.00018475 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00034273 | 5.9609
U | Expaned uncertainty |  |  |  | k = | 2.65 |  |  |  | 0.00091 | 5.9609
[5 empty] | 2.65 |  |  |  |  | 5.9609
Function : | Mass |  |  | Testpoint | 360 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.00036 |  | Rectangular |  | 1.7321 | 1 |  | 0.00020785 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00035572 | 6.9166
U | Expaned uncertainty |  |  |  | k = | 2.52 |  |  |  | 0.0009 | 6.9166
[5 empty] | 2.52 |  |  |  |  | 6.9166
Function : | Mass |  |  | Testpoint | 400 |  | g | Range | 400 | g
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) g | [5 empty] | (+/-) % | (+/-) g
UB1 | Air Buoyancy effect |  | 0.0004 |  | Rectangular |  | 1.7321 | 1 |  | 0.00023094 | ¥
UB2 | Resolution of UUC |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UA1 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 3
Uc | Combined uncertainty |  |  |  | T-Distibution |  |  |  |  | 0.00036968 | 8.0688
U | Expaned uncertainty |  |  |  | k = | 2.37 |  |  |  | 0.00088 | 8.0688
[6 empty] | 2.37 |  |  |  | 8.0688

## Cert of STD
Weight 10 mg - 500 mg |  |  |  | Weight 1g - 5 kg |  |  |  | Weight 10 g - 20 kg |  |  |  | Weight 5 kg - 20 kg
SP-SM-001 |  |  |  | SP-SM-002 |  |  |  | SP-SM-003 |  |  |  | SP-SM-004
Due Date |  |  |  | Due Date |  |  |  | Due Date |  |  |  | No. | Due Date
42550 |  |  |  | 42550 |  |  |  | 42601 |  |  |  |  | 42706
10 | mg | 0.01 | mg | 1 | g | 0.02 | mg | 10 | g | 0.035 | mg | 1 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg | 2 | g* | 0.03 | mg* | 20 | g | 0.042 | mg | 2 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg* | 2 | g | 0.03 | mg | 50 | g | 0.049 | mg | 1 | 10 | kg | 31 | mg
50 | mg | 0.01 | mg | 5 | g | 0.04 | mg | 100 | g | 0.11 | mg | 2 | 10 | kg | 31 | mg
100 | mg | 0.01 | mg | 10 | g | 0.04 | mg | 200 | g | 0.17 | mg | 1 | 20 | kg | 54 | mg
200 | mg | 0.01 | mg | 20 | g | 0.05 | mg | 500 | g | 0.4 | mg | 2 | 20 | kg | 54 | mg
500 | mg | 0.01 | mg | 20 | g* | 0.05 | mg* | 1 | kg | 1.7 | mg | 3 | 20 | kg | 54 | mg
|  |  |  | 50 | g | 0.1 | mg | 2 | kg | 2 | mg | 4 | 20 | kg | 54 | mg
|  |  |  | 100 | g | 0.2 | mg | 10 | kg | 12 | mg | 5 | 20 | kg | 54 | mg
|  |  |  | 200 | g | 0.4 | mg | 20 | kg | 18 | mg | 6 | 20 | kg | 54 | mg
|  |  |  | 200 | g* | 0.4 | mg* |  |  |  |  | 7 | 20 | kg | 54 | mg
|  |  |  | 500 | g | 2 | mg |  |  |  |  | 8 | 20 | kg | 54 | mg
|  |  |  | 1 | kg | 2 | mg |  |  |  |  | 9 | 20 | kg | 54 | mg
|  |  |  | 1 | kg* | 2 | mg |  |  |  |  | 10 | 20 | kg | 54 | mg
|  |  |  | 2 | kg | 4 | mg |  |  |  |  | 11 | 20 | kg | 54 | mg
|  |  |  | 5 | kg | 10 | mg |  |  |  |  | 12 | 20 | kg | 54 | mg
[12 empty] | 13 | 20 | kg | 54 | mg
[12 empty] | 14 | 20 | kg | 54 | mg
[12 empty] | 15 | 20 | kg | 54 | mg
[12 empty] | 16 | 20 | kg | 54 | mg
[12 empty] | 17 | 20 | kg | 54 | mg
[12 empty] | 18 | 20 | kg | 54 | mg
[12 empty] | 19 | 20 | kg | 54 | mg
[12 empty] | 20 | 20 | kg | 54 | mg
[12 empty] | 21 | 20 | kg | 54 | mg
[12 empty] | 22 | 20 | kg | 54 | mg
[12 empty] | 23 | 20 | kg | 54 | mg
[12 empty] | 24 | 20 | kg | 54 | mg
[12 empty] | 25 | 20 | kg | 54 | mg
[12 empty] | 26 | 20 | kg | 54 | mg
[12 empty] | 27 | 20 | kg | 54 | mg
[12 empty] | 28 | 20 | kg | 54 | mg
[12 empty] | 29 | 20 | kg | 54 | mg
[12 empty] | 30 | 20 | kg | 54 | mg
[12 empty] | 31 | 20 | kg | 54 | mg
[12 empty] | 32 | 20 | kg | 54 | mg
[12 empty] | 33 | 20 | kg | 54 | mg
[12 empty] | 34 | 20 | kg | 54 | mg
[12 empty] | 35 | 20 | kg | 54 | mg
[12 empty] | 36 | 20 | kg | 54 | mg
[12 empty] | 37 | 20 | kg | 54 | mg
[12 empty] | 38 | 20 | kg | 54 | mg
[12 empty] | 39 | 20 | kg | 54 | mg
[12 empty] | 40 | 20 | kg | 54 | mg
[12 empty] | 41 | 20 | kg | 54 | mg
[12 empty] | 42 | 20 | kg | 54 | mg
[12 empty] | 43 | 20 | kg | 54 | mg
[12 empty] | 44 | 20 | kg | 54 | mg
[12 empty] | 45 | 20 | kg | 54 | mg
[12 empty] | 46 | 20 | kg | 54 | mg
[12 empty] | 47 | 20 | kg | 54 | mg
[12 empty] | 48 | 20 | kg | 54 | mg
[12 empty] | 49 | 20 | kg | 54 | mg
[12 empty] | 50 | 20 | kg | 54 | mg
Note on B6: MPE Class F1
Note on G6: MPE Class F1
Note on L6: MPE Class E2
Note on Q6: MPE Class M1
Note on B8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B9: Nathaphol Boonmee: Normal Value
Note on D9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G9: Nathaphol Boonmee: Normal Value
Note on I9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L9: Nathaphol Boonmee: Normal Value
Note on N9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R9: Nathaphol Boonmee: Normal Value
Note on T9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง