## Data Record(pitch)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR16010011-1 | [6 empty] | Page |  | 1 |  | of | 1
[9 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  | 20 |  | oC | 50 |  | %RH
Thread Ring Gauge | [8 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  | ABC
Equipment Name : |  |  | Thread Ring Gauge | [8 empty] | Manufacturer : |  |  |  | ISOKU
Model : | 1414 | [7 empty] | Serial No. : |  | S12345 | [7 empty] | ID No : |  | M13
Range : | M15xP1.25 GPII | [5 empty] | M15xP1.25 IPII
Overall Inspection |  |  |  |  | Good |  |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Referance Standard : | [13 empty] | Due Date :
Measurement Result : | [6 empty] | □ | Without Adjustment | [5 empty] | □ | After Adjustment
Symbol of Thread Ring Gauge |  |  | Pitch Diameter | [10 empty] | Minor Diameter
|  |  | Nominal Value (mm) |  |  | Permissible Deviation (μm) | [7 empty] | Nominal Value (mm) |  |  | Permissible Deviation (μm)
[6 empty] | Upper |  |  | Lower |  | Tolerance | [5 empty] | Upper |  |  | Lower |  |  | Tolerance
M15xP1.25 GPII |  |  | 14.232 |  |  | 23 |  |  | 12 |  | 16 |  |  | 14.222 |  |  | 10 |  |  | 11 |  |  | 11
M15xP1.25 IPII |  |  | 14.283 |  |  | 45 |  |  | 10 |  | 35 |  |  | 14.212 |  |  | 10 |  |  | 11 |  |  | 11
1. Measurement Results | [25 empty] | มุมเอียงสันเกลียว ( Thread angle ) : ( ° )
Description : |  |  | M | 10 |  | x | P | 1.25 |  | GPII |  |  |  | Pitch Diameter [E] | [5 empty] | 8.9175 |  |  | mm |  |  | ค่าที่ระบุ | a
Thread angle: |  |  |  | 60 | [9 empty] | 2-Ball Stylus (dD) | [5 empty] | 0.725 |  |  | mm |  |  |  | 29° | 30° | 55° | 60°
[14 empty] | Best Dimeter (do) | [5 empty] | 0.72169 |  |  | mm |  |  |  | 1.9334 | 1.866 | 0.9605 | 0.866
[14 empty] | d๐ = (P/2)(1/COS(α/2) | [12 empty] | 4.9939 | 4.8637 | 3.1657 | 3
Nominal |  | Position |  |  | Standard Reading | [10 empty] | Average |  |  | Pitch Dia [E] |  |  | Repeatability |  |  |  | 2 Ball Stylus of SP
[5 empty] | X1 |  |  | X2 |  | X3 |  |  | X4 | [12 empty] | 0.2 | 1.1
X |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 |  |  | 0 |  |  |  | 0.25 | 1.35
|  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 | [6 empty] | 0.28 | 1.65
|  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 | [6 empty] | 0.455 | 2.05
Y |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 | [6 empty] | 0.53 | 2.55
|  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 | [6 empty] | 0.62 | 3.2
|  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 | [6 empty] | 0.725 | 4
0.895
Description : |  |  | M | 10 |  | x | P | 1.25 |  | IPII |  |  |  | Pitch Diameter [E] | [5 empty] | 8.9175 |  |  | mm |  |  | ค่าที่ระบุ | a
Thread angle: |  |  |  | 60 | [9 empty] | 2-Ball Stylus (dD) | [5 empty] | 1.35 |  |  | mm |  |  |  | 29° | 30° | 55° | 60°
[14 empty] | Best Dimeter (do) | [5 empty] | 0.72169 |  |  | mm |  |  |  | 1.9334 | 1.866 | 0.9605 | 0.866
[14 empty] | d๐ = (P/2)(1/COS(α/2) | [12 empty] | 4.9939 | 4.8637 | 3.1657 | 3
Nominal |  | Position |  |  | Standard Reading | [10 empty] | Average |  |  | Pitch Dia [E] |  |  | Repeatability
[5 empty] | X1 |  |  | X2 |  | X3 |  |  | X4
X |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175 |  |  | 0
|  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175
|  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175
Y |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175
|  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175
|  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 8.9175

## Data Record(minor)
SP METROLOGY SYSTEM THAILAND | [9 empty] | Certificate No. : |  |  | SPR16010011-1 | [7 empty] | Page |  | 1 |  | of | 1
[10 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [9 empty] | Temp & Humiduty : |  |  |  | 20 |  | oC | 50 |  | %RH
Thread Plug Gauge | [9 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  | 0
Equipment Name : |  |  |  | Thread Ring Gauge | [8 empty] | Manufacturer : |  |  |  | ISOKU
Model : |  | 1414 | [7 empty] | Serial No. : |  | S12345 | [7 empty] | ID No : |  | M13
Range : |  | M15xP1.25 GPII | [5 empty] | M15xP1.25 IPII
Overall Inspection | [5 empty] | Good |  |  |  | Not Good
Referance Standard : | [14 empty] | Due Date :
Referance Standard : | [14 empty] | Due Date :
Symbol of Thread Ring Gauge |  |  |  | Pitch Diameter | [10 empty] | Minor Diameter
|  |  |  | Nominal Value (mm) |  |  | Permissible Deviation (μm) | [7 empty] | Nominal Value (mm) |  |  | Permissible Deviation (μm)
[7 empty] | Upper |  |  | Lower |  | Tolerance | [5 empty] | Upper |  |  | Lower |  |  | Tolerance
M15xP1.25 GPII |  |  |  | 14.232 |  |  | 23 |  |  | 12 |  | 16 |  |  | 14.222 |  |  | 10 |  |  | 11 |  |  | 11
M15xP1.25 IPII |  |  |  | 14.283 |  |  | 45 |  |  | 10 |  | 35 |  |  | 14.212 |  |  | 10 |  |  | 11 |  |  | 11
1. Measurement Results
M15xP1.25 GPII
Nominal |  |  | Position |  |  | Standard Reading | [10 empty] | Average |  |  | Repeatability
[6 empty] | X1 |  |  | X2 |  | X3 |  |  | X4
X |  |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 0
|  |  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
Y |  |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
M15xP1.25 IPII
Nominal |  |  | Position |  |  | Standard Reading | [10 empty] | Average |  |  | Repeatability
[6 empty] | X1 |  |  | X2 |  | X3 |  |  | X4
X |  |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10 |  |  | 0
|  |  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
Y |  |  | 1/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 1/2T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
|  |  | 4/5T |  |  | 10 |  |  | 10 |  | 10 |  |  | 10 |  |  | 10
| Calibrated By : |  |  |  | Ms. Arunkamon Raramanus
|  |  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16010011-1 | [6 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Thread Ring Gauge
| Manufacturer |  |  |  | : | ISOKU
| Model |  |  |  | : | 1414
| Serial Number |  |  |  | : | S12345
| ID. Number |  |  |  | : | M13
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 20 °C ± 1 °C |  | Received Date |  |  |  | : | 2016-01-01
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-01-01
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2016-12-31
| Calibration Procedure |  |  |  |  | : | SP-CPT-04-21
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-02 |  |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [8 empty] | 1 | Mr.Sombut Srikampa
[11 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16010011-1 |  |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  |  | Model | Serial No. | Certificate No. | Due. Date
| Universal Length Measuring |  |  |  | LMI 01-680 PC | 050021 | 1000959-1 | 2017-04-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result (Pitch)
Result of Calibration
|  | Certificate Number : |  |  | SPR16010011-1 | [7 empty] | Page : 3 of 4
Symbol of Thread Ring Gauge |  |  | Pitch Diameter | [5 empty] | Major Diameter
|  |  | Nominal Value (mm) |  | Permissible Deviation (μm) |  |  |  | Nominal Value (mm) |  | Permissible Deviation (μm)
[5 empty] | Upper | Lower | Tolerance |  |  |  | Upper | Lower |  | Tolerance
M15xP1.25 GPII |  |  | 14.232 |  | 23 | 12 | 16 |  | 14.222 |  | 10 | 11 |  | 11
M15xP1.25 IPII |  |  | 14.283 |  | 45 | 10 | 35 |  | 14.212 |  | 10 | 11 |  | 11
| Measurement Results Pitch Diameter | [11 empty] | Unit : |  | mm
| Symbol of Thread Ring Gauge |  |  | Positions |  | Measured Values (mm) |  |  |  | Deviation |  | Uncertainty ( ± ) µm
[6 empty] | x-axis |  | y-axis
| M15xP1.25 GPII |  |  | 1/5 T |  | 8.9175 |  | 8.9175 |  | 5.3145 |  | 0.58497
|  |  |  | 1/2 T |  | 8.9175 |  | 8.9175
|  |  |  | 4/5 T |  | 8.9175 |  | 8.9175
| M15xP1.25 IPII |  |  | 1/5 T |  | 8.9175 |  | 8.9175 |  | 5.3655 |  | 0.58497
|  |  |  | 1/2 T |  | 8.9175 |  | 8.9175
|  |  |  | 4/5 T |  | 8.9175 |  | 8.9175

## Result (Minor)
Result of Calibration
|  | Certificate Number : |  | SPR16010011-1 |  |  |  |  | Page : 4 of 4
| Measurement Results Minor Diameter | [7 empty] | Unit : | mm
| Symbol of Thread Plug Gague |  | Positions |  | Measured Values (mm) |  | Deviation | Uncertainty ( ± ) µm
[5 empty] | x-axis | y-axis
| M15xP1.25 GPII |  | 1/5 T |  | 10 | 10 | 4.222 | 0.45118
|  |  | 1/2 T |  | 10 | 10
|  |  | 4/5 T |  | 10 | 10
| M15xP1.25 IPII |  | 1/5 T |  | 10 | 10 | 4.212 | 0.45118
|  |  | 1/2 T |  | 10 | 10
|  |  | 4/5 T |  | 10 | 10
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget(Pitch)
Uncertainty Budget of Thread Ring Gauge
Nominal Value | Repeatability |  | Uncertainty of Rang Gauge |  | Uncertainty of ULM |  | Resolution of ULM |  | Flank Angle |  | Elastic Deformation |  | Temperature Effect |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
8.9175 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00029249 | 0 | ∞ | 2 | 0.58497
8.9175 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00029249 | 0 | ∞ | 2 | 0.58497
8 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 9.2e-05 | 5.3116e-05 | 0.00029131 | 0 | 0.064818 | 2.00 | 0.58263
10 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.000115 | 6.6395e-05 | 0.00029403 | 0 | 0.067264 | 2.00 | 0.58805
12 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.000138 | 7.9674e-05 | 0.00030098 | 0 | 0.073861 | 2.00 | 0.60197
18 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.000207 | 0.00011951 | 0.00031389 | 0 | 0.087366 | 2.00 | 0.62778
20 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.00023 | 0.00013279 | 0.00031918 | 0 | 0.093409 | 2.00 | 0.63836
25 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00027 | 0.000135 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.0002875 | 0.00016599 | 0.00034175 | 0 | 0.12277 | 2.00 | 0.68351
30 | 0 | 0 | 0.00037033 | 0.00018517 | 0.00027 | 0.000135 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.000345 | 0.00019919 | 0.00035905 | 0 | 0.14958 | 2.00 | 0.7181
75 | 0 | 0 | 0.002 | 0.001 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.0008625 | 0.00049796 | 0.0011501 | 0 | 15.7464 | 2.00 | 2.3002
90 | 0 | 0 | 0.002 | 0.001 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.001035 | 0.00059756 | 0.0011966 | 0 | 18.4513 | 2.00 | 2.3932
100 | 0 | 0 | 0.002 | 0.001 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.00115 | 0.00066395 | 0.0012311 | 0 | 20.6731 | 2.00 | 2.4622
125 | 0 | 0 | 0.002 | 0.001 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.0014375 | 0.00082994 | 0.001328 | 0 | 27.9913 | 2.00 | 2.656
150 | 0 | 0 | 0.002 | 0.001 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 3.3376e-05 | 1.9269e-05 | 0.00032647 | 0.00018849 | 0.001725 | 0.00099593 | 0.0014376 | 0 | 38.4387 | 2.00 | 2.8752
[8 empty] | uA2 = | 2WVO
[9 empty] | WVO = | (sin(α / 2))^(-5/3)*(1 / 2)^2/3*Wo
[8 empty] | Wo = | 8.1617e-08 |  | Wvo = | 1.6323e-07
2 Ball Stylus of SP
[7 empty] | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2 |  |  |  | 0.2 | 1.1
[7 empty] | dD = | Ball diameter |  | 4 | [5 empty] | 0.25 | 1.35
[7 empty] | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25 |  |  | 0.28 | 1.65
[7 empty] | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2 |  | 0.455 | 2.05
[16 empty] | 0.53 | 2.55
[16 empty] | 0.62 | 3.2
[16 empty] | 0.725 | 4
0.895
Note on F4: Certificate of Calubration Ring Gauge
Note on H4: Certificate of Calubration ULM
Note on P4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = 2
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3
Note on Q6: Divisor = √3

## Uncertainty Budget(Minor)
Uncertainty Budget of Thread Ring Gauge
Nominal Value | Repeatability |  | Uncertainty of GB |  | Uncertainty of ULM |  | Resolution of ULM |  | Elastic Deformation |  | Temperature Effect |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
8.9175 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00022559 | 0 | ∞ | 2 | 0.45118
8.9175 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.00010255 | 5.9208e-05 | 0.00022559 | 0 | ∞ | 2 | 0.45118
8 | 0 | 0 | 0 | 0 | 0.00021 | 0.000105 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 9.2e-05 | 5.3116e-05 | 0.00022407 | 0 | 0.022686 | 2.00 | 0.44814
10 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.000115 | 6.6395e-05 | 0.00023237 | 0 | 0.026238 | 2.00 | 0.46473
12 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.000138 | 7.9674e-05 | 0.0002365 | 0 | 0.028157 | 2.00 | 0.473
18 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.000207 | 0.00011951 | 0.00025272 | 0 | 0.036713 | 2.00 | 0.50544
20 | 0 | 0 | 0 | 0 | 0.00023 | 0.000115 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.00023 | 0.00013279 | 0.00025927 | 0 | 0.040665 | 2.00 | 0.51853
25 | 0 | 0 | 0 | 0 | 0.00027 | 0.000135 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.0002875 | 0.00016599 | 0.0002866 | 0 | 0.060719 | 2.00 | 0.57319
30 | 0 | 0 | 0 | 0 | 0.00027 | 0.000135 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.000345 | 0.00019919 | 0.00030702 | 0 | 0.079965 | 2.00 | 0.61404
75 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.0008625 | 0.00049796 | 0.00056776 | 0 | 0.93521 | 2.00 | 1.1355
90 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.001035 | 0.00059756 | 0.00065686 | 0 | 1.6754 | 2.00 | 1.3137
100 | 0 | 0 | 0 | 0 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.00115 | 0.00066395 | 0.00071779 | 0 | 2.3891 | 2.00 | 1.4356
125 | 0 | 0 | 0.00018 | 0.00010392 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.0014375 | 0.00082994 | 0.00087977 | 0 | 5.3915 | 2.00 | 1.7595
150 | 0 | 0 | 0.00018 | 0.00010392 | 0.00039 | 0.000195 | 5e-05 | 2.8868e-05 | 0.00032647 | 0.00018849 | 0.001725 | 0.00099593 | 0.0010378 | 0 | 10.4405 | 2.00 | 2.0756
Elastic Deformation
[9 empty] | WVO = | (sin(α / 2))^-5/3*(1 / 2)^2/3*Wo
[8 empty] | Wo = | 8.1617e-08 |  | Wvo = | 1.6323e-07
2 Ball Stylus of SP
[7 empty] | F = | Measuring force (perpendicular to the flat) |  |  |  | 0.2 |  |  |  | 0.2 | 1.1
[7 empty] | dD = | Ball diameter |  | 4 | [5 empty] | 0.25 | 1.35
[7 empty] | V = | Poisson-coefficient |  | Steel = | 0.28 | Ruby = | 0.25 |  |  | 0.28 | 1.65
[7 empty] | E = | Elasticity modulus |  | Steel = | 2*10^11 | Ruby = | 4*10^11 | N/m2 |  | 0.455 | 2.05
[16 empty] | 0.53 | 2.55
[16 empty] | 0.62 | 3.2
[16 empty] | 0.725 | 4
0.895
Note on F4: Certificate of Calubration Gauge Block
Note on H4: Certificate of Calubration ULM
Note on N4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = √3
Note on I6: Divisor = 2
Note on K6: Divisor = √3
Note on M6: Divisor = √3
Note on O6: Divisor = √3

## Cert of STD
Certificate of Calibration (ULM) | [5 empty] | Certificate of Calibration (Ring Gauge) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Ring Gauge) |  |  |  |  | Certificate of Calibration (Ring Gauge)
SP-SD-001 | [5 empty] | SP-SD-007 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-003 |  |  |  |  | SP-SD-004
1 | mm | 0.21 | µm | 0.00021 | mm | 2 | 0.59 | µm | 0.00059 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 13.9962 | 0.37033 | µm | 0.00037033 | mm | 49.9965 | 2 | µm | 0.002 | mm
1.01 | mm | 0.21 | µm | 0.00021 | mm | 3 | 0.59 | µm | 0.00059 | mm | 5.1 | 0.09 | µm | 9e-05 | mm
1.05 | mm | 0.21 | µm | 0.00021 | mm | 4 | 0.7 | µm | 0.0007 | mm | 7.7 | 0.09 | µm | 9e-05 | mm
1.1 | mm | 0.21 | µm | 0.00021 | mm | 5 | 0.7 | µm | 0.0007 | mm | 10.3 | 0.09 | µm | 9e-05 | mm
2 | mm | 0.21 | µm | 0.00021 | mm | 6 | 0.75 | µm | 0.00075 | mm | 12.9 | 0.09 | µm | 9e-05 | mm
5 | mm | 0.21 | µm | 0.00021 | mm | 8 | 0.75 | µm | 0.00075 | mm | 15 | 0.1 | µm | 0.0001 | mm
10 | mm | 0.21 | µm | 0.00021 | mm | 10 | 0.8 | µm | 0.0008 | mm | 17.6 | 0.1 | µm | 0.0001 | mm
20 | mm | 0.23 | µm | 0.00023 | mm | 12 | 0.87 | µm | 0.00087 | mm | 20.2 | 0.1 | µm | 0.0001 | mm
30 | mm | 0.27 | µm | 0.00027 | mm | 16 | 0.99 | µm | 0.00099 | mm | 22.8 | 0.1 | µm | 0.0001 | mm
40 | mm | 0.27 | µm | 0.00027 | mm | 18 | 0.99 | µm | 0.00099 | mm | 25 | 0.11 | µm | 0.00011 | mm
50 | mm | 0.27 | µm | 0.00027 | mm | 20 | 1.1 | µm | 0.0011 | mm | 50 | 0.13 | µm | 0.00013 | mm
60 | mm | 0.32 | µm | 0.00032 | mm | 22 | 1.2 | µm | 0.0012 | mm | 75 | 0.16 | µm | 0.00016 | mm
70 | mm | 0.32 | µm | 0.00032 | mm | 25 | 1.2 | µm | 0.0012 | mm | 100 | 0.18 | µm | 0.00018 | mm
80 | mm | 0.39 | µm | 0.00039 | mm | 28 | 1.4 | µm | 0.0014 | mm
90 | mm | 0.39 | µm | 0.00039 | mm | 30 | 1.4 | µm | 0.0014 | mm
100 | mm | 0.39 | µm | 0.00039 | mm | 75 | 2.8 | µm | 0.0028 | mm
Note on B6: Normal Value (mm)
Note on D6: มาจากค่าความไม่แน่นอนของการวัด ULM
Note on F6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on J6: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Plain Ring Gauge
Note on L6: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on P6: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on R6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on V6: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on X6: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AB6: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on AD6: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน