## Data
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR12345 | [10 empty] | Page 1 of 1
[9 empty] | Receive Date : |  |  |  | 2016-01-11 |  |  |  |  | Calibration Date : |  |  |  | 2016-01-11
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [8 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Adress |  |  | : |  | 88/115
Equipment Name : |  |  |  |  | Digital Thermometer with RTD Probe | [5 empty] | Manufacturer : |  |  |  | SEKI | [5 empty] | Model : |  | T100
Serial No. : |  | 45678 | [5 empty] | ID No : |  | N/A |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution : |  |  | 0.1
Overall Inspection | [5 empty] | Good |  | Not Good | [7 empty] | Sensor Type: |  |  |  | RTD |  |  |  |  | 0.001
Reference Standard |  |  |  |  | : | SP-ST-002 | [6 empty] | Due Date : |  | 2017-09-09 | [9 empty] | 0.1
Reference Standard |  |  |  |  | : | SP-ST-003 | [6 empty] | Due Date : |  | 2017-10-31
Calibration Procedure |  |  |  |  | : | SP-CPT-04-06
| Nominal |  | Number of Measure |  |  | Actual Value | [8 empty] | Average STD |  |  | Average UUC |  |  | Error |  | Repeatability
| Value |  |  |  |  | STD | UUC |  |  | UUC |  | STD | [10 empty] | STD |  | UUC
| 50 |  | 1st |  |  | 50 | 50 |  |  | 50 |  | 50 |  |  | 50 |  |  | 50 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 50 | 50 |  |  | 50 |  | 50
|  |  | 3rd |  |  | 50 | 50 |  |  | 50 |  | 50
|  |  | 4th |  |  | 50 | 50 |  |  | 50 |  | 50
| 100 |  | 1st |  |  | 100 | 100 |  |  | 100 |  | 100 |  |  | 100 |  |  | 100 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 100 | 100 |  |  | 100 |  | 100
|  |  | 3rd |  |  | 100 | 100 |  |  | 100 |  | 100
|  |  | 4th |  |  | 100 | 100 |  |  | 100 |  | 100
| 150 |  | 1st |  |  | 150 | 150 |  |  | 150 |  | 150 |  |  | 150 |  |  | 150 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 150 | 150 |  |  | 150 |  | 150
|  |  | 3rd |  |  | 150 | 150 |  |  | 150 |  | 150
|  |  | 4th |  |  | 150 | 150 |  |  | 150 |  | 150
| 200 |  | 1st |  |  | 200 | 200 |  |  | 200 |  | 200 |  |  | 200 |  |  | 200 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 200 | 200 |  |  | 200 |  | 200
|  |  | 3rd |  |  | 200 | 200 |  |  | 200 |  | 200
|  |  | 4th |  |  | 200 | 200 |  |  | 200 |  | 200
| 250 |  | 1st |  |  | 250 | 250 |  |  | 250 |  | 250 |  |  | 250 |  |  | 250 |  |  | 0 |  | 0 |  | 0
|  |  | 2nd |  |  | 250 | 250 |  |  | 250 |  | 250
|  |  | 3rd |  |  | 250 | 250 |  |  | 250 |  | 250
|  |  | 4th |  |  | 250 | 250 |  |  | 250 |  | 250
Short Teram
[16 empty] | UUC 1 |  |  |  | UUC 2
[16 empty] | 50 |  |  |  | 50
0
Calibrated By : |  |  |  | Mr.Prayoon Topart
Mr.Nirut Loha
Mr.Pakapon Nammontree
Mr.Prayoon Topart
Mr.Santi Thonghlor
Mr.Werayut Jampol

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR12345 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | SP
88/115
| Equipment Name |  |  | : | Digital Thermometer with RTD Probe
| Manufacturer |  |  | : | SEKI
| Model |  |  | : | T100
| Serial Number |  |  | : | 45678
| ID. Number |  |  | : | N/A
Environmental Conditions
| Ambient Temperature |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  | : | 2016-01-11
| Relative Humidity |  |  |  | : | 50% ± 15 % |  | Calibration Date |  | : | 2016-01-11
| Location of Calibration |  |  |  | : | In-Lab |  | Recommended Due Date |  | : | 2017-01-10
| Calibration Procedure |  |  |  | : | SP-CPT-04-06 |  | Date of Issue |  | : | 2016-01-12
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Calibrated by | : | Mr.Prayoon Topart |  |  |  | Approved by :
|  |  | Calibration Officer |  |  |  | 3 |  | ( Mr. Santi Hankitudomsuk )
Authorized Signatory
Note on Q38: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR12345 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  | Model |  | Serial No. | Certificate No. | Due. Date
| Digital Thermometer |  | GT11/3850-40-392 |  | 08000098/100288 | PSL T 572/58 | 09 Sep 2017
with PRT Sensor
| Digital Thermometer |  | 34970A |  | MY44010293 | SPR16100170-2 | 31 Oct 2017
with SRT Sensor
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
|  | Certificate No. : |  |  | SPR12345 | [5 empty] | Page : 3 of 3
|  | Sensor Type : |  | RTD | [5 empty] | Unit : | °C
|  |  | Setting Value |  |  | Standard Reading | UUC Reading | Error | Uncertainty ( ± )
[12 empty] | UUC | Range | CMC
|  |  | 50 |  |  | 50 | 50 | 0 | 0.07 |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 100 |  |  | 100 | 100 | 0 | 0.07 |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 150 |  |  | 150 | 150 | 0 | 0.07 |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 200 |  |  | 200 | 200 | 0 | 0.07 |  |  | 0.066 | -40 to <250 | 0.07
|  |  | 250 |  |  | 250 | 250 | 0 | 0.07 |  |  | 0.066 | -40 to <250 | 0.07
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%.
- End of Certificate -
Mr.Chainarong Matchayamat

## Uncertainty Budget
Uncertainty Budget Dial Thermometer & Temperature Gauge (-40 to 650°C)
Unit: °C
Nominal Value | Uncertainty of STD |  | Short term stability of UUC |  | Resolution of STD |  | Resolution of UUC |  | Repeatability STD |  | Repeatability UUC |  | Uc | Ui | Veff | K95 | U95%
°C | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
50 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
100 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
150 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
200 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
250 | 0.03 | 0.015 | 0 | 0 | 0.0005 | 0.00028868 | 0.05 | 0.028868 | 0 | 0 | 0 | 0 | 0.032533 | 0 | ∞ | 2 | 0.065067
Note on C4: Certificate of Calubration Digithermometer with PRT senror
Note on E4: ได้จากค่า ผลการสอบเทียบที่จุดสอบเทียบแรกครั้งที่1 - ผลการสอบเทียบที่จุดสอบเทียบแรกครั้งที่2
Note on G4: ค่าความละเอียดในการอ่านของ STD = 0.001 C
Note on I4: ค่าความละเอียดในการอ่านของ UUC
Note on D5: Divisor = 2
Note on F5: Divisor = √3
Note on H5: Divisor = √3
Note on J5: Divisor = √3
Note on L5: Divisor = 1
Note on N5: Divisor = 1

## UCB New '2016
ใบแบบการประมาณค่าความไม่แน่นอนของการวัด
Function : | Temperature |  |  | Testpoint | 50 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 100 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 150 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 200 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
Function : | Temperature |  |  | Testpoint | 250 |  | °C | Range
Symbol | Source of uncertainty | Relative | Absolute |  | Probability distribution |  | Divisor | ci | Relative | Absolute | Degree of freedom
|  | (+/-) % | (+/-) °C | [5 empty] | (+/-) % | (+/-) °C
UB1 | Uncertainty of STD |  | 0.03 |  | Normal |  | 2 | 1 |  | 0.015 | ¥
UB2 | Short term stability of UUC |  | 0 |  | Rectangular |  | 1.7321 | 1 |  | 0 | ¥
UB3 | Resolution of STD |  | 0.0005 |  | Rectangular |  | 1.7321 | 1 |  | 0.00028868 | ¥
UB4 | Resolution of UUC |  | 0.05 |  | Rectangular |  | 1.7321 | 1 |  | 0.028868 | ¥
UA1 | Repeatability STD |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
UA2 | Repeatability UUC |  | 0 |  | Normal |  | 1 | 1 |  | 0 | 7
Uc | Combined uncertainty |  |  |  | Normal |  |  |  |  | 0.032533 | ¥
U | Expaned uncertainty |  |  |  | k = | 2 |  |  |  | 0.066 | ¥
2
2

## Cert of STD
Digital Thermometer with TC PRT | [5 empty] | Digital Thermometer with TC SRT
SP-ST-002
Due Date |  |  | 2016-06-05 |  |  | SP-ST-003
Point | Uncert |  | Point | Uncert |  | Due Date | 2016-11-02
°C | CH 1 |  | °C | CH 2 |  | Point | Uncert
-40.0092 | 0.03 | °C | -80.0236 | 0.03 | °C | °C
-0.00145 | 0.03 | °C | -40.0092 | 0.03 | °C | 250 | 0.7 | °C
50.0427 | 0.03 | °C | -0.00145 | 0.03 | °C | 500 | 0.7 | °C
100.0512 | 0.03 | °C | 50.0427 | 0.03 | °C | 750 | 2.6 | °C
150.0512 | 0.03 | °C | 100.0518 | 0.03 | °C | 1200 | 2.6 | °C
200.0524 | 0.03 | °C
250.0772 | 0.03 | °C
Note on B6: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I7: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B8: Nathaphol Boonmee: Normal Value
Note on C8: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on F8: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on I9: Nathaphol Boonmee: Normal Value
Note on J9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง