## Data Record(50)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR16010009
[8 empty] | Receive Date : |  |  |  | 2016-01-01 |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [7 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  | SP
Equipment Name : |  |  |  | Liquid Bath | [5 empty] | Manufacturer : |  |  |  | OKOKOK |  |  |  |  | Model : |  |  | TE-01
Serial No. : |  | 987654 |  |  |  |  | ID No : |  | TT-1 |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution : |  | 0.1
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : | [12 empty] | Due Date :
Calibration and Data Record
Record | Chamber |  | Standard Reading (oC) @ Probe Number
No. | Indicator |  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  |  | #4 |  |  |  |  | #5ref
1 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
2 | 50 |  | 50.272 |  |  |  | 50.218 |  |  |  | 50.184 |  |  |  |  | 50.773 |  |  |  |  | 50.591
3 | 50 |  | 50.465 |  |  |  | 50.636 |  |  |  | 50.959 |  |  |  |  | 50.972 |  |  |  |  | 50.109
4 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
5 | 50 |  | 50.549 |  |  |  | 50.524 |  |  |  | 50.852 |  |  |  |  | 50.016 |  |  |  |  | 50.464
6 | 50 |  | 50.995 |  |  |  | 50.146 |  |  |  | 50.833 |  |  |  |  | 50.564 |  |  |  |  | 50.036
7 | 50 |  | 50.545 |  |  |  | 50.101 |  |  |  | 50.868 |  |  |  |  | 50.815 |  |  |  |  | 50.621
8 | 50 |  | 50.445 |  |  |  | 50.49 |  |  |  | 50.466 |  |  |  |  | 50.743 |  |  |  |  | 50.821
9 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
10 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
11 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
12 | 50 |  | 50.272 |  |  |  | 50.218 |  |  |  | 50.184 |  |  |  |  | 50.773 |  |  |  |  | 50.591
13 | 50 |  | 50.465 |  |  |  | 50.636 |  |  |  | 50.959 |  |  |  |  | 50.972 |  |  |  |  | 50.109
14 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
15 | 50 |  | 50.549 |  |  |  | 50.524 |  |  |  | 50.852 |  |  |  |  | 50.016 |  |  |  |  | 50.464
16 | 50 |  | 50.995 |  |  |  | 50.146 |  |  |  | 50.833 |  |  |  |  | 50.564 |  |  |  |  | 50.036
17 | 50 |  | 50.545 |  |  |  | 50.101 |  |  |  | 50.868 |  |  |  |  | 50.815 |  |  |  |  | 50.621
18 | 50 |  | 50.445 |  |  |  | 50.49 |  |  |  | 50.466 |  |  |  |  | 50.743 |  |  |  |  | 50.821
19 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
20 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
21 | 50 |  | 50.921 |  |  |  | 50.502 |  |  |  | 50.923 |  |  |  |  | 50.029 |  |  |  |  | 50.109
22 | 50 |  | 50.272 |  |  |  | 50.218 |  |  |  | 50.184 |  |  |  |  | 50.773 |  |  |  |  | 50.591
23 | 50 |  | 50.465 |  |  |  | 50.636 |  |  |  | 50.959 |  |  |  |  | 50.972 |  |  |  |  | 50.109
24 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
25 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
26 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
27 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
28 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
29 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
30 | 50 |  | 50.356 |  |  |  | 50.608 |  |  |  | 50.932 |  |  |  |  | 50.495 |  |  |  |  | 50.327
Averg | 50 |  | 50.5643 |  |  |  | 50.469 |  |  |  | 50.8105 |  |  |  |  | 50.4723 |  |  |  |  | 50.323
Repeat |  |  | 0.33297
Summary of Calibration Results
Chamber Uniformity: Maximum difference of measured temp. between any probes and reference probe | [17 empty] | = | 0.959 |  |  | oC
Temperature Stability: One-half of greatest maximum diff. of measured temp. at any probes | [17 empty] | = | 0.478 |  |  | oC
Overall Variation: Difference of maximum and minimum measured temperatures | [17 empty] | = | 0.979 |  |  | oC
Standard Reading (oC) @ Probe Number
|  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  | #4 |  |  |  |  | #5
Min. Temp. |  | 50.272 |  |  |  | 50.101 |  |  |  | 50.184 |  |  |  | 50.016 |  |  |  |  | 50.036
Max. Temp. |  | 50.995 |  |  |  | 50.636 |  |  |  | 50.959 |  |  |  | 50.972 |  |  |  |  | 50.821
Diff. Temp. |  | 0.723 |  |  |  | 0.535 |  |  |  | 0.775 |  |  |  | 0.956 |  |  |  |  | 0.785
Stability |  | 0.3615 |  |  |  | 0.2675 |  |  |  | 0.3875 |  |  |  | 0.478 |  |  |  |  | 0.3925
Difference Temperature Between Any Probes and Ref.
Record | Chamber |  | Diff. Temperature Between Any Probe and Probe
No. | Indicator |  | #1 - #5ref |  |  |  |  | #2 - #5ref |  |  |  |  | #3 - #5ref | [5 empty] | #4 - #5ref
1 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
2 | 50 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
3 | 50 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
4 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
5 | 50 |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
6 | 50 |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
7 | 50 |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
8 | 50 |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
9 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
10 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
11 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
12 | 50 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
13 | 50 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
14 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
15 | 50 |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
16 | 50 |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
17 | 50 |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
18 | 50 |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
19 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
20 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
21 | 50 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
22 | 50 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
23 | 50 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
24 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
25 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
26 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
27 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
28 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
29 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
30 | 50 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
Calibrated By |  |  | : | Mr.Natthaphol Boonmee
|  | 3 | Mr.Natthaphol Boonmee
|  | 11 | Ms. Arunkamon Raramanus

## Data Record(100)
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  | SPR16010009
[8 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [7 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  | SP
Equipment Name : |  |  |  | Liquid Bath | [5 empty] | Manufacturer : |  |  |  | OKOKOK |  |  |  |  | Model : |  |  | TE-01
Serial No. : |  | 987654 |  |  |  |  | ID No : |  | TT-1 |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution :
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Referance Standard : | [12 empty] | Due Date :
Calibration and Data Record
Record | Chamber |  | Standard Reading (oC) @ Probe Number
No. | Indicator |  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  |  | #4 |  |  |  |  | #5ref
1 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
2 | 100 |  | 100.272 |  |  |  | 100.218 |  |  |  | 100.184 |  |  |  |  | 100.773 |  |  |  |  | 100.591
3 | 100 |  | 100.465 |  |  |  | 100.636 |  |  |  | 100.959 |  |  |  |  | 100.972 |  |  |  |  | 100.109
4 | 100 |  | 100.356 |  |  |  | 100.608 |  |  |  | 100.932 |  |  |  |  | 100.495 |  |  |  |  | 100.327
5 | 100 |  | 100.549 |  |  |  | 100.524 |  |  |  | 100.852 |  |  |  |  | 100.016 |  |  |  |  | 100.464
6 | 100 |  | 100.995 |  |  |  | 100.146 |  |  |  | 100.833 |  |  |  |  | 100.564 |  |  |  |  | 100.036
7 | 100 |  | 100.545 |  |  |  | 100.101 |  |  |  | 100.868 |  |  |  |  | 100.815 |  |  |  |  | 100.621
8 | 100 |  | 100.445 |  |  |  | 100.49 |  |  |  | 100.466 |  |  |  |  | 100.743 |  |  |  |  | 100.821
9 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
10 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
11 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
12 | 100 |  | 100.272 |  |  |  | 100.218 |  |  |  | 100.184 |  |  |  |  | 100.773 |  |  |  |  | 100.591
13 | 100 |  | 100.465 |  |  |  | 100.636 |  |  |  | 100.959 |  |  |  |  | 100.972 |  |  |  |  | 100.109
14 | 100 |  | 100.356 |  |  |  | 100.608 |  |  |  | 100.932 |  |  |  |  | 100.495 |  |  |  |  | 100.327
15 | 100 |  | 100.549 |  |  |  | 100.524 |  |  |  | 100.852 |  |  |  |  | 100.016 |  |  |  |  | 100.464
16 | 100 |  | 100.995 |  |  |  | 100.146 |  |  |  | 100.833 |  |  |  |  | 100.564 |  |  |  |  | 100.036
17 | 100 |  | 100.545 |  |  |  | 100.101 |  |  |  | 100.868 |  |  |  |  | 100.815 |  |  |  |  | 100.621
18 | 100 |  | 100.445 |  |  |  | 100.49 |  |  |  | 100.466 |  |  |  |  | 100.743 |  |  |  |  | 100.821
19 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
20 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
21 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
22 | 100 |  | 100.272 |  |  |  | 100.218 |  |  |  | 100.184 |  |  |  |  | 100.773 |  |  |  |  | 100.591
23 | 100 |  | 100.465 |  |  |  | 100.636 |  |  |  | 100.959 |  |  |  |  | 100.972 |  |  |  |  | 100.109
24 | 100 |  | 100.356 |  |  |  | 100.608 |  |  |  | 100.932 |  |  |  |  | 100.495 |  |  |  |  | 100.327
25 | 100 |  | 100.549 |  |  |  | 100.524 |  |  |  | 100.852 |  |  |  |  | 100.016 |  |  |  |  | 100.464
26 | 100 |  | 100.995 |  |  |  | 100.146 |  |  |  | 100.833 |  |  |  |  | 100.564 |  |  |  |  | 100.036
27 | 100 |  | 100.545 |  |  |  | 100.101 |  |  |  | 100.868 |  |  |  |  | 100.815 |  |  |  |  | 100.621
28 | 100 |  | 100.445 |  |  |  | 100.49 |  |  |  | 100.466 |  |  |  |  | 100.743 |  |  |  |  | 100.821
29 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
30 | 100 |  | 100.921 |  |  |  | 100.502 |  |  |  | 100.923 |  |  |  |  | 100.029 |  |  |  |  | 100.109
Averg | 100 |  | 100.639 |  |  |  | 100.4229 |  |  |  | 100.7863 |  |  |  |  | 100.4465 |  |  |  |  | 100.3296
Repeat |  |  | 0.37105
Summary of Calibration Results
Chamber Uniformity: Maximum difference of measured temp. between any probes and reference probe | [17 empty] | = | 0.959 |  |  | oC
Temperature Stability: One-half of greatest maximum diff. of measured temp. at any probes | [17 empty] | = | 0.478 |  |  | oC
Overall Variation: Difference of maximum and minimum measured temperatures | [17 empty] | = | 0.979 |  |  | oC
Standard Reading (oC) @ Probe Number
|  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  | #4 |  |  |  |  | #5
Min. Temp. |  | 100.272 |  |  |  | 100.101 |  |  |  | 100.184 |  |  |  | 100.016 |  |  |  |  | 100.036
Max. Temp. |  | 100.995 |  |  |  | 100.636 |  |  |  | 100.959 |  |  |  | 100.972 |  |  |  |  | 100.821
Diff. Temp. |  | 0.723 |  |  |  | 0.535 |  |  |  | 0.775 |  |  |  | 0.956 |  |  |  |  | 0.785
Stability |  | 0.3615 |  |  |  | 0.2675 |  |  |  | 0.3875 |  |  |  | 0.478 |  |  |  |  | 0.3925
Difference Temperature Between Any Probes and Ref.
Record | Chamber |  | Diff. Temperature Between Any Probe and Probe
No. | Indicator |  | #1 - #5ref |  |  |  |  | #2 - #5ref |  |  |  |  | #3 - #5ref | [5 empty] | #4 - #5ref
1 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
2 | 100 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
3 | 100 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
4 | 100 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
5 | 100 |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
6 | 100 |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
7 | 100 |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
8 | 100 |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
9 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
10 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
11 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
12 | 100 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
13 | 100 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
14 | 100 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
15 | 100 |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
16 | 100 |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
17 | 100 |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
18 | 100 |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
19 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
20 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
21 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
22 | 100 |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
23 | 100 |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
24 | 100 |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
25 | 100 |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
26 | 100 |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
27 | 100 |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
28 | 100 |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
29 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
30 | 100 |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
Calibrated By |  |  | : | Mr.Natthaphol Boonmee
|  | 3 | Mr.Natthaphol Boonmee
|  | 11 | Ms. Arunkamon Raramanus

## Data Record(150)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  | SPR16010009
[9 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [8 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Liquid Bath | [5 empty] | Manufacturer : |  |  |  | OKOKOK |  |  |  |  | Model : |  |  | TE-01
Serial No. : |  | 987654 | [5 empty] | ID No : |  | TT-1 |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution :
Overall Inspection | [5 empty] | Good |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Calibration and Data Record
Record | Chamber |  |  | Standard Reading (oC) @ Probe Number
No. | Indicator |  |  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  |  | #4 |  |  |  |  | #5
1 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
2 | 150 |  |  | 150.272 |  |  |  | 150.218 |  |  |  | 150.184 |  |  |  |  | 150.773 |  |  |  |  | 150.591
3 | 150 |  |  | 150.465 |  |  |  | 150.636 |  |  |  | 150.959 |  |  |  |  | 150.972 |  |  |  |  | 150.109
4 | 150 |  |  | 150.356 |  |  |  | 150.608 |  |  |  | 150.932 |  |  |  |  | 150.495 |  |  |  |  | 150.327
5 | 150 |  |  | 150.549 |  |  |  | 150.524 |  |  |  | 150.852 |  |  |  |  | 150.016 |  |  |  |  | 150.464
6 | 150 |  |  | 150.995 |  |  |  | 150.146 |  |  |  | 150.833 |  |  |  |  | 150.564 |  |  |  |  | 150.036
7 | 150 |  |  | 150.545 |  |  |  | 150.101 |  |  |  | 150.868 |  |  |  |  | 150.815 |  |  |  |  | 150.621
8 | 150 |  |  | 150.445 |  |  |  | 150.49 |  |  |  | 150.466 |  |  |  |  | 150.743 |  |  |  |  | 150.821
9 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
10 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
11 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
12 | 150 |  |  | 150.272 |  |  |  | 150.218 |  |  |  | 150.184 |  |  |  |  | 150.773 |  |  |  |  | 150.591
13 | 150 |  |  | 150.465 |  |  |  | 150.636 |  |  |  | 150.959 |  |  |  |  | 150.972 |  |  |  |  | 150.109
14 | 150 |  |  | 150.356 |  |  |  | 150.608 |  |  |  | 150.932 |  |  |  |  | 150.495 |  |  |  |  | 150.327
15 | 150 |  |  | 150.549 |  |  |  | 150.524 |  |  |  | 150.852 |  |  |  |  | 150.016 |  |  |  |  | 150.464
16 | 150 |  |  | 150.995 |  |  |  | 150.146 |  |  |  | 150.833 |  |  |  |  | 150.564 |  |  |  |  | 150.036
17 | 150 |  |  | 150.545 |  |  |  | 150.101 |  |  |  | 150.868 |  |  |  |  | 150.815 |  |  |  |  | 150.621
18 | 150 |  |  | 150.445 |  |  |  | 150.49 |  |  |  | 150.466 |  |  |  |  | 150.743 |  |  |  |  | 150.821
19 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
20 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
21 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
22 | 150 |  |  | 150.272 |  |  |  | 150.218 |  |  |  | 150.184 |  |  |  |  | 150.773 |  |  |  |  | 150.591
23 | 150 |  |  | 150.465 |  |  |  | 150.636 |  |  |  | 150.959 |  |  |  |  | 150.972 |  |  |  |  | 150.109
24 | 150 |  |  | 150.356 |  |  |  | 150.608 |  |  |  | 150.932 |  |  |  |  | 150.495 |  |  |  |  | 150.327
25 | 150 |  |  | 150.549 |  |  |  | 150.524 |  |  |  | 150.852 |  |  |  |  | 150.016 |  |  |  |  | 150.464
26 | 150 |  |  | 150.995 |  |  |  | 150.146 |  |  |  | 150.833 |  |  |  |  | 150.564 |  |  |  |  | 150.036
27 | 150 |  |  | 150.545 |  |  |  | 150.101 |  |  |  | 150.868 |  |  |  |  | 150.815 |  |  |  |  | 150.621
28 | 150 |  |  | 150.445 |  |  |  | 150.49 |  |  |  | 150.466 |  |  |  |  | 150.743 |  |  |  |  | 150.821
29 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
30 | 150 |  |  | 150.921 |  |  |  | 150.502 |  |  |  | 150.923 |  |  |  |  | 150.029 |  |  |  |  | 150.109
Averg | 150 |  |  | 150.639 |  |  |  | 150.4229 |  |  |  | 150.7863 |  |  |  |  | 150.4465 |  |  |  |  | 150.3296
Repeat |  |  |  | 0.37105
Summary of Calibration Results
Chamber Uniformity: Maximum difference of measured temp. between any probes and reference probe | [18 empty] | = | 0.959 |  |  | oC
Temperature Stability: One-half of greatest maximum diff. of measured temp. at any probes | [18 empty] | = | 0.478 |  |  | oC
Overall Variation: Difference of maximum and minimum measured temperatures | [18 empty] | = | 0.979 |  |  | oC
Standard Reading (oC) @ Probe Number
|  | #1 |  |  |  |  | #2 |  |  |  | #3 |  |  |  | #4 |  |  |  |  | #5
Min. Temp. |  | 150.272 |  |  |  |  | 150.101 |  |  |  | 150.184 |  |  |  | 150.016 |  |  |  |  | 150.036
Max. Temp. |  | 150.995 |  |  |  |  | 150.636 |  |  |  | 150.959 |  |  |  | 150.972 |  |  |  |  | 150.821
Diff. Temp. |  | 0.723 |  |  |  |  | 0.535 |  |  |  | 0.775 |  |  |  | 0.956 |  |  |  |  | 0.785
Stability |  | 0.3615 |  |  |  |  | 0.2675 |  |  |  | 0.3875 |  |  |  | 0.478 |  |  |  |  | 0.3925
Difference Temperature Between Any Probes and Ref.
Record | Chamber |  |  | Diff. Temperature Between Any Probe and Probe
No. | Indicator |  |  | #1 - #5ref |  |  |  |  | #2 - #5ref |  |  |  |  | #3 - #5ref | [5 empty] | #4 - #5ref
1 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
2 | 150 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
3 | 150 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
4 | 150 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
5 | 150 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
6 | 150 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
7 | 150 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
8 | 150 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
9 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
10 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
11 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
12 | 150 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
13 | 150 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
14 | 150 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
15 | 150 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
16 | 150 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
17 | 150 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
18 | 150 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
19 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
20 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
21 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
22 | 150 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
23 | 150 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
24 | 150 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
25 | 150 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
26 | 150 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
27 | 150 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
28 | 150 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
29 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
30 | 150 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
Calibrated By |  |  | : | Mr.Natthaphol Boonmee
|  | 10 |  | Mr.Chainarong Matchayamat
|  | 11 |  | Ms. Arunkamon Raramanus

## Data Record(200)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  | SPR16010009
[9 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  |  |  |  | 2016-01-01
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  | 23 |  | oC | 50 |  | %RH
Temperature | [8 empty] | Location |  |  |  | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Liquid Bath | [5 empty] | Manufacturer : |  |  |  | OKOKOK |  |  |  |  | Model : |  |  | TE-01
Serial No. : |  | 987654 | [5 empty] | ID No : |  | TT-1 |  |  |  | Range : |  |  | 0 |  | to | 200 |  | Resolution :
Overall Inspection | [5 empty] | Good |  |  | Not Good
Referance Standard : | [13 empty] | Due Date :
Calibration and Data Record
Record | Chamber |  |  | Standard Reading (oC) @ Probe Number
No. | Indicator |  |  | #1 |  |  |  | #2 |  |  |  | #3 |  |  |  |  | #4 |  |  |  |  | #5
1 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
2 | 200 |  |  | 200.272 |  |  |  | 200.218 |  |  |  | 200.184 |  |  |  |  | 200.773 |  |  |  |  | 200.591
3 | 200 |  |  | 200.465 |  |  |  | 200.636 |  |  |  | 200.959 |  |  |  |  | 200.972 |  |  |  |  | 200.109
4 | 200 |  |  | 200.356 |  |  |  | 200.608 |  |  |  | 200.932 |  |  |  |  | 200.495 |  |  |  |  | 200.327
5 | 200 |  |  | 200.549 |  |  |  | 200.524 |  |  |  | 200.852 |  |  |  |  | 200.016 |  |  |  |  | 200.464
6 | 200 |  |  | 200.995 |  |  |  | 200.146 |  |  |  | 200.833 |  |  |  |  | 200.564 |  |  |  |  | 200.036
7 | 200 |  |  | 200.545 |  |  |  | 200.101 |  |  |  | 200.868 |  |  |  |  | 200.815 |  |  |  |  | 200.621
8 | 200 |  |  | 200.445 |  |  |  | 200.49 |  |  |  | 200.466 |  |  |  |  | 200.743 |  |  |  |  | 200.821
9 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
10 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
11 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
12 | 200 |  |  | 200.272 |  |  |  | 200.218 |  |  |  | 200.184 |  |  |  |  | 200.773 |  |  |  |  | 200.591
13 | 200 |  |  | 200.465 |  |  |  | 200.636 |  |  |  | 200.959 |  |  |  |  | 200.972 |  |  |  |  | 200.109
14 | 200 |  |  | 200.356 |  |  |  | 200.608 |  |  |  | 200.932 |  |  |  |  | 200.495 |  |  |  |  | 200.327
15 | 200 |  |  | 200.549 |  |  |  | 200.524 |  |  |  | 200.852 |  |  |  |  | 200.016 |  |  |  |  | 200.464
16 | 200 |  |  | 200.995 |  |  |  | 200.146 |  |  |  | 200.833 |  |  |  |  | 200.564 |  |  |  |  | 200.036
17 | 200 |  |  | 200.545 |  |  |  | 200.101 |  |  |  | 200.868 |  |  |  |  | 200.815 |  |  |  |  | 200.621
18 | 200 |  |  | 200.445 |  |  |  | 200.49 |  |  |  | 200.466 |  |  |  |  | 200.743 |  |  |  |  | 200.821
19 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
20 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
21 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
22 | 200 |  |  | 200.272 |  |  |  | 200.218 |  |  |  | 200.184 |  |  |  |  | 200.773 |  |  |  |  | 200.591
23 | 200 |  |  | 200.465 |  |  |  | 200.636 |  |  |  | 200.959 |  |  |  |  | 200.972 |  |  |  |  | 200.109
24 | 200 |  |  | 200.356 |  |  |  | 200.608 |  |  |  | 200.932 |  |  |  |  | 200.495 |  |  |  |  | 200.327
25 | 200 |  |  | 200.549 |  |  |  | 200.524 |  |  |  | 200.852 |  |  |  |  | 200.016 |  |  |  |  | 200.464
26 | 200 |  |  | 200.995 |  |  |  | 200.146 |  |  |  | 200.833 |  |  |  |  | 200.564 |  |  |  |  | 200.036
27 | 200 |  |  | 200.545 |  |  |  | 200.101 |  |  |  | 200.868 |  |  |  |  | 200.815 |  |  |  |  | 200.621
28 | 200 |  |  | 200.445 |  |  |  | 200.49 |  |  |  | 200.466 |  |  |  |  | 200.743 |  |  |  |  | 200.821
29 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
30 | 200 |  |  | 200.921 |  |  |  | 200.502 |  |  |  | 200.923 |  |  |  |  | 200.029 |  |  |  |  | 200.109
Averg | 200 |  |  | 200.639 |  |  |  | 200.4229 |  |  |  | 200.7863 |  |  |  |  | 200.4465 |  |  |  |  | 200.3296
Repeat |  |  |  | 0.31636
Summary of Calibration Results
Chamber Uniformity: Maximum difference of measured temp. between any probes and reference probe | [18 empty] | = | 0.959 |  |  | oC
Temperature Stability: One-half of greatest maximum diff. of measured temp. at any probes | [18 empty] | = | 0.478 |  |  | oC
Overall Variation: Difference of maximum and minimum measured temperatures | [18 empty] | = | 0.979 |  |  | oC
Standard Reading (oC) @ Probe Number
|  | #1 |  |  |  |  | #2 |  |  |  | #3 |  |  |  | #4 |  |  |  |  | #5
Min. Temp. |  | 200.272 |  |  |  |  | 200.101 |  |  |  | 200.184 |  |  |  | 200.016 |  |  |  |  | 200.036
Max. Temp. |  | 200.995 |  |  |  |  | 200.636 |  |  |  | 200.959 |  |  |  | 200.972 |  |  |  |  | 200.821
Diff. Temp. |  | 0.723 |  |  |  |  | 0.535 |  |  |  | 0.775 |  |  |  | 0.956 |  |  |  |  | 0.785
Stability |  | 0.3615 |  |  |  |  | 0.2675 |  |  |  | 0.3875 |  |  |  | 0.478 |  |  |  |  | 0.3925
Difference Temperature Between Any Probes and Ref.
Record | Chamber |  |  | Diff. Temperature Between Any Probe and Probe
No. | Indicator |  |  | #1 - #5ref |  |  |  |  | #2 - #5ref |  |  |  |  | #3 - #5ref | [5 empty] | #4 - #5ref
1 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
2 | 200 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
3 | 200 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
4 | 200 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
5 | 200 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
6 | 200 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
7 | 200 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
8 | 200 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
9 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
10 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
11 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
12 | 200 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
13 | 200 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
14 | 200 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
15 | 200 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
16 | 200 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
17 | 200 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
18 | 200 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
19 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
20 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
21 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
22 | 200 |  |  | 0.319 |  |  |  |  | 0.373 |  |  |  |  | 0.407 | [5 empty] | 0.182
23 | 200 |  |  | 0.356 |  |  |  |  | 0.527 |  |  |  |  | 0.85 | [5 empty] | 0.863
24 | 200 |  |  | 0.029 |  |  |  |  | 0.281 |  |  |  |  | 0.605 | [5 empty] | 0.168
25 | 200 |  |  | 0.085 |  |  |  |  | 0.06 |  |  |  |  | 0.388 | [5 empty] | 0.448
26 | 200 |  |  | 0.959 |  |  |  |  | 0.11 |  |  |  |  | 0.797 | [5 empty] | 0.528
27 | 200 |  |  | 0.076 |  |  |  |  | 0.52 |  |  |  |  | 0.247 | [5 empty] | 0.194
28 | 200 |  |  | 0.376 |  |  |  |  | 0.331 |  |  |  |  | 0.355 | [5 empty] | 0.078
29 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
30 | 200 |  |  | 0.812 |  |  |  |  | 0.393 |  |  |  |  | 0.814 | [5 empty] | 0.08
Calibrated By |  |  | : | Mr.Natthaphol Boonmee
|  | 10 |  | Mr.Chainarong Matchayamat
|  | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR16010009 | [6 empty] | Page : 1 of 3
| Customer |  |  |  | : | SP
| Equipment Name |  |  |  | : | Liquid Bath
| Manufacturer |  |  |  | : | OKOKOK
| Model |  |  |  | : | TE-01
| Serial Number |  |  |  | : | 987654
| ID. Number |  |  |  | : | TT-1
Environmental Conditions
| Ambient Temperature |  |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  |  |  | : | 2016-01-01
| Relative Humidity |  |  |  |  | : | 50% ± 15 % |  | Calibration Date |  |  |  | : | 2016-01-01
| Location of Calibration |  |  |  |  | : | In-Lab |  | Recommended Due Date |  |  |  | : | 2016-12-31
| Calibration Procedure |  |  |  |  | : | SP-CPT-04-04
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-02 |  |  |  | Approved by :
| Calibrated by |  | : | Mr.Natthaphol Boonmee |  |  |  | 3 |  |  |  | ( Mr. Natthaphol Boonmee )
Authorized Signatory
[10 empty] | 1 | Mr.Sombut Srikampa
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q37: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16010009 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Data Acquisition Switch/Unit |  |  | 34970A | MY44010293 | TC/150327 | 2016-11-02
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  | SPR16010009 | [7 empty] | Page : 3 of 3
| 1. Temperature Accuracy in the Measurement Zone. | [10 empty] | Unit : | °C
| UUC Setting | Measured Temperature (°C) @ Probe No. | [8 empty] | Uncertainty (±)
|  | # 1 |  | # 2 | # 3 |  |  | # 4 |  | # 5
| 50 | 50.5643 |  | 50.469 | 50.8105 |  |  | 50.4723 |  | 50.323 | 0.983
| 100 | 100.639 |  | 100.4229 | 100.7863 |  |  | 100.4465 |  | 100.3296 | 0.98487
| 150 | 150.639 |  | 150.4229 | 150.7863 |  |  | 150.4465 |  | 150.3296 | 0.98487
| 200 | 200.639 |  | 200.4229 | 200.7863 |  |  | 200.4465 |  | 200.3296 | 0.98225
| 2. Temperature Uniformity, Stability | [5 empty] | Unit : |  | °C
| UUC Setting | UUC Reading |  | Temperature Stability |  | Temperature Uniformity
| 50 | 50 |  | 0.478 |  | 0.959
| 100 | 100 |  | 0.478 |  | 0.959
| 150 | 150 |  | 0.478 |  | 0.959
| 200 | 200 |  | 0.478 |  | 0.959
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget 0 to 200 C
Uncertainty Budget Oven / Chamber (0 to 200°C)
Unit: °C
Nominal Value | Uncertainty of STD |  | UUC Stability |  | Resolution of STD |  | Resolution of UUC |  | Repeatability |  | Uc | Ui | Veff | K95 | U95%
°C | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
50 | 0.8 | 0.4 | 0.478 | 0.27597 | 0.05 | 0.028868 | 0.05 | 0.028868 | 0.060793 | 0.060793 | 0.49145 | 0.060793 | 12812.6986 | 2.0002 | 0.983
100 | 0.8 | 0.4 | 0.478 | 0.27597 | 0.05 | 0.028868 | 0.05 | 0.028868 | 0.067744 | 0.067744 | 0.49236 | 0.067744 | 8370.9816 | 2.0003 | 0.98487
150 | 0.8 | 0.4 | 0.478 | 0.27597 | 0.05 | 0.028868 | 0.05 | 0.028868 | 0.067744 | 0.067744 | 0.49236 | 0.067744 | 8370.9816 | 2.0003 | 0.98487
200 | 0.8 | 0.4 | 0.478 | 0.27597 | 0.05 | 0.028868 | 0.05 | 0.028868 | 0.05776 | 0.05776 | 0.49108 | 0.05776 | 15676.2092 | 2.0002 | 0.98225
Note on C5: Certificate of Calubration Digithermometer with PRT senror
Note on G5: ค่าความละเอียดในการอ่านของ STD = 0.1 C
Note on I5: ค่าความละเอียดในการอ่านของ UUC
Note on D6: Divisor = 2
Note on F6: Divisor = √3
Note on H6: Divisor = √3
Note on J6: Divisor = √3
Note on L6: Divisor = 1

## Uncert of STD
Digital Thermometer with TC K
SP-ST-001
Due Date | [6 empty] | 2016-11-02
Point |  | Uncerttainty of Chanel
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
0 | °C | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | °C
50 | °C | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | °C
100 | °C | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | °C
150 | °C | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | °C
200 | °C | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | °C
Note on B5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B7: Nathaphol Boonmee: Normal Value