## Data Record (25mm)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  | SPR15120011 | [6 empty] | Page |  | 1 | of | 1
[9 empty] | Receive Date : |  |  | 2016-01-18 |  |  |  |  | Calibration Date : |  | 2016-01-19
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  | 20 | oC | 50 | %RH
Depth Micrometer | [8 empty] | Location |  |  |  |  | In Lab |  |  |  |  | On Site
Customer Name : |  |  |  |  | LG
Equipment Name : |  |  |  |  | Depth Micrometer | [6 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  | 123 | [5 empty] | Serial No. : |  | 456 |  |  |  | ID No :
Range : |  |  | 0 |  | to | 25 |  | mm |  |  | Resolution : | 0.001 |  | mm
Overall Inspection | [6 empty] | Good |  |  | Not Good
Referance Standard : | [10 empty] | Due Date :
Referance Standard : | [10 empty] | Due Date :
Flatness of The Spindle
| Optical Flat | [6 empty] | Fringe Number
| 12 |  |  |  |  | mm |  | 1
Parallelism
| Normal Value |  |  | Position |  |  | UUC Reading |  | Parallelism
| 2.5 |  |  | 1 |  |  | 2.501 |  | 0.001
|  |  |  | 2 |  |  | 2.5 |  | 0
|  |  |  | 3 |  |  | 2.5 |  | 0
|  |  |  | 4 |  |  | 2.5 |  | 0
| Measurement Results | [16 empty] | Unit : |  | mm
| Normal Value |  |  | UUC Reading | [9 empty] | Average |  |  | Repeatability | Error
|  |  |  | X1 |  |  | X2 |  | X3 |  | X4
| 0 |  |  | 0 |  |  | 0 |  | 0 |  | 0 |  |  | 0 |  |  | 0 | 0
| 2.2 |  |  | 2.2 |  |  | 2.2 |  | 2.2 |  | 2.2 |  |  | 2.2 |  |  | 0 | 0
| 4.8 |  |  | 4.8 |  |  | 4.8 |  | 4.8 |  | 4.8 |  |  | 4.8 |  |  | 0 | 0
| 7.4 |  |  | 7.4 |  |  | 7.4 |  | 7.4 |  | 7.4 |  |  | 7.4 |  |  | 0 | 0
| 10 |  |  | 10 |  |  | 10 |  | 10 |  | 10 |  |  | 10 |  |  | 0 | 0
| 12.1 |  |  | 12.1 |  |  | 12.1 |  | 12.1 |  | 12.1 |  |  | 12.1 |  |  | 0 | 0
| 14.7 |  |  | 14.7 |  |  | 14.7 |  | 14.7 |  | 14.7 |  |  | 14.7 |  |  | 0 | 0
| 17.3 |  |  | 17.3 |  |  | 17.3 |  | 17.3 |  | 17.3 |  |  | 17.3 |  |  | 0 | 0
| 19.9 |  |  | 19.9 |  |  | 19.9 |  | 19.9 |  | 19.9 |  |  | 19.9 |  |  | 0 | 0
| 22.5 |  |  | 22.5 |  |  | 22.5 |  | 22.5 |  | 22.5 |  |  | 22.5 |  |  | 0 | 0
| 25 |  |  | 25 |  |  | 25 |  | 25 |  | 25 |  |  | 25 |  |  | 0 | 0
Calibrated by : |  |  |  |  | Ms. Arunkamon Raramanus
| 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120011 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | : | LG
| Equipment Name |  |  |  | : | Depth Micrometer
| Manufacturer |  |  |  | : | Mitutoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 456
| ID. Number |  |  |  | : | 0
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-01-18
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2016-01-19
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2017-01-18
| Calibration Procedure |  |  |  | : | SP-CPD-04-16
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-20 |  |  | Approved by :
| Calibrated by |  | : | Ms. Arunkamon Raramanus |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120011 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR15120011 | [5 empty] | Page : 3 of 3
Flatness of The Spindle
|  | Optical Flat |  |  | Fringe Number
|  | 12 |  | mm | 1
|  | Measurement Results | [5 empty] | Unit : | mm
|  | Normal Value |  |  | UUC Reading | Error | Uncertainty (±) µm
|  | 0 |  |  | 0 | 0 | 0.58046
|  | 2.2 |  |  | 2.2 | 0 | 0.59041
|  | 4.8 |  |  | 4.8 | 0 | 0.59312
|  | 7.4 |  |  | 7.4 | 0 | 0.59782
|  | 10 |  |  | 10 | 0 | 0.59546
|  | 12.1 |  |  | 12.1 | 0 | 0.61323
|  | 14.7 |  |  | 14.7 | 0 | 0.62317
|  | 17.3 |  |  | 17.3 | 0 | 0.63483
|  | 19.9 |  |  | 19.9 | 0 | 0.64812
|  | 22.5 |  |  | 22.5 | 0 | 0.66295
|  | 25 |  |  | 25 | 0 | 0.66966
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Depth Micrometer
Nominal Value | Repeatability |  | Uncert of Gauge Block |  | Temperature Effect |  | Resolution of UUC |  | Uncert of Surface |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 6e-05 | 3e-05 | 0 | 0 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00029023 | 0.00085 | ∞ | 2 | 0.58046
2.2 | 0 | 0 | 0.00012 | 6e-05 | 2.53e-05 | 1.4607e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00029521 | 0.00085 | ∞ | 2 | 0.59041
4.8 | 0 | 0 | 0.00012 | 6e-05 | 5.52e-05 | 3.187e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00029656 | 0.00085 | ∞ | 2 | 0.59312
7.4 | 0 | 0 | 0.00012 | 6e-05 | 8.51e-05 | 4.9133e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00029891 | 0.00085 | ∞ | 2 | 0.59782
10 | 0 | 0 | 6e-05 | 3e-05 | 0.000115 | 6.6395e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00029773 | 0.00085 | ∞ | 2 | 0.59546
12.1 | 0 | 0 | 0.00013 | 6.5e-05 | 0.00013915 | 8.0338e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00030661 | 0.00085 | ∞ | 2 | 0.61323
14.7 | 0 | 0 | 0.00013 | 6.5e-05 | 0.00016905 | 9.7601e-05 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00031158 | 0.00085 | ∞ | 2 | 0.62317
17.3 | 0 | 0 | 0.00013 | 6.5e-05 | 0.00019895 | 0.00011486 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00031741 | 0.00085 | ∞ | 2 | 0.63483
19.9 | 0 | 0 | 0.00013 | 6.5e-05 | 0.00022885 | 0.00013213 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00032406 | 0.00085 | ∞ | 2 | 0.64812
22.5 | 0 | 0 | 0.00013 | 6.5e-05 | 0.00025875 | 0.00014939 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00033147 | 0.00085 | ∞ | 2 | 0.66295
25 | 0 | 0 | 7e-05 | 3.5e-05 | 0.0002875 | 0.00016599 | 0.0005 | 0.00028868 | 0.0017 | 0.00085 | 0.00033483 | 0.00085 | ∞ | 2 | 0.66966
Note on F4: Certificate of Calubration Gauge Block
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3
Note on M6: Divisor = 1

## Uncert of STD
Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Surface Plate)
SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012 |  |  |  |  | SP-SD-014
Due Date |  | 2016-05-13 |  |  | Due Date |  | 2016-11-27 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09 |  |  | Due Date |  | 2016-05-26
2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm | 450 x 600 mm | 1.7 | µm | 0.0017 | mm
5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[5 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[5 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[5 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[5 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[5 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[5 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[5 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[5 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[5 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[5 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[5 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[5 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[5 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[5 empty] | 14 | 0.07 | µm | 7e-05 | mm
[5 empty] | 15 | 0.07 | µm | 7e-05 | mm
[5 empty] | 16 | 0.07 | µm | 7e-05 | mm
[5 empty] | 17 | 0.07 | µm | 7e-05 | mm
[5 empty] | 18 | 0.07 | µm | 7e-05 | mm
[5 empty] | 19 | 0.07 | µm | 7e-05 | mm
[5 empty] | 20 | 0.07 | µm | 7e-05 | mm
[5 empty] | 21 | 0.07 | µm | 7e-05 | mm
[5 empty] | 22 | 0.07 | µm | 7e-05 | mm
[5 empty] | 23 | 0.07 | µm | 7e-05 | mm
[5 empty] | 24 | 0.07 | µm | 7e-05 | mm
[5 empty] | 25 | 0.07 | µm | 7e-05 | mm
[5 empty] | 50 | 0.09 | µm | 9e-05 | mm
[5 empty] | 75 | 0.1 | µm | 0.0001 | mm
[5 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on B4: Grade 0
Note on H4: Grade 0
Note on N4: Grade 0
Note on T4: Grade 0
Note on Z4: Grade 0
Note on D5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on P5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on V5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AB5: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on C6: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on E6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I6: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on K6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O6: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U6: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W6: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AA6: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on AC6: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน