## Data Record(0to2mm)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR15120045-1 | [9 empty] | Page | 1 | of | 1
[9 empty] | Receive Date : |  |  |  | 2016-01-12 |  |  |  |  | Calibration Date : |  |  |  |  | 2015-12-11
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Dial and Digital Indicator | [8 empty] | Location | [5 empty] | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Dial and Digital Indicator | [5 empty] | Manufacturer : |  |  |  | Mittutoyo
Model : |  | 123 | [6 empty] | Serial No. : |  |  | 456 | [7 empty] | ID No : |  | 789
Range : |  |  | 0 |  | to | 10 |  | mm |  | Resolution : |  |  | 0.001 |  | mm
Overall Inspection | [6 empty] | Good |  |  | Not Good
Referance Standard : | [5 empty] | SP-SD-023 | [9 empty] | Due Date :
Referance Standard : | [15 empty] | Due Date :
Repeatability
|  | Function |  |  | UUC Reading | [13 empty] | Repeatability |  |  |  | Max-Min
[5 empty] | X1 |  |  | X2 |  | X3 |  |  | X4 |  |  | X5
|  | Fast |  |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  |  | 0
|  | Slow |  |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
| Norminal Value |  |  | Cs (𝜇m) |  |  |  | True Value (mm) |  |  | UUC Reading (mm) | [14 empty] | Average (mm) |  | 𝜎n-1 (mm) |  | uA (mm) repeatability |  | Correction (mm)
[11 empty] | X1 |  |  |  | X2 |  |  |  | X3 |  |  |  | X4
|  |  |  | F |  | B |  | F | B |  | F |  | B |  | F |  | B |  | F |  | B |  | F | B |  | F | B | F | B | F | B | F | B
| 0 |  |  | 0 |  | 0.4 |  | 0 | 0.0044 |  | 0 |  | 0.004 |  | 0 |  | 0.004 |  | 0 |  | 0.004 |  | 0 | 0.004 |  | 0 | 0.004 | 0 | 0 | 0 | 0 | 0 | 0.0044
| 0.1 |  |  | -0.3 |  | 0.5 |  | 0.1017 | 0.1035 |  | 0.102 |  | 0.103 |  | 0.102 |  | 0.103 |  | 0.102 |  | 0.103 |  | 0.102 | 0.103 |  | 0.102 | 0.103 | 0 | 0 | 0 | 0 | 0.0017 | 0.0035
| 0.2 |  |  | -0.1 |  | 0.4 |  | 0.2019 | 0.2054 |  | 0.202 |  | 0.205 |  | 0.202 |  | 0.205 |  | 0.202 |  | 0.205 |  | 0.202 | 0.205 |  | 0.202 | 0.205 | 0 | 0 | 0 | 0 | 0.0019 | 0.0054
| 0.3 |  |  | -0.1 |  | 0.6 |  | 0.3009 | 0.3056 |  | 0.301 |  | 0.305 |  | 0.301 |  | 0.305 |  | 0.301 |  | 0.305 |  | 0.301 | 0.305 |  | 0.301 | 0.305 | 0 | 0 | 0 | 0 | 0.0009 | 0.0056
| 0.4 |  |  | -0.3 |  | 0.6 |  | 0.4027 | 0.4066 |  | 0.403 |  | 0.406 |  | 0.403 |  | 0.406 |  | 0.403 |  | 0.406 |  | 0.403 | 0.406 |  | 0.403 | 0.406 | 0 | 0 | 0 | 0 | 0.0027 | 0.0066
| 0.5 |  |  | -0.2 |  | 0.3 |  | 0.5018 | 0.5053 |  | 0.502 |  | 0.505 |  | 0.502 |  | 0.505 |  | 0.502 |  | 0.505 |  | 0.502 | 0.505 |  | 0.502 | 0.505 | 0 | 0 | 0 | 0 | 0.0018 | 0.0053
| 0.6 |  |  | -0.4 |  | 0.4 |  | 0.6036 | 0.6074 |  | 0.604 |  | 0.607 |  | 0.604 |  | 0.607 |  | 0.604 |  | 0.607 |  | 0.604 | 0.607 |  | 0.604 | 0.607 | 0 | 0 | 0 | 0 | 0.0036 | 0.0074
| 0.7 |  |  | -0.2 |  | 0.4 |  | 0.6998 | 0.7054 |  | 0.7 |  | 0.705 |  | 0.7 |  | 0.705 |  | 0.7 |  | 0.705 |  | 0.7 | 0.705 |  | 0.7 | 0.705 | 0 | 0 | 0 | 0 | -0.0002 | 0.0054
0.8
0.9
1
1.1
1.2
1.3
1.4
1.5
1.6
1.7
1.8
1.9

## Data Record(2to5mm)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR15120045-1 | [9 empty] | Page | 1 | of | 1
[9 empty] | Receive Date : |  |  |  | 2016-01-12 |  |  |  |  | Calibration Date : |  |  |  |  | 2015-12-11
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Dial and Digital Indicator | [8 empty] | Location | [5 empty] | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Dial and Digital Indicator | [5 empty] | Manufacturer : |  |  |  | Mittutoyo
Model : |  | 123 | [6 empty] | Serial No. : |  |  | 456 | [7 empty] | ID No : |  | 789
Range : |  |  | 0 |  | to | 10 |  | mm |  | Resolution : |  |  | 0.01 |  | mm
Overall Inspection | [6 empty] | Good |  |  | Not Good
Referance Standard : | [5 empty] | SP-SD-023 | [9 empty] | Due Date :
Referance Standard : | [15 empty] | Due Date :
| Norminal Value |  |  | Cs (𝜇m) |  |  |  | True Value (mm) |  |  | UUC Reading (mm) | [14 empty] | Average (mm) |  | 𝜎n-1 (mm) |  | uA (mm) repeatability |  | Correction (mm)
[11 empty] | X1 |  |  |  | X2 |  |  |  | X3 |  |  |  | X4
|  |  |  | F |  | B |  | F | B |  | F |  | B |  | F |  | B |  | F |  | B |  | F | B |  | F | B | F | B | F | B | F | B
| 2 |  |  | -0.7 |  | 0.9 |  | 2.0043 | 2.0089 |  | 2.005 |  | 2.008 |  | 2.005 |  | 2.008 |  | 2.005 |  | 2.008 |  | 2.005 | 2.008 |  | 2.005 | 2.008 | 0 | 0 | 0 | 0 | 0.0043 | 0.0089
| 2.5 |  |  | -0.8 |  | 0.8 |  | 2.5042 | 2.5088 |  | 2.505 |  | 2.508 |  | 2.505 |  | 2.508 |  | 2.505 |  | 2.508 |  | 2.505 | 2.508 |  | 2.505 | 2.508 | 0 | 0 | 0 | 0 | 0.0042 | 0.0088
| 3 |  |  | -0.8 |  | 1 |  | 3.0012 | 3.005 |  | 3.002 |  | 3.004 |  | 3.002 |  | 3.004 |  | 3.002 |  | 3.004 |  | 3.002 | 3.004 |  | 3.002 | 3.004 | 0 | 0 | 0 | 0 | 0.0012 | 0.005
| 3.5 |  |  | -0.8 |  | 0.8 |  | 3.5062 | 3.5108 |  | 3.507 |  | 3.51 |  | 3.507 |  | 3.51 |  | 3.507 |  | 3.51 |  | 3.507 | 3.51 |  | 3.507 | 3.51 | 0 | 0 | 0 | 0 | 0.0062 | 0.0108
| 4 |  |  | -0.8 |  | 0.8 |  | 4.0072 | 4.0098 |  | 4.008 |  | 4.009 |  | 4.008 |  | 4.009 |  | 4.008 |  | 4.009 |  | 4.008 | 4.009 |  | 4.008 | 4.009 | 0 | 0 | 0 | 0 | 0.0072 | 0.0098
| 4.5 |  |  | -0.8 |  | 0.8 |  | 4.505 | 4.5058 |  | 4.505 |  | 4.505 |  | 4.505 |  | 4.505 |  | 4.505 |  | 4.505 |  | 4.505 | 4.505 |  | 4.505 | 4.505 | 0 | 0 | 0 | 0 | 0.005 | 0.0058
| 5 |  |  | -0.8 |  | 1 |  | 5.0052 | 5.009 |  | 5.006 |  | 5.008 |  | 5.006 |  | 5.008 |  | 5.006 |  | 5.008 |  | 5.006 | 5.008 |  | 5.006 | 5.008 | 0 | 0 | 0 | 0 | 0.0052 | 0.009
Calibration by : |  |  |  |  | Ms. Arunkamon Raramanus
[5 empty] | 11 |  | Ms. Arunkamon Raramanus

## Data Record(5to50mm)
SP METROLOGY SYSTEM THAILAND | [8 empty] | Certificate No. : |  |  |  | SPR15120045-1 | [9 empty] | Page | 1 | of | 1
[9 empty] | Receive Date : |  |  |  | 2016-01-12 |  |  |  |  | Calibration Date : |  |  |  |  | 2015-12-11
Equipment Name | [8 empty] | Temp & Humiduty : |  |  |  |  | 20 | oC | 50 | %RH
Dial and Digital Indicator | [8 empty] | Location | [5 empty] | In Lab | [7 empty] | On Site
Customer Name : |  |  |  |  | SP
Equipment Name : |  |  |  |  | Dial and Digital Indicator | [5 empty] | Manufacturer : |  |  |  | Mittutoyo
Model : |  | 123 | [6 empty] | Serial No. : |  |  | 456 | [7 empty] | ID No : |  | 789
Range : |  |  | 0 |  | to | 10 |  | mm |  | Resolution : |  |  | 0.01 |  | mm
Overall Inspection | [6 empty] | Good |  |  | Not Good
Referance Standard : | [5 empty] | SP-SD-023 | [9 empty] | Due Date :
Referance Standard : | [15 empty] | Due Date :
| Norminal Value |  |  | Cs (𝜇m) |  |  |  | True Value (mm) |  |  | UUC Reading (mm) | [14 empty] | Average (mm) |  | 𝜎n-1 (mm) |  | uA (mm) repeatability |  | Correction (mm)
[11 empty] | X1 |  |  |  | X2 |  |  |  | X3 |  |  |  | X4
|  |  |  | F |  | B |  | F | B |  | F |  | B |  | F |  | B |  | F |  | B |  | F | B |  | F | B | F | B | F | B | F | B
| 6 |  |  | -0.8 |  | 0.9 |  | 6.0082 | 6.0149 |  | 6.009 |  | 6.014 |  | 6.009 |  | 6.014 |  | 6.009 |  | 6.014 |  | 6.009 | 6.014 |  | 6.009 | 6.014 | 0 | 0 | 0 | 0 | 0.0082 | 0.0149
| 7 |  |  | -0.8 |  | 0.5 |  | 7.0062 | 7.007 |  | 7.007 |  | 7.01 |  | 7.007 |  | 7.01 |  | 7.007 |  | 7.01 |  | 7.007 | 7.01 |  | 7.007 | 7.01 | 0 | 0 | 0 | 0 | 0.0062 | 0.007
| 8 |  |  | -0.7 |  | 0.3 |  | 8.0073 | 8.0093 |  | 8.008 |  | 8.009 |  | 8.008 |  | 8.009 |  | 8.008 |  | 8.009 |  | 8.008 | 8.009 |  | 8.008 | 8.009 | 0 | 0 | 0 | 0 | 0.0073 | 0.0093
| 9 |  |  | -0.7 |  | 0.5 |  | 9.0113 | 9.0135 |  | 9.012 |  | 9.013 |  | 9.012 |  | 9.013 |  | 9.012 |  | 9.013 |  | 9.012 | 9.013 |  | 9.012 | 9.013 | 0 | 0 | 0 | 0 | 0.0113 | 0.0135
| 10 |  |  | -0.6 |  | 0.5 |  | 10.0124 | 10.0155 |  | 10.013 |  | 10.015 |  | 10.013 |  | 10.015 |  | 10.013 |  | 10.015 |  | 10.013 | 10.015 |  | 10.013 | 10.015 | 0 | 0 | 0 | 0 | 0.0124 | 0.0155
| 11 |  |  | -0.5 |  | 0.4 |  | 11.0095 | 11.0144 |  | 11.01 |  | 11.014 |  | 11.01 |  | 11.014 |  | 11.01 |  | 11.014 |  | 11.01 | 11.014 |  | 11.01 | 11.014 | 0 | 0 | 0 | 0 | 0.0095 | 0.0144
| 12 |  |  | -0.4 |  | 0.4 |  | 0.6036 | 0.6074 |  | 0.604 |  | 0.607 |  | 0.604 |  | 0.607 |  | 0.604 |  | 0.607 |  | 0.604 | 0.607 |  | 0.604 | 0.607 | 0 | 0 | 0 | 0 | -11.3964 | -11.3926
| 13 |  |  | -0.2 |  | 0.4 |  | 0.6998 | 0.7054 |  | 0.7 |  | 0.705 |  | 0.7 |  | 0.705 |  | 0.7 |  | 0.705 |  | 0.7 | 0.705 |  | 0.7 | 0.705 | 0 | 0 | 0 | 0 | -12.3002 | -12.2946
14
15
16
17
18
19
20
25
30
35
40
45
50
Calibration by :
| Calibration by : |  |  |  |  | Ms. Arunkamon Raramanus
[6 empty] | 11 |  | Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120045-1 | [7 empty] | Page : 1 of 3
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Dial and Digital Indicator
| Manufacturer |  |  |  | : | Mittutoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 456
| ID. Number |  |  |  | : | 789
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-01-12
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2015-12-11
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2016-12-10
| Calibration Procedure |  |  |  | : | SP-CPT-04-11
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2015-12-12 |  |  | Approved by :
| Calibrated by |  | : | #REF! |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120045-1 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Gauge Block Set |  |  | N/A | 60711 | MTL142959-2 | 2015-11-28
Traceability
This certification is traceable to the International System of Unit maintained at :
-National Institute of Metrology (Thailand) (NIMT)
-Thailand Institute of Scientific And Technological Research (TISTR)

## Result
Result of Calibration
| Certificate No. : |  |  | #REF! | [5 empty] | Page : 3 of 3
| Repeatability |  | 0 |  | mm
[8 empty] | Unit : | mm
|  | Nominal Value |  |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  | 0 |  |  | 0.004 | 0.004 | 0.001701
|  | 0.1 |  |  | 0.103 | 0.103 | 0.001701
|  | 0.2 |  |  | 0.205 | 0.205 | 0.001701
|  | 0.3 |  |  | 0.305 | 0.305 | 0.001701
|  | 0.4 |  |  | 0.406 | 0.406 | 0.001701
|  | 0.5 |  |  | 0.505 | 0.505 | 0.001701
|  | 0.6 |  |  | 0.607 | 0.607 | 0.001701
|  | 0.7 |  |  | 0.705 | 0.705 | 0.001701
|  | 0.8 |  |  | 0 | - | 0.001701
|  | 0.9 |  |  | 0 | - | 0.001701
|  | 1 |  |  | 0 | - | 0.001701
|  | 1.1 |  |  | 0 | - | 0.001701
|  | 1.2 |  |  | 0 | - | 0.0017011
|  | 1.3 |  |  | 0 | - | 0.0017011
|  | 1.4 |  |  | 0 | - | 0.0017011
|  | 1.5 |  |  | 0 | - | 0.0017011
|  | 1.6 |  |  | 0 | - | 0.0017011
|  | 1.7 |  |  | 0 | - | 0.0017011
|  | 1.8 |  |  | 0 | - | 0.0017011
|  | 1.9 |  |  | 0 | - | 0.0017012
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Result (2)
Result of Calibration
| Certificate No. : |  | #REF! | [5 empty] | Page : 3 of 3
[7 empty] | Unit : | mm
|  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± ) µm
|  | 2 |  | 2.008 | 0.008 | 0.0017012 |  |  |  | Data Record(2to5mm)
|  | 2.5 |  | 2.508 | 0.008 | 0.0017013
|  | 3 |  | 3.004 | 0.004 | 0.0017014
|  | 3.5 |  | 3.51 | 0.01 | 0.0017016
|  | 4 |  | 4.009 | 0.009 | 0.0017018
|  | 4.5 |  | 4.505 | 0.005 | 0.001702
|  | 5 |  | 5.008 | 0.008 | 0.0017023
|  | 6 |  | 6.014 | -3.992 | 0.0017028 |  |  |  | Data Record(5to50mm)
|  | 7 |  | 7.01 | -4.492 | 0.0017035
|  | 8 |  | 8.009 | -4.996 | 0.0017043
|  | 9 |  | 9.013 | -5.49 | 0.0017052
|  | 10 |  | 10.015 | -5.991 | 0.0017062
|  | 11 |  | 11.014 | -6.495 | 0.0017072
|  | 12 |  | 0.607 | -6.992 | 0.0017084
|  | 13 |  | 0.705 | -6.986 | 0.0017097
|  | 14 |  | - | - | 0.0017111
|  | 15 |  | - | - | 0.0017126
|  | 16 |  | - | - | 0.0017142
|  | 17 |  | - | - | 0.0017159
|  | 18 |  | - | - | 0.0017177
|  | 19 |  | - | - | 0.0017196
|  | 20 |  | - | - | 0.0017216
|  | 25 |  | - | - | 0.0017331

## Result (3)
Result of Calibration
| Certificate No. : |  | #REF! | [5 empty] | Page : 3 of 3
[7 empty] | Unit : | mm
|  | Nominal Value |  | UUC Reading | Error | Uncertainty ( ± )
|  | 30 |  | - | - | 0.001747
|  | 35 |  | - | - | 0.0017633
|  | 40 |  | - | - | 0.001782
|  | 45 |  | - | - | 0.0018029
|  | 50 |  | - | - | 0.001826
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Dial and Digital Indicator
Nominal Value | Repeatability |  | Uncertainty of STD |  | Temperature Effect |  | Resolution of UUC |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 | 0 | 0 | 0.0008 | 0.0008 | 0 | 0 | 0.0005 | 0.00028868 | 0.00085049 | 0 | ∞ | 2 | 0.001701
0.1 | 0 | 0 | 0.0008 | 0.0008 | 1.15e-06 | 6.6395e-07 | 0.0005 | 0.00028868 | 0.00085049 | 0 | ∞ | 2 | 0.001701
0.2 | 0 | 0 | 0.0008 | 0.0008 | 2.3e-06 | 1.3279e-06 | 0.0005 | 0.00028868 | 0.00085049 | 0 | ∞ | 2 | 0.001701
0.3 | 0 | 0 | 0.0008 | 0.0008 | 3.45e-06 | 1.9919e-06 | 0.0005 | 0.00028868 | 0.00085049 | 0 | ∞ | 2 | 0.001701
0.4 | 0 | 0 | 0.0008 | 0.0008 | 4.6e-06 | 2.6558e-06 | 0.0005 | 0.00028868 | 0.00085049 | 0 | ∞ | 2 | 0.001701
0.5 | 0 | 0 | 0.0008 | 0.0008 | 5.75e-06 | 3.3198e-06 | 0.0005 | 0.00028868 | 0.0008505 | 0 | ∞ | 2 | 0.001701
0.6 | 0 | 0 | 0.0008 | 0.0008 | 6.9e-06 | 3.9837e-06 | 0.0005 | 0.00028868 | 0.0008505 | 0 | ∞ | 2 | 0.001701
0.7 | 0 | 0 | 0.0008 | 0.0008 | 8.05e-06 | 4.6477e-06 | 0.0005 | 0.00028868 | 0.0008505 | 0 | ∞ | 2 | 0.001701
0.8 | 0 | 0 | 0.0008 | 0.0008 | 9.2e-06 | 5.3116e-06 | 0.0005 | 0.00028868 | 0.00085051 | 0 | ∞ | 2 | 0.001701
0.9 | 0 | 0 | 0.0008 | 0.0008 | 1.035e-05 | 5.9756e-06 | 0.0005 | 0.00028868 | 0.00085051 | 0 | ∞ | 2 | 0.001701
1 | 0 | 0 | 0.0008 | 0.0008 | 1.15e-05 | 6.6395e-06 | 0.0005 | 0.00028868 | 0.00085052 | 0 | ∞ | 2 | 0.001701
1.1 | 0 | 0 | 0.0008 | 0.0008 | 1.265e-05 | 7.3035e-06 | 0.0005 | 0.00028868 | 0.00085052 | 0 | ∞ | 2 | 0.001701
1.2 | 0 | 0 | 0.0008 | 0.0008 | 1.38e-05 | 7.9674e-06 | 0.0005 | 0.00028868 | 0.00085053 | 0 | ∞ | 2 | 0.0017011
1.3 | 0 | 0 | 0.0008 | 0.0008 | 1.495e-05 | 8.6314e-06 | 0.0005 | 0.00028868 | 0.00085053 | 0 | ∞ | 2 | 0.0017011
1.4 | 0 | 0 | 0.0008 | 0.0008 | 1.61e-05 | 9.2953e-06 | 0.0005 | 0.00028868 | 0.00085054 | 0 | ∞ | 2 | 0.0017011
1.5 | 0 | 0 | 0.0008 | 0.0008 | 1.725e-05 | 9.9593e-06 | 0.0005 | 0.00028868 | 0.00085055 | 0 | ∞ | 2 | 0.0017011
1.6 | 0 | 0 | 0.0008 | 0.0008 | 1.84e-05 | 1.0623e-05 | 0.0005 | 0.00028868 | 0.00085056 | 0 | ∞ | 2 | 0.0017011
1.7 | 0 | 0 | 0.0008 | 0.0008 | 1.955e-05 | 1.1287e-05 | 0.0005 | 0.00028868 | 0.00085056 | 0 | ∞ | 2 | 0.0017011
1.8 | 0 | 0 | 0.0008 | 0.0008 | 2.07e-05 | 1.1951e-05 | 0.0005 | 0.00028868 | 0.00085057 | 0 | ∞ | 2 | 0.0017011
1.9 | 0 | 0 | 0.0008 | 0.0008 | 2.185e-05 | 1.2615e-05 | 0.0005 | 0.00028868 | 0.00085058 | 0 | ∞ | 2 | 0.0017012
2 | 0 | 0 | 0.0008 | 0.0008 | 2.3e-05 | 1.3279e-05 | 0.0005 | 0.00028868 | 0.00085059 | 0 | ∞ | 2 | 0.0017012
2.5 | 0 | 0 | 6e-05 | 0.0008 | 2.875e-05 | 1.6599e-05 | 0.0005 | 0.00028868 | 0.00085065 | 0 | ∞ | 2 | 0.0017013
3 | 0 | 0 | 6e-05 | 0.0008 | 3.45e-05 | 1.9919e-05 | 0.0005 | 0.00028868 | 0.00085072 | 0 | ∞ | 2 | 0.0017014
3.5 | 0 | 0 | 0.00012 | 0.0008 | 4.025e-05 | 2.3238e-05 | 0.0005 | 0.00028868 | 0.00085081 | 0 | ∞ | 2 | 0.0017016
4 | 0 | 0 | 6e-05 | 0.0008 | 4.6e-05 | 2.6558e-05 | 0.0005 | 0.00028868 | 0.0008509 | 0 | ∞ | 2 | 0.0017018
4.5 | 0 | 0 | 0.00012 | 0.0008 | 5.175e-05 | 2.9878e-05 | 0.0005 | 0.00028868 | 0.00085101 | 0 | ∞ | 2 | 0.001702
5 | 0 | 0 | 6e-05 | 0.0008 | 5.75e-05 | 3.3198e-05 | 0.0005 | 0.00028868 | 0.00085114 | 0 | ∞ | 2 | 0.0017023
6 | 0 | 0 | 0.00012 | 0.0008 | 6.9e-05 | 3.9837e-05 | 0.0005 | 0.00028868 | 0.00085142 | 0 | ∞ | 2 | 0.0017028
7 | 0 | 0 | 6e-05 | 0.0008 | 8.05e-05 | 4.6477e-05 | 0.0005 | 0.00028868 | 0.00085176 | 0 | ∞ | 2 | 0.0017035
8 | 0 | 0 | 6e-05 | 0.0008 | 9.2e-05 | 5.3116e-05 | 0.0005 | 0.00028868 | 0.00085215 | 0 | ∞ | 2 | 0.0017043
9 | 0 | 0 | 6e-05 | 0.0008 | 0.0001035 | 5.9756e-05 | 0.0005 | 0.00028868 | 0.00085259 | 0 | ∞ | 2 | 0.0017052
10 | 0 | 0 | 6e-05 | 0.0008 | 0.000115 | 6.6395e-05 | 0.0005 | 0.00028868 | 0.00085308 | 0 | ∞ | 2 | 0.0017062
11 | 0 | 0 | 7e-05 | 0.0008 | 0.0001265 | 7.3035e-05 | 0.0005 | 0.00028868 | 0.00085362 | 0 | ∞ | 2 | 0.0017072
12 | 0 | 0 | 7e-05 | 0.0008 | 0.000138 | 7.9674e-05 | 0.0005 | 0.00028868 | 0.00085421 | 0 | ∞ | 2 | 0.0017084
13 | 0 | 0 | 7e-05 | 0.0008 | 0.0001495 | 8.6314e-05 | 0.0005 | 0.00028868 | 0.00085486 | 0 | ∞ | 2 | 0.0017097
14 | 0 | 0 | 7e-05 | 0.0008 | 0.000161 | 9.2953e-05 | 0.0005 | 0.00028868 | 0.00085555 | 0 | ∞ | 2 | 0.0017111
15 | 0 | 0 | 7e-05 | 0.0008 | 0.0001725 | 9.9593e-05 | 0.0005 | 0.00028868 | 0.0008563 | 0 | ∞ | 2 | 0.0017126
16 | 0 | 0 | 7e-05 | 0.0008 | 0.000184 | 0.00010623 | 0.0005 | 0.00028868 | 0.0008571 | 0 | ∞ | 2 | 0.0017142
17 | 0 | 0 | 7e-05 | 0.0008 | 0.0001955 | 0.00011287 | 0.0005 | 0.00028868 | 0.00085795 | 0 | ∞ | 2 | 0.0017159
18 | 0 | 0 | 7e-05 | 0.0008 | 0.000207 | 0.00011951 | 0.0005 | 0.00028868 | 0.00085885 | 0 | ∞ | 2 | 0.0017177
19 | 0 | 0 | 7e-05 | 0.0008 | 0.0002185 | 0.00012615 | 0.0005 | 0.00028868 | 0.00085979 | 0 | ∞ | 2 | 0.0017196
20 | 0 | 0 | 7e-05 | 0.0008 | 0.00023 | 0.00013279 | 0.0005 | 0.00028868 | 0.00086079 | 0 | ∞ | 2 | 0.0017216
25 | 0 | 0 | 7e-05 | 0.0008 | 0.0002875 | 0.00016599 | 0.0005 | 0.00028868 | 0.00086654 | 0 | ∞ | 2 | 0.0017331
30 | 0 | 0 | 0.00011 | 0.0008 | 0.000345 | 0.00019919 | 0.0005 | 0.00028868 | 0.0008735 | 0 | ∞ | 2 | 0.001747
35 | 0 | 0 | 0.00017 | 0.0008 | 0.0004025 | 0.00023238 | 0.0005 | 0.00028868 | 0.00088167 | 0 | ∞ | 2 | 0.0017633
40 | 0 | 0 | 0.0002 | 0.0008 | 0.00046 | 0.00026558 | 0.0005 | 0.00028868 | 0.00089099 | 0 | ∞ | 2 | 0.001782
45 | 0 | 0 | 0.00018 | 0.0008 | 0.0005175 | 0.00029878 | 0.0005 | 0.00028868 | 0.00090144 | 0 | ∞ | 2 | 0.0018029
50 | 0 | 0 | 0.00013 | 0.0008 | 0.000575 | 0.00033198 | 0.0005 | 0.00028868 | 0.00091299 | 0 | ∞ | 2 | 0.001826
Note on F4: Certificate of Calubration Gauge Block
Note on H4: (Normal Value * 11.5*10^-6 * 1)
Note on E6: Divisor = 1
Note on G6: Divisor = 2
Note on I6: Divisor = √3
Note on K6: Divisor = √3

## Cert of STD
Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | Certificate of Calibration (Gauge Block) |  |  |  |  | SP-SD-023
SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012 |  |  |  |  | Due Date |  | 2016-04-28
Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09 |  |  | 1 | mm | 0.0008 | mm
2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.37 | µm | 0.00037 | mm | 2 | mm | 0.0008 | mm
5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.39 | µm | 0.00039 | mm | 3 | mm | 0.0008 | mm
7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.41 | µm | 0.00041 | mm | 4 | mm | 0.0008 | mm
10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.43 | µm | 0.00043 | mm | 5 | mm | 0.0008 | mm
12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.49 | µm | 0.00049 | mm | 10 | mm | 0.0008 | mm
15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.57 | µm | 0.00057 | mm | 15 | mm | 0.0008 | mm
17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.74 | µm | 0.00074 | mm | 20 | mm | 0.0008 | mm
20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 0.91 | µm | 0.00091 | mm | 30 | mm | 0.0008 | mm
22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm | [5 empty] | 40 | mm | 0.0008 | mm
25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm | [5 empty] | 50 | mm | 0.0008 | mm
[5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
[5 empty] | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
[5 empty] | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[10 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[10 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[10 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[10 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[10 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[10 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[10 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[10 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[10 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[10 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[10 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[10 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[10 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[10 empty] | 14 | 0.07 | µm | 7e-05 | mm
[10 empty] | 15 | 0.07 | µm | 7e-05 | mm
[10 empty] | 16 | 0.07 | µm | 7e-05 | mm
[10 empty] | 17 | 0.07 | µm | 7e-05 | mm
[10 empty] | 18 | 0.07 | µm | 7e-05 | mm
[10 empty] | 19 | 0.07 | µm | 7e-05 | mm
[10 empty] | 20 | 0.07 | µm | 7e-05 | mm
[10 empty] | 21 | 0.07 | µm | 7e-05 | mm
[10 empty] | 22 | 0.07 | µm | 7e-05 | mm
[10 empty] | 23 | 0.07 | µm | 7e-05 | mm
[10 empty] | 24 | 0.07 | µm | 7e-05 | mm
[10 empty] | 25 | 0.07 | µm | 7e-05 | mm
[10 empty] | 50 | 0.09 | µm | 9e-05 | mm
[10 empty] | 75 | 0.1 | µm | 0.0001 | mm
[10 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on AF2: Universal Length Measuring
Note on B3: Grade 0
Note on H3: Grade 0
Note on N3: Grade 0
Note on T3: Grade 0
Note on Z3: Grade 0
Note on AH3: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on D4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on P4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on V4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AB4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AF4: Normal Value (mm)
Note on AH4: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on C5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on E5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on K5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AA5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AC5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน