## Data Record
SP METROLOGY SYSTEM THAILAND | [9 empty] | Certificate No. : |  |  | SPR15120066-1 | [7 empty] | Page |  | 1 | of | 1
[10 empty] | Receive Date : |  |  | 2016-01-01 |  |  |  |  | Calibration Date : |  | 2016-01-02
Equipment Name | [9 empty] | Temp & Humiduty : |  |  |  | 20 | oC | 50 | %RH
Depth Micro Checker | [9 empty] | Location | [5 empty] | In Lab |  |  |  |  | On Site
Customer Name : |  |  |  |  | SP METROLOGY SYSTEM (THAILAND) CO.,LTD.
Equipment Name : |  |  |  |  | Depth Micro Checker | [5 empty] | Manufacturer : |  |  |  | Mitutoyo
Model : |  | 123 | [6 empty] | Serial No. : |  | 456 |  |  |  |  | ID No : |  | 789
Range : |  | 0 |  | to | 300 |  | mm |  | Resolution : |  | 0.01 |  | mm.
Overall Inspection | [6 empty] | Good |  |  |  |  | Not Good
Referance Standard : | [14 empty] | Due Date :
Referance Standard : | [14 empty] | Due Date :
Flatness of Anvil Block
| Optical Flat | [6 empty] | Fringe Number
| 12 |  |  |  |  | mm |  | 1
Accuracy of Anvil Block
Nominal Lengh |  |  | Position Point |  |  | STD Reading | [10 empty] | Average |  | Error |  |  | Repeatability
[6 empty] | X1 |  |  | X2 |  | X3 |  |  | X4
25 |  |  | C |  |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 25 |  | 0 |  |  | 2.7386e-05
|  |  | 1 |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 25.0001 |  | 0.0001
|  |  | 2 |  |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 25 |  | 0
|  |  | 3 |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 25.0001 |  | 0.0001
|  |  | 4 |  |  | 0 |  |  | 0 |  | 0 |  |  | 0 |  |  | 25 |  | 0
Deviation of length
Nominal Value |  |  | Positions |  |  | Standard Reading | [10 empty] | Average |  | Error |  | Repeatability
[6 empty] | X1 |  |  | X2 |  | X3 |  |  | X4
100 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 100.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 99.9999 |  | -0.0001
125 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 125.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 124.9999 |  | -0.0001
150 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 150.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 149.9999 |  | -0.0001
175 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 175.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 174.9999 |  | -0.0001
200 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 200.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 199.9999 |  | -0.0001
250 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 250.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 249.9999 |  | -0.0001
300 |  |  | Left |  |  | 0.0001 |  |  | 0.0001 |  | 0.0001 |  |  | 0.0001 |  |  | 300.0001 |  | 0.0001 |  | 0
|  |  | Right |  |  | -0.0001 |  |  | -0.0001 |  | -0.0001 |  |  | -0.0001 |  |  | 299.9999 |  | -0.0001
Calibrated By : |  |  |  |  | Ms. Arunkamon Raramanus
Mr.Prayuth Thongauam
Mr. Vichan Ananta
Mr.Kittikorn Kingmali
Ms. Arunkamon Raramanus

## Certificate
Certificate of Calibration
| Certificate Number |  |  |  | : | SPR15120066-1 | [7 empty] | Page : 1 of 4
| Customer |  |  |  | :
| Equipment Name |  |  |  | : | Depth Micro Checker
| Manufacturer |  |  |  | : | Mitutoyo
| Model |  |  |  | : | 123
| Serial Number |  |  |  | : | 456
| ID. Number |  |  |  | : | 789
Environmental Conditions
| Ambient Temperature |  |  |  | : | 20 °C ± 1 °C | Received Date |  |  |  |  | : | 2016-01-01
| Relative Humidity |  |  |  | : | 50% ± 15 % | Calibration Date |  |  |  |  | : | 2016-01-02
| Location of Calibration |  |  |  | : | In-Lab | Recommended Due Date |  |  |  |  | : | 2017-01-01
| Calibration Procedure |  |  |  | : | SP-CPT-04-15
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Date of Issue |  | : | 2016-01-03 |  |  | Approved by :
| Calibrated by |  | : | 0.02 |  |  |  | 3 | ( Mr. Natthaphol Boonmee )
Authorized Signatory
| 11 | Ms. Arunkamon Raramanus | [7 empty] | 1 | Mr.Sombut Srikampa
[10 empty] | 3 | Mr. Natthaphol Boonmee
Note on Q36: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR15120066-1 |  |  |  | Page : 2 of 4
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Linear Height Master |  |  | TESA-HITE M+ | 5C062602 | MTO 150453-2 | 2016-06-26
| Long Gauge Block Set |  |  | ML8 | 11083P | CC558-09-0010 | 2016-09-07
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  |  |  | SPR15120066-1 |  |  |  |  | Page : 3 of 4
Flatness of Anvil Block
| Optical Flat |  |  |  |  | Fringe Number
|  |  | 12 | mm |  | 1
|  | Anvil Block | [6 empty] | Unit : | mm
|  | Nominal Lengh |  |  | Position Point |  | STD Reading | Error | Uncertainty ( ± ) µm
|  | 25 |  |  | C |  | 150.0001 | -0.0001 | 2.0289
[5 empty] | 1 |  | 150 | 0
[5 empty] | 2 |  | 149.9999 | 0.00015
[5 empty] | 3 |  | 150 | 5e-05
[5 empty] | 4 |  | 150.0001 | -0.0001

## Result (2)
Result of Calibration
|  | Certificate No. : |  | SPR15120066-1 |  |  |  | Page : 4 of 4
[7 empty] | Unit : | mm
|  |  | Nominal Value | Position | STD Reading | Error | Uncertainty ( ± ) µm
|  |  | 100 | Left | 100.0001 | 0.0001 | 2.4081
|  |  |  | Right | 99.9999 | -0.0001
|  |  | 125 | Left | 125.0001 | 0.0001 | 2.6334
|  |  |  | Right | 124.9999 | -0.0001
|  |  | 150 | Left | 150.0001 | 0.0001 | 2.8621
|  |  |  | Right | 149.9999 | -0.0001
|  |  | 175 | Left | 175.0001 | 0.0001 | 3.1086
|  |  |  | Right | 174.9999 | -0.0001
|  |  | 200 | Left | 200.0001 | 0.0001 | 3.3703
|  |  |  | Right | 199.9999 | -0.0001
|  |  | 250 | Left | 250.0001 | 0.0001 | 3.927
|  |  |  | Right | 249.9999 | -0.0001
|  |  | 300 | Left | 300.0001 | 0.0001 | 4.5141
|  |  |  | Right | 299.9999 | -0.0001
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying the
standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95 %.
- End of Certificate -

## Uncertainty Budget
Uncertainty Budget of Anvil Block
Nominal Value | Repeatability |  | Uncertainty of Lenier |  | Contact Deformation |  | Temperature Effect |  | Resolution of STD |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
25 | 2.7386e-05 | 2.7386e-05 | 0.002 | 0.001 | 0 | 0 | 0.0002875 | 0.00016599 | 5e-05 | 2.8868e-05 | 0.0010145 | 2.7386e-05 | 5648638.4312 | 2 | 2.0289
Contact Deformation
[5 empty] | STD | UUC
[5 empty] | Steel | Steel | 0 | µm
[5 empty] | Steel | Ceramic | 0.00575 | µm
[5 empty] | Steel | Tungsten | -0.2026 | µm
Uncertainty Budget of Depth Micro Checker
Nominal Value | Repeatability |  | Uncertainty of Lenier |  | Uncertainty of GB |  | Contact Deformation |  | Temperature Effect |  | Resolution of STD |  | Uc | Ui | Veff | K95 | U95%
(mm) | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) |  | (mm) | [5 empty] | (µm)
Value | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
100 | 0 | 0 | 0.002 | 0.001 | 0.00018 | 9e-05 | 0 | 0 | 0.00115 | 0.00066395 | 5e-05 | 2.8868e-05 | 0.0012041 | 0 | ∞ | 2 | 2.4081
125 | 0 | 0 | 0.002 | 0.001 | 0.00042 | 0.00021 | 0 | 0 | 0.0014375 | 0.00082994 | 5e-05 | 2.8868e-05 | 0.0013167 | 0 | ∞ | 2 | 2.6334
150 | 0 | 0 | 0.002 | 0.001 | 0.00047 | 0.000235 | 0 | 0 | 0.001725 | 0.00099593 | 5e-05 | 2.8868e-05 | 0.0014311 | 0 | ∞ | 2 | 2.8621
175 | 0 | 0 | 0.002 | 0.001 | 0.00051 | 0.000255 | 0 | 0 | 0.0020125 | 0.0011619 | 5e-05 | 2.8868e-05 | 0.0015543 | 0 | ∞ | 2 | 3.1086
200 | 0 | 0 | 0.002 | 0.001 | 0.00055 | 0.000275 | 0 | 0 | 0.0023 | 0.0013279 | 5e-05 | 2.8868e-05 | 0.0016852 | 0 | ∞ | 2 | 3.3703
250 | 0 | 0 | 0.002 | 0.001 | 0.00063 | 0.000315 | 0 | 0 | 0.002875 | 0.0016599 | 5e-05 | 2.8868e-05 | 0.0019635 | 0 | ∞ | 2 | 3.927
300 | 0 | 0 | 0.002 | 0.001 | 0.00071 | 0.000355 | 0 | 0 | 0.00345 | 0.0019919 | 5e-05 | 2.8868e-05 | 0.0022571 | 0 | ∞ | 2 | 4.5141
Contact Deformation
[7 empty] | STD | UUC
[7 empty] | Steel | Steel | 0 | µm
[7 empty] | Steel | Ceramic | 0.00575 | µm
[7 empty] | Steel | Tungsten | -0.2026 | µm
Note on F5: Certificate of Calubration Linear Height Gauge
Note on J5: (Normal Value * 11.5*10^-6 * 1)
Note on E7: Divisor = 1
Note on G7: Divisor = 2
Note on I7: Divisor = √3
Note on K7: Divisor = √3
Note on M7: Divisor = √3
Note on F17: Certificate of Calubration Linear Height Gauge
Note on H17: Certificate of Calubration Gauge Block
Note on L17: (Normal Value * 11.5*10^-6 * 1)
Note on E19: Divisor = 1
Note on G19: Divisor = 2
Note on I19: Divisor = √3
Note on K19: Divisor = √3
Note on M19: Divisor = √3
Note on O19: Divisor = √3

## Uncert of STD
Certificate of Calibration (Linear) |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration |  |  |  |  | Certificate of Calibration
SP-SD-006 |  |  |  |  | SP-SD-008 |  |  |  |  | SP-SD-009 |  |  |  |  | SP-SD-010 |  |  |  |  | SP-SD-011 |  |  |  |  | SP-SD-012
Due Date |  | 2016-06-26 |  |  | Due Date |  | 2015-11-29 |  |  | Due Date |  | 2016-05-13 |  |  | Due Date |  | 2016-11-27 |  |  | Due Date |  | 2016-05-12 |  |  | Due Date |  | 2016-06-09
0 | 2 | µm | 0.002 | mm | 2.5 | 0.06 | µm | 6e-05 | mm | 2.5 | 0.08 | µm | 8e-05 | mm | 1.005 | 0.06 | µm | 6e-05 | mm | 1 | 0.08 | µm | 8e-05 | mm | 125 | 0.42 | µm | 0.00042 | mm
20 | 2 | µm | 0.002 | mm | 5.1 | 0.06 | µm | 6e-05 | mm | 5.1 | 0.09 | µm | 9e-05 | mm | 1.01 | 0.06 | µm | 6e-05 | mm | 1.005 | 0.08 | µm | 8e-05 | mm | 150 | 0.47 | µm | 0.00047 | mm
50 | 2 | µm | 0.002 | mm | 7.7 | 0.06 | µm | 6e-05 | mm | 7.7 | 0.09 | µm | 9e-05 | mm | 1.02 | 0.06 | µm | 6e-05 | mm | 1.01 | 0.08 | µm | 8e-05 | mm | 175 | 0.51 | µm | 0.00051 | mm
100 | 2 | µm | 0.002 | mm | 10.3 | 0.07 | µm | 7e-05 | mm | 10.3 | 0.09 | µm | 9e-05 | mm | 1.03 | 0.06 | µm | 6e-05 | mm | 1.02 | 0.08 | µm | 8e-05 | mm | 200 | 0.55 | µm | 0.00055 | mm
150 | 2 | µm | 0.002 | mm | 12.9 | 0.07 | µm | 7e-05 | mm | 12.9 | 0.09 | µm | 9e-05 | mm | 1.04 | 0.06 | µm | 6e-05 | mm | 1.03 | 0.08 | µm | 8e-05 | mm | 250 | 0.63 | µm | 0.00063 | mm
200 | 2 | µm | 0.002 | mm | 15 | 0.07 | µm | 7e-05 | mm | 15 | 0.1 | µm | 0.0001 | mm | 1.05 | 0.06 | µm | 6e-05 | mm | 1.04 | 0.08 | µm | 8e-05 | mm | 300 | 0.71 | µm | 0.00071 | mm
250 | 2 | µm | 0.002 | mm | 17.6 | 0.07 | µm | 7e-05 | mm | 17.6 | 0.1 | µm | 0.0001 | mm | 1.06 | 0.06 | µm | 6e-05 | mm | 1.05 | 0.08 | µm | 8e-05 | mm | 400 | 0.89 | µm | 0.00089 | mm
300 | 2 | µm | 0.002 | mm | 20.2 | 0.07 | µm | 7e-05 | mm | 20.2 | 0.1 | µm | 0.0001 | mm | 1.07 | 0.06 | µm | 6e-05 | mm | 1.06 | 0.08 | µm | 8e-05 | mm | 500 | 1.1 | µm | 0.0011 | mm
400 | 2 | µm | 0.002 | mm | 22.8 | 0.07 | µm | 7e-05 | mm | 22.8 | 0.1 | µm | 0.0001 | mm | 1.08 | 0.06 | µm | 6e-05 | mm | 1.07 | 0.08 | µm | 8e-05 | mm
500 | 2 | µm | 0.002 | mm | 25 | 0.07 | µm | 7e-05 | mm | 25 | 0.11 | µm | 0.00011 | mm | 1.09 | 0.06 | µm | 6e-05 | mm | 1.08 | 0.08 | µm | 8e-05 | mm
600 | 2 | µm | 0.002 | mm | [5 empty] | 50 | 0.13 | µm | 0.00013 | mm | 1.1 | 0.06 | µm | 6e-05 | mm | 1.09 | 0.08 | µm | 8e-05 | mm
700 | 2 | µm | 0.002 | mm | Digital | 2 |  |  |  | 75 | 0.16 | µm | 0.00016 | mm | 1.2 | 0.06 | µm | 6e-05 | mm | 1.1 | 0.08 | µm | 8e-05 | mm
[5 empty] | Analog | 5 |  |  |  | 100 | 0.18 | µm | 0.00018 | mm | 1.3 | 0.06 | µm | 6e-05 | mm | 1.2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.4 | 0.06 | µm | 6e-05 | mm | 1.3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.5 | 0.06 | µm | 6e-05 | mm | 1.4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.6 | 0.06 | µm | 6e-05 | mm | 1.5 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.7 | 0.06 | µm | 6e-05 | mm | 1.6 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.8 | 0.06 | µm | 6e-05 | mm | 1.7 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1.9 | 0.06 | µm | 6e-05 | mm | 1.8 | 0.08 | µm | 8e-05 | mm
[15 empty] | 1 | 0.06 | µm | 6e-05 | mm | 1.9 | 0.08 | µm | 8e-05 | mm
[15 empty] | 2 | 0.06 | µm | 6e-05 | mm | 2 | 0.08 | µm | 8e-05 | mm
[15 empty] | 3 | 0.06 | µm | 6e-05 | mm | 3 | 0.08 | µm | 8e-05 | mm
[15 empty] | 4 | 0.06 | µm | 6e-05 | mm | 4 | 0.08 | µm | 8e-05 | mm
[15 empty] | 5 | 0.06 | µm | 6e-05 | mm | 5 | 0.09 | µm | 9e-05 | mm
[15 empty] | 6 | 0.06 | µm | 6e-05 | mm | 6 | 0.09 | µm | 9e-05 | mm
[15 empty] | 7 | 0.06 | µm | 6e-05 | mm | 7 | 0.09 | µm | 9e-05 | mm
[15 empty] | 8 | 0.06 | µm | 6e-05 | mm | 8 | 0.09 | µm | 9e-05 | mm
[15 empty] | 9 | 0.06 | µm | 6e-05 | mm | 9 | 0.09 | µm | 9e-05 | mm
[15 empty] | 10 | 0.06 | µm | 6e-05 | mm | 10 | 0.09 | µm | 9e-05 | mm
[15 empty] | 11 | 0.07 | µm | 7e-05 | mm | 20 | 0.1 | µm | 0.0001 | mm
[15 empty] | 12 | 0.07 | µm | 7e-05 | mm | 30 | 0.11 | µm | 0.00011 | mm
[15 empty] | 13 | 0.07 | µm | 7e-05 | mm | 50 | 0.13 | µm | 0.00013 | mm
[15 empty] | 14 | 0.07 | µm | 7e-05 | mm
[15 empty] | 15 | 0.07 | µm | 7e-05 | mm
[15 empty] | 16 | 0.07 | µm | 7e-05 | mm
[15 empty] | 17 | 0.07 | µm | 7e-05 | mm
[15 empty] | 18 | 0.07 | µm | 7e-05 | mm
[15 empty] | 19 | 0.07 | µm | 7e-05 | mm
[15 empty] | 20 | 0.07 | µm | 7e-05 | mm
[15 empty] | 21 | 0.07 | µm | 7e-05 | mm
[15 empty] | 22 | 0.07 | µm | 7e-05 | mm
[15 empty] | 23 | 0.07 | µm | 7e-05 | mm
[15 empty] | 24 | 0.07 | µm | 7e-05 | mm
[15 empty] | 25 | 0.07 | µm | 7e-05 | mm
[15 empty] | 50 | 0.09 | µm | 9e-05 | mm
[15 empty] | 75 | 0.1 | µm | 0.0001 | mm
[15 empty] | 100 | 0.12 | µm | 0.00012 | mm
Note on H3: Grade 0
Note on N3: Grade 0
Note on T3: Grade 0
Note on Z3: Grade 0
Note on AF3: Grade 0
Note on D4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on J4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on P4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on V4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AB4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on AH4: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on C5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Linear Hight Gauge
Note on E5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on I5: Nathaphol Boonmee: มาจากค่าความไม่แน่นอนของการวัด Gauge Block
Note on K5: Nathaphol Boonmee: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on O5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on Q5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on U5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on W5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AA5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AC5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน
Note on AG5: ค่าความไม่แน่นอนของการวัดจากใบรับรองผลการสอบเทียบ Gauge Block
Note on AI5: แปลงค่าความไม่แน่ของการวัดให้เป็นหน่วยใช้งาน เดียวกัน