## Data
SP METROLOGY SYSTEM THAILAND | [7 empty] | Certificate No. : |  |  |  | SPR16030112-16 | [5 empty] | Page |  | 1 | of | 1
[8 empty] | Receive Date : |  |  |  | 2016-03-30 |  |  |  | Calibration Date : |  |  |  | 2016-03-30
Equipment Name | [7 empty] | Temp & Humiduty : |  |  |  |  | 23 |  | oC | 50 | %RH
Balance | [7 empty] | Location |  |  |  | In Lab | [6 empty] | On Site
Customer Name : |  |  |  | TTB INDUSTRY
Adress |  | : |  | 88/115
Equipment Name : |  |  |  | Electronic Balance | [6 empty] | Manufacturer : |  |  |  |  | Balance
Model : | Max-Min | [6 empty] | Serial No. : |  |  | 1234567 | [5 empty] | ID No : |  | M-001
Range : | 0 |  | to | 600 |  | g |  |  | Resolution : | 0.001 |  |  |  | g
Overall Inspection |  |  |  |  | Good |  |  | Not Good
Reference Standard |  |  |  | : | SP-SM-001 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-002 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-002-1 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-003 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-003-1 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-003-2 | [7 empty] | Due Date : |  |  | 2018-02-10
Reference Standard |  |  |  | : | SP-SM-004 | [7 empty] | Due Date : |  |  | 2017-12-01
Reference Standard |  |  |  | : | SP-SM-026 | [7 empty] | Due Date : |  |  | 2017-11-22
Calibration Procedure |  |  |  | : | SP-CPM-04-01
Measurement Result
Repeatability ( n = 10 number of measurement ) | [8 empty] | (Max Capacity)
Value | 1 |  |  |  | 2 |  |  | 3 |  |  | 4 |  |  |  | 5 |  |  | STDEV
600
| 6 |  |  |  | 7 |  |  | 8 |  |  | 9 |  |  |  | 10
Departure of indication from nominal Value | [8 empty] | (Step up to 10%)
Nominal | UUC Reading | [13 empty] | Average |  |  | Repeatability |  |  | Error
g | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
0
60
120
180
240
300
360
420
480
540
600
Off - Center Loading | [6 empty] | (1/4 of Capacity) | [13 empty] | Page 2 of 2
Load test | lndicator
150 | X1 |  |  |  | X2 |  |  | X3 |  |  | X4
Center
Front
Back
Left
Right
Diff. Error Max. |  | 0
Calibrated By : |  |  | Mr.Pakapon Nammontree
Mr.Nirut Loha
Mr.Pakapon Nammontree
Mr.Prayoon Topart
Mr.Santi Thonghlor
Mr.Werayut Jampol

## Certificate
Certificate of Calibration
| Certificate Number |  |  | : | SPR16030112-16 |  |  |  |  | Page : 1 of 3
| Customer |  |  | : | TTB INDUSTRY
88/115
| Equipment Name |  |  | : | Electronic Balance
| Manufacturer |  |  | : | Balance
| Model |  |  | : | Max-Min
| Serial Number |  |  | : | 1234567
| ID. Number |  |  | : | M-001
Environmental Conditions
| Ambient Temperature |  |  |  | : | 23 °C ± 2 °C |  | Received Date |  | : | 42459
| Relative Humidity |  |  |  | : | 50% ± 15 % |  | Calibration Date |  | : | 42459
| Location of Calibration |  |  |  | : | In-Lab |  | Recommended Due Date |  | : | 42824
| Calibration Procedure |  |  |  | : | SP-CPM-04-01 |  | Date of Issue |  | : | 42460
Method of Calibration
This certifies that the above instrument was calibrated in compliance with the calibration system
requirement of ISO/IEC 17025:2005 in accordance with reference procedure. Standards used to perform
this calibration are certified by to NIST or equivalent, National metrology institute, Natural physical constants,
consensus standards. The result reported herein apply only to the calibration of the item described above.
All calibrations are performed within manufacture's specifications.The calibration certificate shall not be
reproduced except in full,without written approval of SP Metrology System (Thailand).
| Calibrated by | : | Mr.Pakapon Nammontree |  |  |  | Approved by :
|  |  | Calibration Officer |  |  |  | 3 |  | (Mr. Santi Hankitudomsuk)
Authorized Signatory
Note on Q38: ระบุ ID ประจำตัวพนักงาน

## Report
Certificate Report
| Certificate Number |  | : | SPR16030112-16 |  |  |  | Page : 2 of 3
Reference Standards
| Equipment Name |  |  | Model | Serial No. | Certificate No. | Due. Date
| Standard Weight |  |  | N/A | N/A | SPR16110161-3 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | SPR16110161-4 | 01 Dec 2017
| Standard Weight |  |  | 5 kg | N/A | 60-210057-1 | 10 Feb 2018
| Standard Weight Hook |  |  | N/A | N/A | SPR16110161-5 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | 60-210057-2 | 10 Feb 2018
| Standard Weight |  |  | N/A | N/A | 60-210057-3 | 10 Feb 2018
| Standard Weight |  |  | N/A | N/A | SPR16110161-2 | 01 Dec 2017
| Standard Weight |  |  | N/A | N/A | SS6K3324-16 | 22 Nov 2017
Traceability
This certification is traceable to the International System of Unit maintained at :
-The National Institute of Metrology ( Thailand ), NIMT.

## Result
Result of Calibration
| Certificate No. : |  | SPR16030112-16 | [7 empty] | Page : 3 of 3
Repeatability ( n = 10 number of measurement )
|  | Standard Weight |  | Standard Deviation
|  | 600 | g |  |  | g
| Departure of indication from nominal Value | [7 empty] | Unit : | g
|  | Nominal Value |  |  | UUC Reading |  | Error | Uncertainty ( ± )
|  | 0 |  |  |  |  | #VALUE! | #VALUE!
|  | 60 |  |  |  |  | #VALUE! | #VALUE!
|  | 120 |  |  |  |  | #VALUE! | #VALUE!
|  | 180 |  |  |  |  | #VALUE! | #VALUE!
|  | 240 |  |  |  |  | #VALUE! | #VALUE!
|  | 300 |  |  |  |  | #VALUE! | #VALUE!
|  | 360 |  |  |  |  | #VALUE! | #VALUE!
|  | 420 |  |  |  |  | #VALUE! | #VALUE!
|  | 480 |  |  |  |  | #VALUE! | #VALUE!
|  | 540 |  |  |  |  | #VALUE! | #VALUE!
|  | 600 |  |  |  |  | #VALUE! | #VALUE!
Off - Center Loading
|  | Center |  | 0 |  | g
|  | Front |  | 0 |  | g
|  | Back |  | 0 |  | g
|  | Left |  | 0 |  | g
|  | Right |  | 0 |  | g
|  | Maximum difference |  | 0 |  | g
Measurement Uncertainty
The reported uncertainty of measurement is the expanded uncertainty obtained by multiplying
the standard uncertainty with the coverage factor k = 2.00, providing a level of confidence approximately 95%
- End of Certificate -

## Uncert Budget
Uncertainty Budget of Electronic Balance
Nominal | Uncertainty |  |  |  |  | Air Buoyancy effect |  | Resolution of |  | Repeatability |  | u(δIecc) |  | u(δItemp) |  | Uc | Ui | Veff | K95 | U95%
| Standard Weight. |  |  |  |  | ± 1 ppm 10^6 |  | UUC |  | of UUC
g | g |  |  |  |  | g |  | g |  | g
Value | Value |  |  |  | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Ui | Value | Value | Value | Value | Value
0 |  |  |  |  | 0 | 0 | 0 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0 | 0 | 0.00028868 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
60 |  |  |  |  | 0 | 6e-05 | 3.4641e-05 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00024 | 6.9282e-05 | 0.00029889 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
120 |  |  |  |  | 0 | 0.00012 | 6.9282e-05 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00048 | 0.00013856 | 0.00032762 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
180 |  |  |  |  | 0 | 0.00018 | 0.00010392 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00072 | 0.00020785 | 0.00037059 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
240 |  |  |  |  | 0 | 0.00024 | 0.00013856 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00096 | 0.00027713 | 0.00042348 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
300 |  |  |  |  | 0 | 0.0003 | 0.00017321 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.0012 | 0.00034641 | 0.00048305 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
360 |  |  |  |  | 0 | 0.00036 | 0.00020785 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00144 | 0.00041569 | 0.00054711 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
420 |  |  |  |  | 0 | 0.00042 | 0.00024249 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00168 | 0.00048497 | 0.00061427 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
480 |  |  |  |  | 0 | 0.00048 | 0.00027713 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00192 | 0.00055426 | 0.00068362 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
540 |  |  |  |  | 0 | 0.00054 | 0.00031177 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.00216 | 0.00062354 | 0.00075454 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
600 |  |  |  |  | 0 | 0.0006 | 0.00034641 | 0.0005 | 0.00028868 |  |  | 0 | 0 | 0.0024 | 0.00069282 | 0.00082664 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Temp = | 23
Humidity = | 50
Range | 0-600 g. | Resolution | 0.001 | g.
ρa | 1.1777
ปี 2012
Nominal |  |  |  |  | Variation
Length |  |  |  |  | of GB.
( mm ) |  |  |  |  | ( mm )
1 |  |  |  |  | 0.03
1.005 |  |  |  |  | 0.05
1.01 |  |  |  |  | 0.04
1.02 |  |  |  |  | 0.07
1.03 |  |  |  |  | 0.04
1.04 |  |  |  |  | 0.04
1.05 |  |  |  |  | 0.06
1.06 |  |  |  |  | 0.07
1.07 |  |  |  |  | 0.08
1.08 |  |  |  |  | 0.04
1.09 |  |  |  |  | 0.04
1.1 |  |  |  |  | 0.04
1.2 |  |  |  |  | 0.14
1.3 |  |  |  |  | 0.03
1.4 |  |  |  |  | 0.07
1.5 |  |  |  |  | 0.02
1.6 |  |  |  |  | 0.07
1.7 |  |  |  |  | 0.07
1.8 |  |  |  |  | 0.06
1.9 |  |  |  |  | 0.04
2 |  |  |  |  | 0.03
3 |  |  |  |  | 0.08
4 |  |  |  |  | 0.06
5 |  |  |  |  | 0.07
6 |  |  |  |  | 0.05
7 |  |  |  |  | 0.07
8 |  |  |  |  | 0.01
9 |  |  |  |  | 0.07
10 |  |  |  |  | 0.06
11 |  |  |  |  | 0.06
12 |  |  |  |  | 0.02
13 |  |  |  |  | 0.04
14 |  |  |  |  | 0.05
15 |  |  |  |  | 0.05
16 |  |  |  |  | 0.07
17 |  |  |  |  | 0.04
18 |  |  |  |  | 0.05
19 |  |  |  |  | 0.09
20 |  |  |  |  | 0.08
21 |  |  |  |  | 0.06
22 |  |  |  |  | 0.04
23 |  |  |  |  | 0.04
24 |  |  |  |  | 0.13
25 |  |  |  |  | 0.05
50 |  |  |  |  | 0.14
75 |  |  |  |  | 0.14
100 |  |  |  |  | 0.08
125 |  |  |  |  | 0.03
150 |  |  |  |  | 0
175 |  |  |  |  | 0
200 |  |  |  |  | 0
250 |  |  |  |  | 0
300 |  |  |  |  | 0
400 |  |  |  |  | 0
500 |  |  |  |  | 0
600 |  |  |  |  | 0
700 |  |  |  |  | 0
800 |  |  |  |  | 0
900 |  |  |  |  | 0
1000 |  |  |  |  | 0
Note on N3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การวางน้ำหนักไม่ตรงกลาง
Note on P3: SP Metrology: การประเมินค่าความไม่แน่นอนที่อาจเกิดจาก_การเปลี่ยนแปลงอุณภูมิ
Note on B6: (1ppm*W)/SQRT3
Note on C6: ค่าจากใบรายงานผลการสอบเทียบ STD Weight
Note on H6: (1ppm*W)/SQRT3
Note on J6: ค่าความละเอียดในการอ่านของเครื่องชั่ง
Note on L6: ได้จากการวัดซ้ำ
Note on B22: SP Metrology: The current density of an Air.

## Cert of STD
Weight 10 mg - 500 mg |  |  |  | Weight 1g - 5 kg |  |  |  | Weight 10 g - 20 kg |  |  |  | Weight 5 kg - 20 kg
SP-SM-001 |  |  |  | SP-SM-002 |  |  |  | SP-SM-003 |  |  |  | SP-SM-004
Due Date |  |  |  | Due Date |  |  |  | Due Date |  |  |  | No. | Due Date
42550 |  |  |  | 42550 |  |  |  | 42601 |  |  |  |  | 42706
10 | mg | 0.01 | mg | 1 | g | 0.02 | mg | 10 | g | 0.035 | mg | 1 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg | 2 | g* | 0.03 | mg* | 20 | g | 0.042 | mg | 2 | 5 | kg | 16 | mg
20 | mg | 0.01 | mg* | 2 | g | 0.03 | mg | 50 | g | 0.049 | mg | 1 | 10 | kg | 31 | mg
50 | mg | 0.01 | mg | 5 | g | 0.04 | mg | 100 | g | 0.11 | mg | 2 | 10 | kg | 31 | mg
100 | mg | 0.01 | mg | 10 | g | 0.04 | mg | 200 | g | 0.17 | mg | 1 | 20 | kg | 54 | mg
200 | mg | 0.01 | mg | 20 | g | 0.05 | mg | 500 | g | 0.4 | mg | 2 | 20 | kg | 54 | mg
500 | mg | 0.01 | mg | 20 | g* | 0.05 | mg* | 1 | kg | 1.7 | mg | 3 | 20 | kg | 54 | mg
|  |  |  | 50 | g | 0.1 | mg | 2 | kg | 2 | mg | 4 | 20 | kg | 54 | mg
|  |  |  | 100 | g | 0.2 | mg | 10 | kg | 12 | mg | 5 | 20 | kg | 54 | mg
|  |  |  | 200 | g | 0.4 | mg | 20 | kg | 18 | mg | 6 | 20 | kg | 54 | mg
|  |  |  | 200 | g* | 0.4 | mg* |  |  |  |  | 7 | 20 | kg | 54 | mg
|  |  |  | 500 | g | 2 | mg |  |  |  |  | 8 | 20 | kg | 54 | mg
|  |  |  | 1 | kg | 2 | mg |  |  |  |  | 9 | 20 | kg | 54 | mg
|  |  |  | 1 | kg* | 2 | mg |  |  |  |  | 10 | 20 | kg | 54 | mg
|  |  |  | 2 | kg | 4 | mg |  |  |  |  | 11 | 20 | kg | 54 | mg
|  |  |  | 5 | kg | 10 | mg |  |  |  |  | 12 | 20 | kg | 54 | mg
[12 empty] | 13 | 20 | kg | 54 | mg
[12 empty] | 14 | 20 | kg | 54 | mg
[12 empty] | 15 | 20 | kg | 54 | mg
[12 empty] | 16 | 20 | kg | 54 | mg
[12 empty] | 17 | 20 | kg | 54 | mg
[12 empty] | 18 | 20 | kg | 54 | mg
[12 empty] | 19 | 20 | kg | 54 | mg
[12 empty] | 20 | 20 | kg | 54 | mg
[12 empty] | 21 | 20 | kg | 54 | mg
[12 empty] | 22 | 20 | kg | 54 | mg
[12 empty] | 23 | 20 | kg | 54 | mg
[12 empty] | 24 | 20 | kg | 54 | mg
[12 empty] | 25 | 20 | kg | 54 | mg
[12 empty] | 26 | 20 | kg | 54 | mg
[12 empty] | 27 | 20 | kg | 54 | mg
[12 empty] | 28 | 20 | kg | 54 | mg
[12 empty] | 29 | 20 | kg | 54 | mg
[12 empty] | 30 | 20 | kg | 54 | mg
[12 empty] | 31 | 20 | kg | 54 | mg
[12 empty] | 32 | 20 | kg | 54 | mg
[12 empty] | 33 | 20 | kg | 54 | mg
[12 empty] | 34 | 20 | kg | 54 | mg
[12 empty] | 35 | 20 | kg | 54 | mg
[12 empty] | 36 | 20 | kg | 54 | mg
[12 empty] | 37 | 20 | kg | 54 | mg
[12 empty] | 38 | 20 | kg | 54 | mg
[12 empty] | 39 | 20 | kg | 54 | mg
[12 empty] | 40 | 20 | kg | 54 | mg
[12 empty] | 41 | 20 | kg | 54 | mg
[12 empty] | 42 | 20 | kg | 54 | mg
[12 empty] | 43 | 20 | kg | 54 | mg
[12 empty] | 44 | 20 | kg | 54 | mg
[12 empty] | 45 | 20 | kg | 54 | mg
[12 empty] | 46 | 20 | kg | 54 | mg
[12 empty] | 47 | 20 | kg | 54 | mg
[12 empty] | 48 | 20 | kg | 54 | mg
[12 empty] | 49 | 20 | kg | 54 | mg
[12 empty] | 50 | 20 | kg | 54 | mg
Note on B6: MPE Class F1
Note on G6: MPE Class F1
Note on L6: MPE Class E2
Note on Q6: MPE Class M1
Note on B8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R8: บันทึกวันหมดอายุการใช้งานของSTD ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on B9: Nathaphol Boonmee: Normal Value
Note on D9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on G9: Nathaphol Boonmee: Normal Value
Note on I9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on L9: Nathaphol Boonmee: Normal Value
Note on N9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง
Note on R9: Nathaphol Boonmee: Normal Value
Note on T9: Update Uncertainty ที่ได้จากการสอบเทียบล่าสุดทุกครั้ง